Technical SOP

| Title | OSMOLALITY: Serum and Urine |  |  |  |
| --- | --- | --- | --- | --- |
| Prepared by | Wendell R. McMillan II |  | Date: 03/24/2015 |  |
| Owner | David Rajamonickam |  | Date: 04/02/2015 |  |
| Laboratory Approval | Local Effective Date: |  |  |  |
| Print Name and Title |  | Signature |  | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |  |  |
| Review |  |  |  |  |
| Print Name |  | Signature |  | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	ANALYTICAL PRINCIPLE	3

3.	SPECIMEN REQUIREMENTS	3

4.	REAGENTS	4

5.	CALIBRATORS/STANDARDS	4

6.	QUALITY CONTROL	6

7.	EQUIPMENT and SUPPLIES	8

8.	PROCEDURE	8

9.	CALCULATIONS	9

10.	REPORTING RESULTS AND REPEAT CRITERIA	9

11.	EXPECTED VALUES	10

12.	CLINICAL SIGNIFICANCE	10

13.	PROCEDURE NOTES	11

14.	LIMITATIONS OF METHOD	11

15.	SAFETY	11

16.	RELATED DOCUMENTS	12

17.	REFERENCES	12

18.	REVISION HISTORY	12

19.	ADDENDA	12

# TEST INFORMATION

| Assay | Method/Instrument | Order Code | Local Code |
| --- | --- | --- | --- |
| OSMOLALITY – Serum, Urine | Freezing point measurement – Osmometer Model 3320 | OSMO UOSM | OSMO |

Synonyms/Abbreviations

OSMO

Department

Chemistry

# ANALYTICAL PRINCIPLE

When a solute is dissolved in a solvent, four colligative properties of the solution are changed in a roughly linear response to the amount of solute added.  One of these properties is the freezing point.  The resultant change in freezing point is proportional only to the molar concentration.  In other words, the lowering of the freezing point is a function of the number of particles, molecules, or ions in a solution.  It is upon this property and response that the osmolality of serum or urine is measured with this method.  The concentration of free particles in the serum or urine is determined by measuring the depression in the freezing point since the osmolality is proportional to the freezing point.  Advanced Osmometers utilize high-precision thermistors to sense the sample temperature, control the degree of supercooling and freeze induction, and to measure the freezing point of the sample.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None. |
| Specimen Collection and/or Timing | Serum: Blood should be collected by venipuncture with a minimum of stasis, and the serum should be separated by centrifugation soon after collection. If the serum is not to be analyzed soon after centrifuging, it should be refrigerated or frozen. Heparinized plasma is also satisfactory, but oxalated plasma is not. Urine: Early morning urine should be collected in clean, dry container without preservatives and centrifuged at sufficiently high speed to remove all gross particulate material. If the analysis can not be carried out soon after collection, the specimen should be refrigerated. |
| Special Collection Procedures | See above |
| Other | None |

3.2	Specimen Type & Handling

| Criteria |  |  |  |  |
| --- | --- | --- | --- | --- |
| Type -Preferred | Plasma / Serum, Urine |  |  |  |
| -Other Acceptable | Heparinized plasma |  |  |  |
| Collection Container | Plasma /Serum: BD Vacutainer Urine: Clean, dry container w/o preservatives |  |  |  |
| Volume | Serum/Plasma |  | Urine |  |
| - Optimum | 6.0 mL |  | 3.0 mL |  |
| - Minimum | 3.0 mL |  | 1.0 mL |  |
| Transport Container and Temperature | Clean, plastic vial/container, room temperature, refrigerated or frozen. |  |  |  |
| Stability & Storage Requirements |  | Serum/Plasma |  | Urine |
|  | Room Temperature: | 8 hours |  | 8 hours |
|  | Refrigerated: 2-8ºC | 1 Week |  | 2 Weeks |
|  | Frozen: <20ºC | 1 Month |  | 6 Months |
| Timing Considerations | N/A |  |  |  |
| Unacceptable Specimens & Actions to Take | Oxalated plasma is unacceptable. Request a new sample. |  |  |  |
| Compromising Physical Characteristics | None |  |  |  |
| Other Considerations | Before analysis, refrigerated specimens should be warmed to aid the complete solution of any precipitated substances. |  |  |  |

# REAGENTS

Not applicable

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

| Calibrator | Supplier and Catalog Number |
| --- | --- |
| 50 mOsn/kg Calibration Standard | Advanced Instrument, Inc 3MA005 |
| 850 mOsm/kg Calibration Standard | Advanced Instrument, Inc 3MA085 |
| 1500 mOsm/kg Calibration Standard | Advanced Instrument, Inc 3LA151 |

5.2	Calibrator Preparation and Storage

NOTE:    Date and initial all calibrators upon opening.  Each container must be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibration Standards |  |
| --- | --- |
| Preparation | Ready to use in 2.0mL ampules. |
| Storage/Stability | Store at room temperature until expiration date on label. |

5.3	Calibration Procedure

| Criteria | Special Notations |
| --- | --- |
| Frequency | See User Guide troubleshooting table on Appendix A. |
| Tolerance Limits | See User Guide troubleshooting table on Appendix A. |
| Procedure | Calibration of the Model 3300 is a procedure that requires no adjustment of the instrument by the user. Run the menu-driven calibration program requiring the testing of standards at each of two calibration points. If the repeatability is acceptable on these calibration points, the instrument automatically performs internal calibration. The calibration procedure is as follows. Note that maintaining and loading accurate, uncontaminated standard samples is extremely important. 1. At "Osmometer Ready", press the left soft key, labeled [CALIB] to initiate the calibration process. The calibration program will prompt you to run a sample of the 50 mOsm calibration standard. 2. Run the 50 mOsm calibration standard as you would an actual sample. When the instrument completes each test and reports the results (not necessarily the exact standard value), raise the operating cradle, remove the sampler, and clean the chamber-2. 3. The calibration program will prompt you for a second 50-mOsm calibration standard. Run the second 50 mOsm calibration standard. When the test finishes, your instrument will display a result and prompt you for an 850-mOsmcalibration standard. 4. Run your first 850-mOsm-calibration standard as described above. The program will report a result and request another 850-mOsm-calibration standard. Continue in this manner until the instrument display reads "Calibration complete". The program can require any-where between three and six 850 standard samples, depending on the repeatability of the results. 5. Verify the calibration by running a Clinitrol 290 reference solution before running unknown samples. If the instrument display shows "Calibration Not Complete" after six samples at either calibration level, the calibration procedure has failed and you should carefully repeat the procedure. If you experience "Calibration Not Complete" more than once, you have a repeatability problem and should consult Chapter 4 for troubleshooting and service information. 6. Fill out the OSMO Calibration Log. |
| Dilutions | N/A |
| - Graph Type - Point of Origin - Type of Paper | N/A |

# QUALITY CONTROL

6.1	Controls Used

| Controls | Supplier and Catalog Number |
| --- | --- |
| Clinitrol TM 290 Reference Solution | Advanced Instrument, Inc 3MA029 |
| Urine Chemistry Control Level I And Level II | Bio-Rad - 390 |
| Unassayed Chemistry Control (serum) Level I And Level II | Bio-Rad – 690 X |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

| Control | Clinitrol TM 290 Reference Solution |
| --- | --- |
| Preparation | Ready to use in 2.0 mL ampules. |
| Storage/Stability | Store at room temperature until expiration date on label. |

| Control | Urine Chemistry Control, Level I And Level II |
| --- | --- |
| Preparation | Controls are supplied ready to use. No additional preparation required |
| Storage/Stability | Until expiration date printed on the label when stored at 2oC – 8oC |

| Control | Unassayed Chemistry Control (serum), Level I And Level II |
| --- | --- |
| Preparation | Controls are supplied ready to use. No additional preparation required |
| Storage/Stability | Stable until expiration date when stored unopened at -20oC to -70oC and stable 15 days opened at 2oC – 8oC |

6.3	Frequency

Quality control is performed daily on each shift.

6.4	Tolerance Limits

6.4.1	All quality control results are recorded on the Osmo QC Log.

6.4.2	Any quality control failures will be investigated using the Providence Hospital Lab Quality Control Procedure for a guide.

- The first step in the investigation will be repeating the test.

- Confirm the expiration date and the lot# of all controls

- Check the performance of a known positive control material.  If the proper result is not obtained, discard and retest with fresh product.

For all Rejection criteria see SOP PHQ.020. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the Westgard Rules.

6.5	Review Patient Data

Review patient results for unusual, trends or distributions, such as an unusually high percentage of abnormal results.

6.6	Documentation

Refer to the local policies and procedures for QC documentation and for record retention requirements.

6.7	Quality Assurance Program

    - Refer to   polices and procedures for quality assurance activities applicable to this procedure in QUEST QA manual.

    - The laboratory participates in CAP proficiency testing and in other State and Local agency PT programs when applicable.

    - CAP Proficiency Challenge

    - Each new lot number of reagent must be tested with external control material or previously analyzed sample and result must be recorded in the parallel check log book.

    - All personnel performing the assay must complete annual competency and training process and results must be kept on file.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

Advanced tm Micro-Osmometer, Model 3300 20-microliter sampler

Advanced TM Micro-Osmometer, Model 3MO 20-microliter Sampler

7.2	Equipment

Sampler Plunger Wires

Replacement Sample Probe

7.3	Supplies

Disposable Sampler Tips Model 3300 20-microliter sampler)

Disposable Chamber Cleaners (for Model 3300 20-microliter sampler)

Thermal Paper Roll

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Instrument Set-up Protocol |
| --- | --- |
| 1. | The Advanced TM 20 microliter Sampling System consists of a Sampler with a mounted sample cell. The 3MO measures the sample osmolality inside the tip of the sample cell. Do not inject the sample into the system. |
| 2. | The Advanced TM 20 microliter Sampling System is specially designed for compatibility with the Advanced TM Micro-Osmometer, and should not be used in conjunction with any other laboratory procedure. |

| 8.2 | Test Run |
| --- | --- |
| 1. | Secure and Advanced TM 20 microliter Sampler and disposable sample cell. |
| 2. | Mount one of the sample cells firmly on the Sampler. |
| 3. | Swirl and carefully open an ampule of Clinitrol TM 290 Reference Solution. |
| 4. | Firmly depress the sampler plunger, insert the tip of the mounted sample cell into The ampule about 1/4" below the fluid surface and release the plunger to load 20 uL of sample. |
| 5. | Remove the sampling system from the ampule and blot the sides of the loaded sample cell with absorbent tissue to remove any fluid protruding beyond the end of the tip but be careful not to wick out the sample. The meniscus remaining may be slightly concave. |
| 6. | Remove the camber cleaner form the Sample Port and discard (if no clean, dry chamber cleaner is found in the Sample Port, please refer to the Chamber Cleaning Section below). |
| 7. | Holding the sampler by the barrel and letting the filled sample tip follow the outside groove into the Sample Port within the operative cradle and beneath the cradle top. |
| 8. | To start the test, push the entire operating cradle down until reaches a positive stop; your instrument will run the test for approximately one minute and display the result in the format "Osmolality xxx mOsm. |
| 9. | When the display reads "Osmolality XXXmOsm", record the test results and disengages the Sampling System from the Sample Port. |
| 10. | Firmly insert a clean dry chamber cleaner all the way into Sample Port, rotate it 2-3 times and remove it. Insert another clean, dry chamber cleaner (or the unused end of the first tip). Rotate this 2-3 times, and leave it in place until the next sample. |
| 11. | Remove the used disposable sample cell from the Sampler and discard. |
| 12. | Wipe the tip of the Sampler with a lint-free tissue and install a fresh disposable sample cell. |
| 13. | If the value of the Clinitrol TM 290 Reference Solution is 290+2, proceed to test unknown samples using exactly the same test procedure as for the Reference Solution. |
| 14. | When finished with a series of tests, do not allow the last sample to dry and concentrate in the sample port. Clean and dry the sample chamber as suggested so that the presence of a clean, dry chamber in the Sample Port will assure that the chamber is read for the next test. |

# CALCULATIONS

Not applicable

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

Results read directly in mOsm/Kg from digital display or printout of Osmometer.

10.2	Rounding

N/A

  - Units of Measure

mOsm/kg H2O

10.4	Clinically Reportable Range (CRR)

None established

  - Repeat Criteria

None

# EXPECTED VALUES

11.1	Reference Ranges

Urine    300 – 1200

Serum    280 – 300

11.2	Priority 1 & 2 Limits

None established

11.3	Priority 3 Limit(s)

None established

# CLINICAL SIGNIFICANCE

The term osmolality refers to the osmotic concentration of a fluid. The osmolality of serum, urine, or any other body fluid depends on the number of active ions or molecules in a solution. In laboratory reports, osmolality is expressed as "so many" milliosmoles per kilogram of water (mOsm/kg water). With a standard measurement of osmoles and of milliosmoles for clinical studies, the precise concentration of active solutes in the serum and urine can be calculated. Tests of both serum and urine osmolality can yield important information about a patient's ability to maintain a normal fluid balance status.

Kidneys that are healthy will excrete water in relation to the amount the person consumes. Patients with impaired renal function may not be able to concentrate urine. As a result, the urine osmolality will resemble that of plasma, approximately 290mOsm/Kg. A urine osmolality test may be done on an early morning urine sample as water depletion during the night should concentrate the urine. The test may also be done using multiple timed samples or on a cumulative sample collected over a 24 hour period.

Sodium, blood urea nitrogen, and blood glucose levels are major factors in determining serum osmolality.

Urine osmolality is a more accurate measurement of urine concentration than specific gravity, and urine osmolality can be compared with the serum osmolality to obtain an accurate picture of a patient's fluid balance.

# PROCEDURE NOTES

- FDA Status: Approved

- Validated Test Modifications: None

# LIMITATIONS OF METHOD

14.1	Analytical Measurement Range (AMR)

0 to 2000 mOsm/kg H2O

14.2	Precision

N/A

14.3	Interfering Substances

N/A

  - Clinical Sensitivity/Specificity/Predictive Values

N/A

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- If you drive on company business (in a company car, a rental, or your own personal vehicle) you must comply with the requirements in the "Driver Safety Manual".  Please remember that driving requires your FULL attention and multi-tasking while driving (e.g., talking on a cell phone) is an unacceptable risk that you must not take.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

Not applicable

# REFERENCES

Advanced® Micro-Osmometer Model 3300 User’s Guide 3305 Rev19 042508 2002

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 0 | 2/4/2016 | 6.3 | Updated Frequency | David Rajamonickam | Wendell McMillan |

# ADDENDA

| Addendum | Title |
| --- | --- |
| 1 | OSMO CALIBRATION LOG |
| 2 | QUALITY CONTROL – URINE OSMO |
| 3 | QUALITY CONTROL – SERUM OSMO |

Addendum 1:  OSMO CALIBRATION LOG

| Date | 290 STD (+/- 2) |  |  |  | Calibration/Recalibration |  |  |  | Corrective Action / STD/CAL Accepted  | Tech |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Lot | Exp. | 1st Run | 2nd Run | 50 STD Lot/Exp. | 850 STD Lot/Exp | 2000 STD Lot/Exp. | Calibration Code |  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |
|  |  |  |  |  | / | / | / |  | STD/CAL Accepted  |  |

Calibration (Codes): 	1-New Lot		2-Instrument Related	3-QC Problem	4-Calibration Due

| Weekly review: | Weekly review: | Weekly review: |
| --- | --- | --- |
| Weekly review: | Weekly review: | Monthly review: |

Addendum 2:  QUALITY CONTROL LOG – URINE OSMO

Addendum 3:  QUALITY CONTROL LOG – SERUM OSMO