Technical SOP

| Title | QUICK VUE INFLUENZA TEST |  |
| --- | --- | --- |
| Prepared by | Omar Dualeh | Date: 09/20/2008 |
| Owner | David Rajamonickam | Date: 09/10/2014 |
| Laboratory Approval | Local Effective Date: |  |
|  | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	5

5.	Calibrators/Standards	6

6.	Quality Control	6

7.	Equipment And Supplies	8

8.	Procedure	9

9.	Calculations	10

10.	Reporting Results And Repeat Criteria	10

11.	Expected Values	11

12.	Clinical Significance	11

13.	Procedure Notes	11

14.	Limitations Of Method	12

15.	Safety	13

16.	Related Documents	13

17.	References	14

18.	Revision History	14

19.	Addenda	14

# TEST INFORMATION

| Assay | Method/Instrument | Order Code | Local Code |
| --- | --- | --- | --- |
| QuickVue Influenza Test | A Lateral-Flow Immunoassay | INFP | INFP |

Synonyms/Abbreviations

N/A

Department

Chemistry

# ANALYTICAL PRINCIPLE

The QuickVue Influenza Test allows for the rapid detection of influenza type A and type B antigens directly from nasal swab; nasal wash and/or nasal aspirate specimens. The test is intended for use as an aid in the rapid diagnosis of acute influenza virus infection. The test is not intended to detect Influenza C antigens. Negative test results should be confirmed by cell culture. Influenza antigens may be detected in clinical specimens by immunoassay. The QuickVue Influenza Test is a lateral-flow immunoassay using highly sensitive monoclonal antibodies that are specific for influenza antigens. The test is specific for influenza types A and B antigens with no known cross-reactivity to normal flora or other known respiratory pathogens.

The QuickVue Influenza Test involves the extraction of influenza A and B viral antigens. The patient specimen is placed in the Extraction Reagent Tube, during which time the virus particles in the specimen are disrupted, exposing internal viral nucleoproteins. After extraction, the Test Strip is placed in the Extraction Reagent tube where nucleoproteins in the specimen will react with the reagents in the Test Strip. If the extracted specimen contains influenza antigens, a pink-to-red Test Line along with a blue procedural Control Line will appear on the Test Strip indicating a positive result. If influenza type A or B antigens are not present, or are present at very low levels, only a blue procedural Control Line will appear.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | None |
| Specimen Collection and/or Timing | Samples must be tested as soon as possible after collection. |
| Special Collection Procedures | NASAL SWAB: - Insert the sterile swab into the nostril that presents the most secretion under visual inspection. - Using gentle rotation, push the swab until resistance is met at the level of the turbinates (less than one inch into the nostril). - Rotate the swab a few times against the nasal wall. |
| Special Collection Procedures, cont . | NASAL WASH: - With the patient's head hyper-extended. - Instill about 2.5 mL of sterile, normal saline into one nostril with a syringe. - To collect the wash, place a clean dry specimen container directly under the nose with slight pressure on the upper lip. - Tilt the head forward and allow the fluid to run out of the nostril into the specimen container. - Repeat for the other nostril and collect the fluid into the same specimen container. NASAL ASPIRATE: - Fill an aspiration bulb or bulb syringe with 2.5 mL of sterile, normal saline (depending on the size of the child). - Instill into one nostril while the head is tilted back. - Release the pressure on the bulb to aspirate the specimen back into the bulb. - Transfer the specimen into a clean, dry specimen container. - Repeat the process for the child's other nostril and transfer the specimen into the same specimen container. |
| Other | None |

3.2	Specimen Type & Handling

| Criteria |  |  |
| --- | --- | --- |
| Type -Preferred | Nasal Swab |  |
| -Other Acceptable | Nasal Aspirate and Wash |  |
| Collection Container | Sterile FOAM Tipped Applicator Swap (Provided with the Kit) |  |
| Volume - Optimum | N/A |  |
| - Minimum | N/A |  |
| Transport Container and Temperature | Sterile FOAM Tipped Applicator Swap (Provided with the Kit) 15° to 30°C: |  |
| Stability & Storage Requirements | Room Temperature 22° to 28°C: | 8 hours |
|  | Refrigerated 2-8°C: | 8 hours |
|  | Frozen: | Note: Only Nasal Aspirate and Wash can be stored at (-70 °C) up to one month |
| Timing Considerations | Samples must be tested as soon as possible after collection. |  |
| Unacceptable Specimens & Actions to Take | - Dry swab or frozen specimens are unacceptable. Call and request another sample. - Do not use any other type of transport media to collect, store, or transfer samples. If received please ask for recollection. |  |
| Compromising Physical Characteristics | N/A |  |
| Other Considerations | N/A |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

4.1	Reagent Summary

| Reagents / Kits | Supplier & Catalog Number |
| --- | --- |
| QuickVue Influenza A+B Test | Quidel Corporation, Cat. #20183 |

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| Test Strips | Quidel Corporation, 0569104 | 25 Strips |
| Extraction Reagent Solution | Quidel Corporation, 1013100 | 25 Solutions |
| Extraction Tubes | Quidel Corporation, 1051600 | 25 Tubes |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

| Reagent | Extraction Reagent Solution: (250 pL each) Normal Saline |
| --- | --- |
| Container | Glass Tubes |
| Storage | 15° to 30°C |
| Stability | Stable till expiration date printed on the outer box. |
| Preparation | Ready for use. |

| Reagent | Extraction Tubes: Lyophilized buffer with detergents and reducing agents. |
| --- | --- |
| Container | Glass Tubes |
| Storage | 15° to 30°C |
| Stability | Stable till expiration date printed on the outer box. |
| Preparation | Dispense all of the Extraction Reagent solution in to extraction Tube, and gently swirl the Extraction Tube to dissolve its contents. |

| Reagent | Test Strips: Mouse monoclonal anti-influenza A and B antibodies |
| --- | --- |
| Container | Individually packaged PROTECTED FOIL POUCH |
| Storage | 15° to 30°C |
| Stability | Stable till expiration date printed on the outer box |
| Preparation | Ready for use. |

| Reagent | POSITIVE Influenza Type A Control Swab POSITIVE Influenza Type B Control Swab NEGATIVE Control Swab |
| --- | --- |
| Container | Swab |
| Storage | 15° to 30°C |
| Stability | Stable till expiration date printed on the outer box. |
| Preparation | Ready for use. |

# CALIBRATORS/STANDARDS

N/A

# QUALITY CONTROL

6.1	Controls Used

| External Controls | Supplier and Catalog Number |
| --- | --- |
| POSITIVE Influenza Type A Control Swab: Swab is coated with non-infectious recombinant Influenza A antigen. | Kit component. Refer to section 4.1. |
| POSITIVE Influenza Type B Control Swab: Swab is coated with non-infectious recombinant Influenza B antigen. | Kit component. Refer to section 4.1. |
| NEGATIVE Control Swab: Swab is coated with formalin-inactivated, non-infectious Streptococcus C antigen | Kit component. Refer to section 4.1. |

| Internal Controls | Supplier and Catalog Number |
| --- | --- |
| The QuickVue Influenza Test contains built-in positive and negative control features. See section 6.4.1. No preparation required. | Part of Kit component. Refer to section 4.1. |

6.2	Control Preparation and Storage

NOTE:    Date and initial all controls upon opening.  Each container should be labeled with (1) substance name, (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

Refer to section 4.2.

6.3	Frequency

Internal Quality Control is assessed and recorded for each test.  External Quality Control is tested each day of testing,  each new lot received and each new shipment of kits, regardless of the lot number.

- Positive and negative external controls are run and documented on each day of testing.

- Positive and negative external controls are run and documented with each new kit lot number.

- Positive and negative external controls are run and documented with each new shipment is received.

- Positive and negative external controls are run and documented if the test kit storage falls outside of 15°-30°C.

  - Tolerance Limits

    - INTERNAL QUALITY CONTROL:

The QuickVue Influenza Test contains built-in procedural control features. The daily quality control is to document the built-in procedural controls for each test run. The two-color result format provides a single interpretation for positive and negative results. The appearance of a sufficient capillary flow has occurred and the functional integrity of the Test Strip was maintained.

NOTE: IF THE BLUE PROCEDURAL CONTROL LINE FAILS TO DEVELOP AT 10 MINUTES, THE TEST IS CONSIDERED INVALID.

A built-in negative control is provided by the clearing of red background, verifying that the test has been performed correctly. Within 10 minutes, the result area should be white to light pink and allow the clear interpretation of the test result.

NOTE: IF BACKGROUND COLOR APPEARS AND INTERFERES WITH INTERPRETATION OF THE TEST RESULT, THE RESULT IS CONSIDERED INVALID.

6.4.2   EXTERNAL QUALITY CONTROL:

External Controls are run upon each day of testing / new lot number of test kit and/or each shipment of test kits. The external controls are used to demonstrate that the reagents and assay procedure performed properly. External positive and negative Control Swabs are supplied in the Test Kit and should be tested using the swab procedure.

    - If the controls do not perform as expected, repeat the test and/or contact a Laboratory Supervisor for assistance.

    - DO NOT TEST PATIENTS UNTIL EXPECTED QC RESULTS ARE OBTAINED.

For all Rejection criteria see SOP PHQ.020. Runs where QC rules are violated will be rejected. It is the technologists’ responsibility to evaluate the Westgard Rules.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

6.6	Documentation

Control results are automatically filed. Refer to Record Retention SOP for record retention requirements.

  - Review of QC

6.7.1.	Weekly and Monthly QC must be reviewed by the supervisor on weekly and monthly basis.

    - Monthly QC must be presented to Medical Director or designee for review and signature.

  - Quality Assurance Program

    - Refer to corporate and local polices and procedures for other quality assurance activities applicable to this procedure.

    - Each new lot number of reagent must be tested with external control materials and/or previously analyzed samples.  Performance of the new lot must be equivalent to the previous lot and must be document in the parallel check book.

    - All persons performing this assay must complete training and are assessed annually for competency, the results are documented on file.

    - This procedure must be incorporated into the departmental competency assessment program.

    - The laboratory participates in CAP proficiency testing.

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

N/A

7.2	Equipment

N/A

7.3	Supplies

N/A

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Test Run |
| --- | --- |
| 1. | NASAL SWABS NOTE: Samples must be tested as soon as possible after collection. Do not use any type of transport media to store or transfer Samples. Samples may be stored refrigerated (2-8°C) In a clean, dry, closed container for up to 0 n e (1) hour. 1. Dispense all of the Extraction Reagent solution from the yellow-capped Reagent Tube, gently swirl the Extraction Tube to dissolve its contents. 2. Place the patient swab sample into the Extraction Tube. Roll the swab at least three (3) Times while pressing the head against the bottom and side of the Extraction Tube. 3. Roll the swab head against the inside of the Extraction Tube as you remove it. Dispose of the swab into the biohazard waste container. 4. Place the Test Strip into the Extraction Tube with the arrows on the Test Strip pointing down. Do not handle Or move the Test Strip until the test is complete and ready for reading. 5. Read result at ten (10) minutes. Some positive results may appear sooner. |
| 2. | NASAL WASH/NASAL ASPIRATE NOTE: Samples must be tested as soon as possible after collection. Samples may be stored refrigerated (2-8°C) in a clean, dry, closed container for up to one (1) hour. 1. Fill the dropper to the top/uppermost notch with nasal wash or aspirate sample. 2. Add entire contents of the dropper to the Extraction Tube. Swirl the Extraction Tube gently to dissolve its contents. 3. Place the Test Strip into the Extraction Tube with the arrows on the Test Strip Pointing down. Do not handle or move the Test Strip until the test is complete and ready for reading. 4. Read result at ten (10) minutes. Some positive results may appear sooner. |

| 8.2 | Special Handling |
| --- | --- |
| 1. | If at ten (10) minutes, the blue procedural Control Line does not appear, even if any shade of pink-to-red Test Line appears, the result is considered invalid. If the test is invalid, a new test should be performed with a new patient sample and a new Test Strip. |

# CALCULATIONS

N/A

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

POSITIVE RESULT:

At ten minutes, ANY shade of pink-to-red Test Line forms AND the appearance of a blue procedural Control Line indicates a positive result for the presence of influenza A and/or B viral antigen.

NEGATIVE RESULT:

At ten minutes, the appearance of ONLY the blue procedural Control Line indicates the sample is negative for influenza A and! or B viral antigen. A negative result should be reported as a presumptive negative for the Presence of influenza antigen.

10.2	Rounding

N/A

10.3	Units of Measure

N/A

10.4	Clinically Reportable Range (CRR)

N/A

  - Repeat Criteria and Resulting

| IF the result is … | THEN… |
| --- | --- |
| If at ten (10) minutes, the blue procedural Control Line does not appear, even if any shade of pink -to-red Test Line appears, the result is considered invalid. If the test is invalid | A new test should be performed with a new patient sample and a new Test Strip. |

| Results | Result Code |
| --- | --- |
| Positive | !P |
| Negative | !N |
| Invalid | !INV |

# EXPECTED VALUES

11.1	Reference Ranges

Negative

11.2    Priority 1, 2 and 3 Limits

None established

# CLINICAL SIGNIFICANCE

Influenza is a highly contagious, acute viral infection of the respiratory tract. The causative agents of the disease are immunologically diverse, single-strand RNA viruses known as influenza viruses. There are three types of influenza viruses: A, B, and C. Type A viruses are the most prevalent and are associated with the most serious epidemics. Type B viruses produce a disease that is generally milder than that caused by type A. Type C viruses have never been associated with a large epidemic of human disease. Both type A and B viruses can circulate simultaneously, but usually one type is dominant during a given season.

# PROCEDURE NOTES

FDA Status: Approved

Validated Test Modifications: None

# LIMITATIONS OF METHOD

- The contents of this kit are to be used for the qualitative detection of Influenza A and B antigen from nasal swab, nasal wash and nasal aspirate. This test does not differentiate between influenza types A and B.

- Test Results must be evaluated in conjunction with other clinical data available to the physician.

- A negative test result may occur if the level of antigen in a sample is below the detection limit of the test, or from improper sample collection.

- Negative test results are not intended to rule-out other non-influenza viral infections.

Accuracy Rate:

Nasal Swabs - 87%

Nasal Wash! Aspirate - 94%

14.1	Analytical Measurement Range (AMR)

N/A

14.2	Precision

A panel containing two different levels of influenza A antigen (Johanneburg/82/96; weak positive and strong positive) and two different levels of influenza B antigen (Harbin/7/94; weak positive and strong positive) were repeated five times with a single lot of QuickVue influenza test on three different days. One hundred percent (100%) accuracy was obtained for all specimens tested.

14.3	Interfering Substances

Whole blood, and several over-the-counter (OTC) products and common chemicals were evaluated and did not interfere with the QuickVue Influenza A+B test at the levels tested: whole blood (2%); three OTC mouthwashes (25%); three OTC throat drops (25%); three  OTC nasal sprays (10%); 4-Acetamidophenol (10 mg/mL); Acetylsalicylic Acid (20 mg/mL); Chlorpheniramine (5 mg/mL); Dextromethorphan (10 mg/mL); Diphenhydramine (5 mg/mL); Ephedrine (20 mg/mL); Guaiacol glyceryl ether (20 mg/mL); Oxymetazoline (10 mg/mL); Phenylephrine (100 mg/mL); and Phenylpropanolamine (20 mg/mL).

  - Clinical Sensitivity/Specificity/Predictive Values

QuickVue influenza Nasal Swab Results versus Culture (by Age Group)

| <18 years of age; N=16 |  |  | >/= 18 years of age; N=214 |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensitivity | Specificity | Accuracy | Sensitivity | Specificity | Accuracy |
| 96% (24/25) | 92% (33/36) | 93% (57/61) | 66% (55/83) | 97% (127/131) | 85% (182/214) |

QuickVue influenza Nasal Wash/Nasal Aspirate Results versus Culture (by Age Group)

| <18 years of age; N=22 |  |  | >/= 18 years of age; N=74 |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensitivity | Specificity | Accuracy | Sensitivity | Specificity | Accuracy |
| 80% (4/5) | 94% (16/17) | 91% (20/22) | 82% (18/22) | 100% (52/52) | 95% (70/74) |

# SAFETY

You, the employee, have direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

- Warnings of other specific hazards are noted in this procedure.  Please comply with the requirements to reduce your risk of injury."

Report all accidents and injuries to your supervisor or the Environmental, Health and Safety Coordinator.

# RELATED DOCUMENTS

Current Package Insert

# REFERENCES

- Murphy, B.R. and R.G. Webster, 1996. Orthomyxoviruses, pp. 1397-1445. In: Fields

Virology, 3rd Edition, B.N. Fields, D.M. Knipe, P.M. Howley, (eds.) Lippincott-Raven,

Philadelphia.

- Organon Teknika Simplate Manufacturer’s Package Insert, Issued November, 1989

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| .000 | 9/15/10 | Page 1 | Updated Owner | A. Foster | R. Levy |
|  |  | 8.0 | Updated Procedure to reflect removal of package inserts |  |  |
|  |  | 14.2 | Added precision studies results |  |  |
|  |  | 14.2 | Added Tables to show comparative studies reflecting ; sensitivity, specificity and accuracy |  |  |
|  |  | 16.0 | Updated related documents to include current product insert |  |  |
|  |  | 19.0 | Updated Addenda to reflect removal of addendum 1 |  |  |
| .001 | 1/18/2011 | Page 1 | Updated Laboratory Approval | A. Foster | R. Levy |
|  |  | 10.5 | - Added Result codes |  |  |
| .002 | 06/24/2011 | 6.3 | Quality Control frequency updated. | A Foster | R. Levy |
| .003 | 9/7/2011 | Page 1 | Updated Laboratory approval | A. Foster | R. Levy |
|  |  | 6.4 | Added Reference of PHQ.020 to tolerance limit |  |  |
| .004 | 10/31/2012 | 3.2 | Changed Room Temperature | W. McMillan | R. Levy |
|  |  | 6.7.2 | Added “or designee” |  |  |
| .005 | 9-10-2014 | Page 1 | Updated Owner, BU Name in Header, Removed Annual from the Review table, Added PRO in SOP Number in footer and removed leading zeros to the Version. | David Rajamonickam | Wendell McMillan |
| 6 | 10-6-2016 | 6.3 | Updated Frequency | David Rajamonickam | Wendell McMillan |

# ADDENDA

| Addendum | Title |
| --- | --- |