- Technical SOP

| Title | Coulter LH750 Operation For Complete Blood Count and Reticulocyte Automated Tests |  |
| --- | --- | --- |
| Prepared by | Solomon Tsehay | Date: 05/25/2006 |
| Owner | Wendell R. McMillan II | Date: 03/09/2012 |
| Laboratory Approval | Local Effective Date: |  |
| Printed Name | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| Review |  |  |
| Print Name | Signature | Date |

TABLE OF CONTENTS

1.	TEST INFORMATION	2

2.	ANALYTICAL PRINCIPLE	3

3.	SPECIMEN REQUIREMENTS	4

4.	REAGENTS	6

5.	CALIBRATORS/STANDARDS	6

6.	QUALITY CONTROL	9

7.	EQUIPMENT and SUPPLIES	14

8.	PROCEDURE	14

9.	CALCULATIONS	16

10.	REPORTING RESULTS AND REPEAT CRITERIA	16

11.	EXPECTED VALUES	17

12.	CLINICAL SIGNIFICANCE	17

13.	PROCEDURE NOTES	18

14.	LIMITATIONS OF METHOD	24

15.	SAFETY	25

16.	RELATED DOCUMENTS	25

17.	REFERENCES	25

18.	REVISION HISTORY	26

19.	ADDENDA	27

# TEST INFORMATION

| Assay | Method/Instrument | Local Code |
| --- | --- | --- |
| Hemogram (WBC, RBC, HGB, HCT, MCV, MCH, MCHC ,RDW PLT, MPV) | Coulter Automated Hematology Analyzer, LH750 | CBCND |
| Hemogram & diff (WBC,RBC,HGB,HCT,MCV,MCH,MCHC,RDW, PLT, MPV, differential) |  | CBC |
| Differential count only |  | DIFF |
| Platelet Count |  | PLTC |
| Reticulocyte Count |  | RETA |

| Abbreviation | Term | Abbreviation | Term |
| --- | --- | --- | --- |
| WBC | White Blood Cell | MCHC | Mean Corpuscular Hemoglobin Concentration |
| RBC | Red Blood Cell |  |  |
| HGB | Hemoglobin | RDW | Red Cell distribution Width |
| HCT | Hematocrit | DIFF | Differential Count |
| MCV | Mean Cell Volume | PLT | Platelet |
| MCH | Mean Corpuscular Hemoglobin | MPV | Mean Platelet Volume |
| RETIC | Reticulocyte Count |  |  |
| Department |  |  |  |
| Hematology |  |  |  |

# ANALYTICAL PRINCIPLE

CBC

The Coulter principle employs electronic counting and sizing of particles using the Hematology analyzers.  WBC Differential analysis and classification are based on simultaneous measuring of cell volume, high frequency conductivity and laser light Scatter.  Hemoglobin, released by hemolysis, is converted to a stable cyanide containing pigment and measured by photometric absorbance.

Reticulocyte

Red blood cell (RBC) RNA is stained with the vital stain new methylene blue.  The dye precipitates the RNA found in reticulated RBC.  Hemoglobin is removed from the RBC leaving the precipitated dye-RNA complex by adding a sulfuric acid solution.  Reticulocyte percent and number are measured by analysis of the total RBC population for volume, conductivity and light scatter.

2.1	Determination of Parameters

| Type of Measurement | Parameter | Source of Data |
| --- | --- | --- |
| Direct | RBC (Red Blood Cell) | Coulter principle |
|  | WBC (White Blood Cell) | Coulter principle |
|  | HGB (Hemoglobin) | Photometric absorbance |
|  | MCV (Mean Cell Volume) | Coulter principle |
|  | PLT (Platelet ) | Coulter principle |
|  | Automated Differential, five-part | Light scatter, volume & conductivity (VCS technology) |
|  | RET% (Reticulocyte) | VCS Technology |
| Derived from Histograms | RDW (RBC Distribution Width) | RBC Histogram |
|  | MPV (Mean Platelet Volume) | PLT Histogram |
|  | NRBC% | WBC Histogram and VCS technology |
| Calculated | HCT (Hematocrit) | HCT = RBC x MCV 10 |
|  | MCH (Mean Corpuscular Hemoglobin) | MCH = HGB x 10 RBC |
|  | MCHC (Mean Hemoglobin Concentration) | MCHC = HGB x 100 HCT |
|  | DIFF # parameters | DIFF as % x WBC (i.e. 0.77 x 5800) |
|  | Absolute Neut | (Neut% + Band% + Meta% + Myelo% + Promyelo %) X WBC |
|  | Absolute Lymph | (Lymph% + Reactive Lymph %) X WBC |
|  | Absolute Monocytes | Mono% X WBC |
|  | Absolute Eosinophils | Eos% X WBC |
|  | Absolute Basophils | Baso% X WBC |

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | Not applicable |
| Specimen Collection and/or Timing | None defined |
| Special Collection Procedures | None defined |

- 3.2	Specimen Type & Handling

| Criteria |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Type -Preferred -Other Acceptable | K3EDTA or K2EDTA Whole Blood Sodium Citrate – for platelet counts only |  |  |  |  |  |  |  |
| Collection Container | Lavender Top Tube Tri-Potassium or Di-Potassium EDTA Anticoagulant |  |  |  |  |  |  |  |
| Volume | Tube |  |  |  | Minimum |  |  | Optimum |
|  | K3EDTA or K2EDTA (non-pediatric) |  |  |  | 1.0mL |  |  | Full tube |
|  | Pediatric K3EDTA or K2EDTA tube |  |  |  | 0.5mL |  |  | Full tube |
|  | Microtainer tube |  |  |  | 0.5mL |  |  |  |
| Transport Container and Temperature | Same as above. Transport at room temperature or refrigerated. |  |  |  |  |  |  |  |
| Stability & Storage Requirements | Room Temperature (18-25ºC): 48 hours Refrigerated: After analysis, specimens are stored for a minimum of 2 days at 2-8C. Frozen (-20ºC and below): Not Acceptable |  |  |  |  |  |  |  |
| Timing Considerations | N/A |  |  |  |  |  |  |  |
| Specimen Quality Table | Condition | Slight |  | Moderate |  |  | Marked |  |
|  | Icterus | OK |  | OK |  |  | Orange-Brown = see section 13.8 |  |
|  | Hemolysis | Slight pink OK |  | Pink OK |  |  | Cherry Red Unacceptable |  |
|  | Lipemia | OK |  | OK |  |  | Milky = see section 13.8 |  |
| Other Interfering Specimens Factors | CBC Indicated by CBC results (see Addendum 2) Fibrin, bacterial contamination, platelet clumps, abnormal proteins, cold agglutinins, extreme temperature conditions, resistant hemoglobin, abnormal chemistries and specimens older than 48 hours. RETIC Extreme temperatures, other erythrocyte inclusions that stain by new methylene blue dye, some hemoglobinopathies (SS, SC), and specimens older than 72 hours. |  |  |  |  |  |  |  |
| Actions to Take for | Condition |  | Code |  |  | Comment |  |  |
| Rejected Specimens Message Codes & Notes | QNS (Less than the minimum volume in Section 3.2) |  | QNS |  |  | Quantity not sufficient to perform test. Notify caregiver. (Document in the LIS) |  |  |
|  | Clotted |  | CLT |  |  | Specimen is clotted, unable to perform test. Notify caregiver. (Document in the LIS) |  |  |
|  | Spurious results that will not duplicate |  | INT or UNSAT |  |  | Possible interfering substance. or Unsatisfactory specimen. Notify caregiver. (Document in the LIS) |  |  |
|  | Gross hemolysis |  | HMT |  |  | Markedly hemolyzed. Notify caregiver. (Document in the LIS) |  |  |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information.

- 4.1	Reagent Summary

| Reagents | Stability (Opened) | Supplier & Catalog Number | Size |
| --- | --- | --- | --- |
| Lyse S III diff | 60 days | Beckman Coulter – Cat # 8546796 | 5 Liter |
| Coulter Clenz | 90 days | Beckman Coulter – Cat # 8546931 | 10 Liter |
| LH 700 Series Diluent | 60 days | Beckman Coulter - PN # 8547194 | 20 Liter |
| LH 700 Series PAK | 60 days | Beckman Coulter - PN # 8547195 | N/A |
| LH 700 Series Retic PAK | 60 days | Beckman Coulter - PN # 8547196 | N/A |

- 4.2	Reagent Preparation and Storage

NOTE:  Date and initial all reagents upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation.

Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

All reagents are stored at room temperature and stable (when unopened) until expiration date on label. All reagents are received ready for use.

# CALIBRATORS/STANDARDS

- 5.1	Calibrators/Standards Used

| Calibrator | Supplier & Catalog Number |
| --- | --- |
| S-Cal (for RBC Parameters) | Beckman Coulter- LH750 PN 7508116-A |

Caution: Calibrator contains sodium azide (<0.1%). Contains potentially biohazardous materials. Use with good laboratory practices to avoid skin/eye contact or ingestion.  Consult MSDS for a complete list of hazards.

- 5.2	Calibrator Preparation and Storage

NOTE:  Date and initial all calibrators upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech (6) any special storage instructions; check for visible signs of degradation.

| Calibrator | Coulter® S-Cal® Calibrator Kit |
| --- | --- |
| Preparation | Bring to room temperature prior to testing. Use within one hour. |
| Storage/Stability | Store refrigerated (2-8C). Use within expiration date from manufacturer. For further details refer to the package insert. Check for visible signs of degradation prior to use, i.e. color change or clotting. |

- 5.3	Calibration Procedure

| Criteria | Special Notations |  |
| --- | --- | --- |
| Frequency | At least every six months, and when indicated by the following: - New installation. - New set of apertures, BSV and electronics are installed. - When recommended by the manufacturer & when ambient temperature changes more than 10 F from calibrating temperature. - When multiple levels of commercial controls are consistently out or biased for one or more parameters. NOTE: Calibration must be verified for sampling modes, opened and closed (cap-piercer). When any parameter is adjusted, the change must be made or verified for both sampling modes. |  |
| Calibration Preparation | Before Calibration: - Has instrument had PM in the last 6 months (Consult Supervisor) - Verify all routine maintenance is up-to-date. - Clean the Baths. - Perform shutdown. - Ensure you have sufficient supply of reagents to complete the procedure. - Perform Startup. - Perform Reproducibility: 1-If the CV% for any parameter is greater than those listed; you might have an instrument problem. Call your Coulter Representative. 2-Review each parameter for trending (a gradual and consistent increase or decrease in values). If you think a trend exists, you might have an instrument problem; call your Coulter Representative. - Perform Carryover Check: (Validate carryover (%) for each parameter against manufacturer acceptability guidelines; if exceeded, call your Coulter Representative). - If all of the above are determined to be acceptable, then proceed with S-Calibration. Otherwise, correct the deficiency and repeat the reproducibility & carryover procedures. Follow the S-Cal preparation, handling, and procedural instructions. |  |
| Tolerance Limits | IF … | Then … |
|  | - results fall within the specifications, - calibration status is displayed as acceptable - Quality Control (QC) values are within acceptable limits. | Proceed with analysis. |
|  | - results fall outside of specifications - the calibration status is displayed as failed - the QC values are outside acceptable limits. | Troubleshoot the assay and/or instrument and repeat the calibration. |
| Procedure | - Follow instructions in the S-Cal package. - For step-by-step procedure, refer to the LH750 Calibration Screen Help Procedure. |  |
|  | Calibration Verification of CBC Parameters with S-Cal Immediately following the calibration steps outlined and the instrument is functioning properly; - Obtain 5 normal, fresh whole-blood specimens. - Cycle each of the 5 samples two consecutive times in CBC Automatic aspiration mode. - Review results and adjust CBC parameters as needed. - Verify calibration by cycling each level of COULTER 5C series cell control in Automatic mode. |  |

5.4	Documentation

All Calibration and/or Calibration Verification processes (with commercial material) are documented.  Calibration and/or Calibration Verification processes are signed and dated by performing staff.  Calibration and/or Calibration Verification documents are reviewed, dated, and signed by supervisory staff.  Calibration and/or Calibration Verification documents are QC documents and maintained according to guidelines published in the Quest Diagnostics Records Management Program Reference Guide.

# QUALITY CONTROL

6.1	Controls Used

A.  Commercial Quality Control Materials

Caution:  Controls contain sodium azide (<0.1%) and potential biohazardous materials.  Use with good laboratory practices to avoid skin/eye contact or ingestion.  Consult MSDS for a complete list of hazards.

| Control | Supplier & Product Number |
| --- | --- |
| 5C Abnormal I 5C Normal 5C Abnormal II | Beckman Coulter # 7547116 4 x 3.3 mL each level |
| RETIC – C (Level I, II, III) | Beckman Coulter # 7547125 3 x 3.3 mL each level |
| Latron 1 (primer) | Beckman Coulter # 7546915 5 x 16 mL each |
| Latron 2 (control) | Beckman Coulter # 7546914 5 x 16 mL each |

- 6.2	Control Preparation and Storage

NOTE:  No control preparation is necessary.  Date and initial all controls upon opening.  Each container should be labeled with (1) substance name, (2) lot number, (3) date of preparation, (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation. Follow the procedure in the Reagent Parallel testing procedure for checking new lots or shipments of reagents.

| Control | Storage & Stability |
| --- | --- |
| 5C Abnormal I 5C Normal 5C Abnormal II RETIC-C (Levels I, II, III) | - Store refrigerated at 2-8C. - Bring to room temperature prior to testing. - Observe expiration date. - Open vial stability = 13 days Coulter 5C Cell Control = 15 days Retic-C Cell Control |
| Latron 1 (primer) | - Store at 2-30C. - Bring to room temperature prior to testing. - Observe expiration date. - Open vial stability = 30 days |
| Latron 2 (control) | - Store at 2-30C. - Bring to room temperature prior to testing. - Open vial stability = 30 days |
| Patient Control | Follow the specimen guidelines for EDTA specimens listed in Section 3.0 |

- 6.3	Frequency

A.	All three commercial control levels must be tested ach shift.  Data from this control run is part of the Coulter IQAP program.

B.	Latron Control will be tested as part of start-up procedure only. Refer to Coulter LH750 help screen for Control for procedure.

C.	Multi-mode sampling

- Each day both closed-mode and open-mode sampling using the commercial control must be performed on each shift, as per CLIA and CAP requirements.

- Typically, most testing is done in closed mode.

- When samples are tested using the closed-mode, a control (Commercial) must be tested for each shift.

D.	New Reagent or diluent: One level of Commercial control will be run when all new LH-750 reagents and diluents are changed. A background check MUST be done and the report stored in the reagent change QC folder in hematology.

E.	See Addendum 4 for the Daily Quality Control Schedule for the Beckman Coulter LH750.

- 6.4	Tolerance Limits

The laboratory’s QC program is set up with mean values provided in the package insert for the respective lot# of QC that have been verified per laboratory procedure. For tracking QC in the LH database, Standard Deviations (SDs) used for acceptable limits must not exceed the Max SD or the SDc (determined from the Coulter QC Range), whichever is greater.

| QC Level | Parameter | Max. Total Allowable Error | Max CV, % | Max SD | SDc = Coulter Range / 3 |
| --- | --- | --- | --- | --- | --- |
| Abnormal I | WBC | +/- 15% | 3.0 | 0.60 | 0.37 |
|  | RBC | +/- 6% | 1.2 | 0.05 | 0.04 |
|  | Hemoglobin | +/- 7% | 1.4 | 0.18 | 0.13 |
|  | Hematocrit | +/- 6% | 1.4 | 0.52 | 0.63 |
|  | MCV | +/- 6% | 1.2 | 1.0 | 1.0 |
|  | Platelet Count | +/- 25% | 5.0 | 17 | 13 |
|  | Neutrophils (%) | +/- 2SD | 1.5 | 1.0 | 1.67 |
|  | Lymphocyte (%) | +/- 2SD | 3.0 | 0.43 | 1.67 |
|  | Monocytes (%) | +/- 2 SD | 4.0 | 0.6 | 1.0 |

| QC Level | Parameter | Max. Total Allowable Error | Max CV, % | Max SD | SDc = Coulter Range / 3 |
| --- | --- | --- | --- | --- | --- |
| Normal | WBC | +/- 15% | 3.0 | 0.27 | 0.27 |
|  | RBC | +/- 6% | 1.2 | 0.06 | 0.06 |
|  | Hemoglobin | +/- 7% | 1.4 | 0.23 | 0.20 |
|  | Hematocrit | +/- 6% | 1.4 | 0.67 | 0.90 |
|  | MCV | +/- 6% | 1.2 | 1.0 | 1.0 |
|  | Platelet Count | +/- 25% | 5.0 | 8.7 | 8.5 |
|  | Neutrophils (%) | +/- 2SD | 1.5 | 0.8 | 1.67 |
|  | Lymphocyte (%) | +/- 2SD | 3.0 | 0.79 | 1.67 |
|  | Monocytes (%) | +/- 2 SD | 5.0 | 0.5 | 1.0 |

| QC Level | Parameter | Max. Total Allowable Error | Max CV, % | Max SD | SDc = Coulter Range / 3 |
| --- | --- | --- | --- | --- | --- |
| Abnormal II | WBC | +/- 15% | 3.0 | 0.10 | 0.13 |
|  | RBC | +/- 6% | 1.2 | 0.022 | 0.027 |
|  | Hemoglobin | +/- 7% | 1.4 | 0.07 | 0.10 |
|  | Hematocrit | +/- 6% | 1.4 | 0.21 | 0.50 |
|  | MCV | +/- 6% | 1.2 | 1.0 | 1.0 |
|  | Platelet Count | +/- 25% | 5.0 | 3.0 | 5.0 |
|  | Neutrophils (%) | +/- 2SD | 2.4 | 1.0 | 1.67 |
|  | Lymphocyte (%) | +/- 2SD | 2.2 | 1.0 | 2.0 |
|  | Monocytes (%) | +/- 2 SD | 8.0 | 0.7 | 1.0 |

| QC Level | Parameter | Max. Total Allowable Error | Max CV % | Max SD | SDc = Coulter Range / 3 |
| --- | --- | --- | --- | --- | --- |
| Level I | Retic | +/- 2SD | 15.4 | 0.17 | 0.2 |
| Level II | Retic | +/- 2SD | 3.2 | 0.1 | 0.4 |
| Level III | Retic | +/- 2SD | 4.2 | 0.4 | 0.8 |

Maximum total allowable error is based on CLIA 88 criteria, which also are the CAP evaluation criteria.

Max CV is established by QC BPT to be consistent with recommended QC rules (see part c, below) in order to detect changes in the assay that would cause an error that exceeded the maximum allowable total error.

Max SD is determined by multiplying the maximum CV assay value. The assay value changes slightly for each new lot, however, it is expected that the precision will remain constant for each new lot of material...

RUN REJECT CRITERIA: The QC procedure for this assay will employ the 1-2SD Westgard rules. The 2SD limit will be identical to the Coulter QC limit. Runs where this QC rule is violated will be rejected.

Each time one control exceeds the criteria for rejection, the run is out of

control (failed), and patient results must not be reported. The steps on the QC

flow chart must be followed to resolve the problem. Refer to SOP PHQ.020 for

corrective action(s) if the values fall outside of the published limits

- Corrective Action

- Rejected runs must be effectively addressed by corrective action. Steps taken in response to QC failures must be documented.  Patient samples in failed analytical runs must be reanalyzed according to the Laboratory QC protocol. Supervisor may override rejection of partial or complete runs only with detailed documentation that follows criteria that is approved by the Medical Director.

- Corrective action documentation must include the following: QC rule(s) violated, the root cause of the problem, steps taken to correct the problem, how patient samples were handled, resolution, and the date and initials of the person recording the information.

- Review of QC

- Upon weekly and monthly review of QC by the Lead Technologist or designee, if the Quality Control data is showing a shift or trend,  investigate the cause for the imprecision and document corrective actions. Monthly QC files are printed, compiled in a log and reviewed by the department supervisor/manager or designee.

- All daily shift QC must be submitted with 5 days of outdating to Coulter’s IQAP program for interlaboratory comparison.

- 6.6	Review Patient Data

Review patient results for unusual patterns, trends or distributions, looking for an unusually high percentage of abnormal results.

- 6.7	Documentation

    - QC results for each control level and each test mode for an instrument are recorded and stored in the instrument.

    - QC records are printed monthly and maintained and available for a minimum of two (2) years.

    - Patient results are reviewed and released to the patient file via the LIS system.

- 6.8	Quality Assurance Program

    - Refer to the QA / QC policy for other quality assurance activities applicable to this procedure.

    - Training must be successfully completed and documented prior to performing this test.

    - The laboratory participates in CAP proficiency testing.

    - All techs that perform differential reviews will be required to perform automated and manual differential correlations

- 6.9	Other QA Tools – XB Moving Averages

| IF … | THEN … |
| --- | --- |
| XB shows a characteristic pattern of an upward or downward drift | - Check patient population to eliminate the possibility of an increased number of patients with a specific disease state. If this is found, then continue to run instrument. - Check patient population, if an increased number of patients with a specific disease state are not found; check commercial material for similar trends/shifts. |
| If commercial control material is in control | The instrument can continue to be operated. |
| If the commercial control material shows a similar trend/shift | Troubleshoot the instrument and calibrate if necessary. |

| Hints For XB Troubleshooting |  |  |  |
| --- | --- | --- | --- |
| When Measurement | Then … |  |  |
|  | MCV | MCH | MCHC |
| HGB Decreased | No change | Decreased | Decreased |
| HGB Increased | No change | Increased | Increased |
| RBC Decreased | Increased | Increased | No change |
| RBC Increased | Decreased | Decreased | No change |
| HCT Decreased | Decreased | No change | Increased |
| HCT Increased | Increased | No change | Decreased |

# EQUIPMENT and SUPPLIES

- 7.1	Assay Platform

- Beckman Coulter LH750

| Brand | Instrument Model | Distributor |
| --- | --- | --- |
| Coulter Electronics | LH750 | Beckman-Coulter, Inc. 200 S Kraemer Blvd, Brea, CA 92822 1-800-697-0099 Technical Support 1-800-526-7694 |

- 7.2	Equipment

| Item | Supplier and Catalog Number |
| --- | --- |
| Microscope | None specified |
| Slide Stainer | None specified |
| Refrigerator, 2-8C | None specified |

- 7.3	Supplies

| Item | Supplier and Catalog Number |
| --- | --- |
| Biohazard wipes | None specified |
| Immersion Oil | None specified |
| Printer | None specified |
| Applicator sticks | None specified |
| Glass Slides | None specified |
| Lens Paper | None specified |
| Optical lens cleaner | None specified |

# PROCEDURE

NOTE:  For all procedures involving specimens, buttoned lab coats, gloves and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor.

8.1	Preventive Maintenance

Daily and weekly maintenance will be performed and documented on the maintenance log by assigned personnel in the day shift.  Please refer to the LH750 Help Screen.

8.2	Start-up/Shut down Procedure

Refer to the LH750 Help Screen.

  - Loading Cassettes

Prior to loading cassettes, mix specimens on a mechanical rocker for 5-10 minutes

| Step | Action |  |
| --- | --- | --- |
| 1. | IF: Specimen received in standard tube containing optimum amount. | THEN: Load cassettes making sure all bar code labels are positioned appropriately. |
| 2. | IF: Specimen received in Microtainer tube or contains minimum amount of blood. | THEN: Run the specimen in the open mode. |
| 3. | Place the cassettes on the loading bay. |  |
| 4. | The instrument will start automatically all counts ordered will be performed. |  |
| 5.. | When load is completed, remove cassettes from instrument. |  |

  - Review of Patient Result

| Step | Action |
| --- | --- |
| 1. | Using the “RULES” 2 function in the LIS, review each patient result before it is released. |
| 2. | Check for delta check flags and critical values. |
| 3 | Call and document all critical values |
| 4. | Release all values that do not need to be repeated for delta values, critical values, or are not flagged on the LH print out for review. |
| 5 | Pull all specimens that need rerun or slide scan as indicated in Addendum 2. |
| 6. | Store normal specimens. |
| 7. | Rerun tests (those needing repeat analysis as indicated in Addendum 2). |
| 8. | For those specimens that are flagged for scan smear or manual diff, release the hemogram after taking out the manual diff. Refer to Addendum 4. |
| 9. | Make slide for scan smear. |

  - Pathologist/ (or designee)/ slide review

| Abnormality | Path. |
| --- | --- |
| Prolymphs > 5% | X |
| Reactive and/or atypical lymphocytes >20% | X |
| Bands > 25% | X |
| Meta/Myelos/Promyelo >10% | X |
| Any blast cell | X |
| Any unidentifiable cell | X |
| Any parasite or micro organism | X |
| Lymphocyte > 70% | X |
| Review by physician request (PRSM) | X |
| WBC < 1.5 > 40.0 with abnormal cells | X |
| NRBC with varying stages of maturity and /or morphological variances | X |
| RBC morphology: 3+ spherocytes, RBC inclusions, suspect HGB-C, crystals, Pappenheimer bodies and bizarre morphology. | X |

NOTE: The above guidelines are for new and recurring patients performed initially and

Over each subsequent hospital encounter (ED visit, OP visit or admission).

# CALCULATIONS

MCV, MCH, MCHC and absolute differential results are released from the LH analyzer.

The absolute differential results are released from either the analyzer or the LIS, depending upon the differential type:

  - Automated differentials have the absolute values calculated by the LH750.

  - Manual differentials have the absolute values calculated by the LIS.

There are instances when results are above assay range or interfering substances require manual correction of assay parameters. These calculations are verified at least annually as well as whenever a change is made to the LIS that could impact a calculation. See Addendum 5 for calculation formulas.

# REPORTING RESULTS AND REPEAT CRITERIA

- 10.1	Interpretation of Data

None required

- 10.2	Rounding

Any result rounded is performed at the interface level.

10.3	Units of Measure

| Parameter | Units |
| --- | --- |
| WBC | 103/µL |
| RBC | 106 /µL |
| HGB | g/dL |
| HCT | % |
| MCV | FL |
| MCH | Pg |
| MCHC | g/dL |
| PLT | 103/µL |
| MPV | FL |
| RDW | % |
| Differential Absolute Values | 103/µL |
| Differential Counts | % |
| Reticulocyte | % |

- 10.4	Clinically Reportable Range (CRR)

| Parameter | Clinical Reportable Range |
| --- | --- |
|  | LH750 |
| WBC | 0-800 x 103 |
| RBC | 0-16.00 x 106 |
| HGB | 0-25 |
| HCT | Calculated and limited by direct measurement reportable ranges |
| MCV | 0-150 |
| MCH | Calculated and limited by direct measurement reportable ranges |
| MCHC | Calculated and limited by direct measurement reportable ranges |
| PLT | 0-3,000 x 103 |
| % NEUTS | 0-100 |
| % LYMPHS | 0-100 |
| % MONO | 0-100 |
| % EOS | 0-100 |
| % BASO | 0-100 |
| Reticulocyte Percentage | 0.0-30.0 % |

- 10.5	Repeat Criteria and Resulting (Refer to Addendum 2)

| Parameter | Repeat Tolerance Limits |
| --- | --- |
| WBC | ± 0.8 |
| RBC | ± 0.25 |
| HGB | ± 0.6 |
| HCT | ± 1.7 |
| MCV | ± 3.0 |
| MCH | ± 1.2 |
| MCHC | ± 1.2 |
| PLT | ± 10% |
| NE% | ± 5.0 |
| LY% | ± 5.0 |
| MO% | ± 3.0 |
| EO% | ± 2.0 |
| BA% | ± 1.0 |

# EXPECTED VALUES

- 11.1	Reference Ranges

Refer to Addendum 1.

11.2	Priority 1, 2 & 3 Results

Refer to the Critical Value Policy and Procedure.

# CLINICAL SIGNIFICANCE

- CBC – The quantitative and qualitative analysis of the cellular elements of blood will identify imbalance between cell production, cell release, cell survival, or cell loss.  This information increases the accuracy and specificity of diagnosis based on pathogenesis and is also used to monitor the effectiveness of therapy.

- Automated Differential – The Differential distribution of white blood cells will, when correlated with the absolute white cell count, identify imbalances between cell production, cell release, cell survival and/or cell loss.  This information increases the accuracy and specificity of diagnosis based on pathogenesis and is also used to monitor the effectiveness of therapy.

- Platelet Count – Platelets must be present in adequate numbers and have proper function to aid in hemostasis.  A normal bleeding time is dependent on adequate platelet number and function.

- Reticulocyte Count - The enumeration of reticulocytes provides an effective means of Determining red cell production and regeneration.   Elevation is seen in patients with Hemolytic anemia, hemorrhage (acute and chronic), treatment of iron-deficiency anemia and megaloblastic anemia and uremia.  Decreased counts may be seen in aplastic anemia, Aplastic crisis of hemolytic anemias and ineffective erythropoiesis as seen in thalassemia, pernicious anemia and sideroblastic anemia.

# PROCEDURE NOTES

- FDA Status: Approved

Validated Test Modifications: None

- 13.1	Manual versus Automated Differential Counts 95% Comparison Confidence Limits

- This table can be used for two purposes:

- To show the tolerance limits of a manual differential at various levels of counting (100-cell diff, 200-cell diff, etc.)

- To determine the tolerance allowed for a technologist performing a 100-cell diff verifying an automated differential with a 95% confidence limit. (If the instrument reports 20% monocytes, the technologist would be expected to find 12-30% monocytes in the100-cell differential count in order to verify the instrument count.)

- “A” is the percentage of cell type counted, e.g. lymphocytes.

- “N” is the size of the manual differential performed.

| A = % of a cell type | N = 100 | NN = 200 | N = 500 |
| --- | --- | --- | --- |
| 0 | 0 - 4 | 0 - 2 | 0 - 1 |
| 1 | 0 - 8 | 0 - 4 | 0 - 3 |
| 2 | 0 - 8 | 0 - 6 | 0 - 4 |
| 3 | 0 - 9 | 1 - 7 | 1 - 5 |
| 4 | 1 - 10 | 1 - 8 | 2 - 7 |
| 5 | 1 - 12 | 2 - 10 | 3 - 8 |
| 6 | 2 - 13 | 3 - 11 | 4 - 9 |
| 7 | 2 - 14 | 3 - 12 | 4 - 10 |
| 8 | 3 - 16 | 4 - 13 | 5 - 11 |
| 9 | 4 - 17 | 5 - 14 | 6 - 12 |
| 10 | 4 - 18 | 6 - 16 | 7 - 13 |
| 15 | 8 - 24 | 10 - 21 | 11 - 19 |
| 20 | 12 - 30 | 14 - 27 | 16 - 24 |
| 25 | 16 - 35 | 19 - 32 | 21 - 30 |
| 30 | 21 - 40 | 23 - 37 | 26 - 35 |
| 35 | 25 - 46 | 28 - 43 | 30 - 40 |
| 40 | 30 - 51 | 33 - 48 | 35 - 45 |
| 45 | 35 - 56 | 37 - 53 | 40 - 50 |
| 50 | 39 - 61 | 42 - 58 | 45 - 55 |
| 55 | 44 - 65 | 47 - 63 | 50 - 60 |
| 60 | 49 - 70 | 52- 67 | 55 - 65 |
| 65 | 54 - 75 | 57 - 72 | 60 - 70 |
| 70 | 60 - 79 | 63 - 77 | 65 - 74 |
| 75 | 65 - 84 | 68 - 81 | 70 - 79 |
| 80 | 70 - 88 | 73 - 86 | 76 - 84 |
| 85 | 76 - 92 | 79 - 90 | 81 - 89 |
| 90 | 82 - 96 | 84 - 94 | 87 - 93 |
| 91 | 83 - 96 | 86 - 95 | 88 - 94 |
| 92 | 84 - 97 | 87 - 96 | 89 - 95 |
| 93 | 86 - 98 | 88 - 97 | 90 - 96 |
| 94 | 87 - 98 | 89 - 99 | 91 - 96 |
| 95 | 88 - 99 | 90 - 98 | 92 - 97 |
| 96 | 90 - 99 | 92 - 99 | 93 - 98 |
| 97 | 91 - 100 | 93 - 99 | 95 - 99 |
| 98 | 92 - 100 | 94 - 100 | 96 - 100 |
| 99 | 94 - 100 | 96 - 100 | 97 - 100 |
| 100 | 96 - 100 | 98 - 100 | 99 - 100 |

- 13.2	WBC Estimate

| IF | Then |
| --- | --- |
| Using the 50X objective | Calculate the average WBC in 10 fields. Multiply by 3,000. |
| - cellular interference is flagged - WBC estimate does not equal the Coulter WBC within 20% - Platelet clumps are present on diff slide review - NRBCs and/or megakaryocytes or giant platelets present - No apparent cause | - Investigate the cause by performing a WBC estimate. - add to the report - perform the manual diff and correct the WBC - Have the test redrawn. |

13.3	Platelet Estimate

| IF | Then |
| --- | --- |
| In the presence of a platelet flag, | Perform a platelet estimate using the 100X objective Count the PLT in each of 10 microscopic fields in areas of the slide where the RBCs are evenly dispersed. Divide the total # of platelets by 10 to establish the mean and multiply by 20,000. |
| The Coulter platelet count and the platelet estimate do not agree within 20% | Repeat the platelet estimate and/or platelet count. If counts still do not agree, consult the supervisor or designee. |

- 13.4	RBC Morphology

- Microcytosis, Macrocytosis and Anisocytosis will be quantitated using the LH criteria.

- The morphology will be quantitated by smear evaluation.

- All clinically significant findings such as specific cell types, inclusions, polychromasia, etc., will be reported from the smear evaluation.

| IF | Then |
| --- | --- |
| There are NO clinically significant findings to be added to a patient report. | Result as Normal. |
| There are ANY additions to be made to the patient report, such as RBC morphology, cell differential, PLT morphology, etc. | Report all clinically significant findings. |

- For consistent morphological reporting, the following criteria are recommended.  They serve only as a guideline for evaluating slight, moderate, or marked degrees of abnormal morphology.

| Variation | Mean Range per 10 Fields (100x) of RBCs | Then |
| --- | --- | --- |
| Poikilocytosis | 0 | Normal |
|  | 1-5 | 1+ |
|  | 6-15 | 2+ |
|  | Over 15 | 3+ |
| Anisocytosis | 0-5 | Normal |
|  | 6-15 | 1+ |
|  | 15-30 | 2+ |
|  | Over 30 | 3+ |
| Polychromasia | 0-2 | Normal |
|  | 3-4 | 1+ |
|  | 5 - 6 | 2+ |
|  | Over 6 | 3+ |
| Hypochromia | 0-5 | Normal |
|  | 6-15 | 1+ |
|  | 16-30 | 2+ |
|  | Over 30 | 3+ |

Listed below is a guideline for abnormal shapes based on mean range/10 fields of RBCs.

| Abnormal Shape | Normal | 1+ | 2+ | 3+ |
| --- | --- | --- | --- | --- |
| Spherocyte, Acanthocyte Sickle cell, Rouleaux | 0 | 1-5 | 6-15 | Over 15 |
| Helmet cell | 0-1 | 1-5 | 6-15 | Over 15 |
| Tear drop, Target cell, Schistocyte, Ovalocyte, Elliptocyte, Burr cell, Stomatocyte, Blister cell | 0-1 | 2-5 | 6-15 | Over 15 |

13.5	Potential Causes of Erroneous Results with Automated Cell Counter

| Parameter | Causes of Spurious Increase | Causes of Spurious Decrease |
| --- | --- | --- |
| WBC | Cryoglobulin, Cryofibrinogen, Heparin, Monoclonal Proteins, Nucleated RBC, PLT Clumps, Lyse-resistant RBC | Clotting, Smudge Cells, Uremia, Immunosuppressants |
| RBC | Cryoglobulin, Cryofibrinogen, Giant PLTs, High WBC (>50,000/L) | Auto-agglutination, Clotting, in vitro Hemolysis, Microcytic RBC |
| Hemoglobin | Carboxyhemoglobin (>10%), Cryoglobulin, Cryofibrinogen, in vitro Hemolysis, Heparin, High WBC (>50,000/L), Hyperbilirubinemia, Lipemia, Monoclonal Proteins | Clotting, Sulfhemoglobin |
| Hematocrit (Automated) | Cryoglobulin, Cryofibrinogen, Giant PLTs, High WBC (>50,000/L), Hyperglycemia (Glucose >600 mg/dL) | Autoagglutination, Clotting, in vitro Hemolysis, Microcytic RBC |
| MCV | Cryofibrinogen, Autoagglutination, High WBC (>50,000/L), Hyperglycemia, Reduced RBC Deformability | Cryoglobulin, Giant Platelets, in vitro Hemolysis, Microcytic RBC, Swollen RBC |
| MCH | High WBC (>50,000/L), Spuriously High HGB, Spuriously Low RBC | Spuriously Low HGB, Spuriously High RBC |
| MCHC | Auto-agglutination, Clotting, Lipemia, in vitro Hemolysis, Spuriously High HGB, Spuriously Low HCT | High WBC (>50,000/L), Spuriously Low HGB, Spuriously High HCT |
| Platelets | Cryoglobulin, Cryofibrinogen, Hemolysis (in vitro and in vivo), Microcytic RBC, RBC Inclusions, WBC Fragments | Clotting, Giant PLT, Heparin, PLT Clumping, PLT Satellitosis |

  - Platelet Clumps

| If | Then |
| --- | --- |
| If PLT count 130 with significant PLT clumps found during slide scan. | Remove the PLT count number and result with the comment !PLC = Clumped platelet Request for a redraw on a blue top. |
| If the redraw is still clumped | Remove the platelet count number and comment, “unable to quantitate due to significant platelet clumping, slide estimate appears increased or decreased depending on the technologist review”. |

- 13.7	Sodium Citrate for Platelet Count

Usually reserved for patients who are known to have a platelet clumping phenomena associated with EDTA anticoagulant.  The specimen of choice is both an EDTA and a sodium citrate tube.  The EDTA is used for the CBC results.  The sodium citrate tube is used for the citrate Platelet count.  Run samples as per the LH750 protocol.  Multiply the Na citrate platelet count by 1.1 to correct for dilution effects.

13.8	MCHCs greater than 36.5 or less than 29.0

If the MCHC is 29.0 or 36.5, it should be repeated one more time on the LH750 to rule out random error.  If the MCHC is still greater than 36.5, a slide should be made and examined to determine the integrity of the specimen.  The smear review should indicate to the technologist which category the specimen falls into cold agglutinin, lipemia, hemolysis, or the situation where the results are accurate due to the presence of spherocytes.

| IF | Then |  |
| --- | --- | --- |
| Spherocytes are noted on the slide scan. | Report the MCHC with a comment reflecting the presence of spherocytes. |  |
| Resistant hemoglobin, marked sickle cells or target cells noted on the slide scan | Specimens with lyse resistant RBCs should be repeated on dilution using bottled, distilled water. Prepare a 1:2 dilution with equal parts of blood and water. Allow to sit three minutes. Resuspend and process through the analyzer. Using the HGB result, multiply the results by 2 to determine the corrected hemoglobin result. Use the corrected HGB to recalculate the MCH and the MCHC. |  |
| If significant RBC clumping is noted on the slide scan. | Warm specimen in a 37°C water bath or heat block for 30 minutes and rerun. If necessary, make a warmed slide for morphology evaluation |  |
|  | IF After Incubation | Then |
|  | The MCHC is within normal range | Report results with the appropriate comment. |
|  | The MCHC is still outside 36.5 (strong cold agglutinins) | Perform the following Plasma Replacement Procedure |
| If hemolysis is suspected on the slide scan, ie schistocytes | Examine the specimen for visual hemolysis. If gross hemolysis is observed, cancel the specimen with the appropriate comment |  |
| If lipemia is suspected on the slide scan. | Examine the specimen for visual lipemia. If lipemia is observed perform a plasma hemoglobin blank. If there is sufficient specimen, mix well and pour off a portion into a plastic specimen tube. Spin the tube for 5-10 minutes at 2000 rpm. If the specimen is short, spin the lavender tube for 5-10 minutes at 2000 rpm. In secondary mode run an Isoton blank. Verify a “0” hemoglobin value. In the secondary mode, aspirate plasma portion of spun specimen to determine the plasma hemoglobin blank value. Using the following formula: Correct Hgb = OH – [PB x (1 – HCT/100) ] Where OH = original hemoglobin PB = plasma hemoglobin blank HCT = original hematocrit Calculate corrected HGB. Enter the corrected HGB on the report and recalculate the indices and enter the correct results... With comment RESULTS WERE OBTAINED BY REPEAT ANALYSIS TO INCLUDE RUNNING A PLASMA BLANK TO ELIMINATE INTERFERENCES CAUSED BY EITHER WBCS, LIPEMIA, OR PROTEIN ENTITIES. |  |

- 13.9	Correction for Nucleated RBCs and/or Megakaryocytes and/or when a Cellular        Interference flag is received.

- Whenever the LH instrument gives a cellular interference flag a slide WBC estimate has to be done. If the estimate does not match within 20% of the LH WBC count a WBC correction has to be done. See section 13.2. Use the following calculation if this correction has to be done manually:

Corrected WBC

Uncorrected WBC x 100

100 + # NRBCS and/or Megakaryocytes

=

- Whenever the LH instrument numerates NRBC’S a slide MUST be reviewed for the presents of NRBC. If no NRBC is seen on the smear the Coulter LH NRBC count should be removed.

13.10	Slide Preparation

When making a smear always check the specimen for clots.  This can be done by visual inspection or by the use of an applicator stick when appropriate.

- 13.11	Coulter Repeats

(See Addendum 2) Results must be reported with the comment.

- REP = RESULTS CONFIRMED, TEST REPEATED.

13.12	Smear Scan

If immature/abnormal cells are noted or in the presence of a “Diff verification” flag, manual differential must be performed and reported.

- 13.13	Correction of RBC & HGB in the presence of a WBC counts greater than 400,000.

a)	Subtract the WBC count from the RBC count to obtain a corrected RBC Count.

b)	Spin an aliquot of specimen for 3-5 minutes at 1500 rpm.

c)	Remove an aliquot of the red cell portion, and dilute it 1:2 with Isoton.

d)	Run the “RBC only” suspension as a sample on the Coulter to obtain an accurate MCH and MCV.

e)	Calculate the corrected HGB:  Hgb = MCH x RBC (corrected) / 10

f)	Calculate the HCT:  Hct = MCV x RBC (corrected) / 10

g)	Calculate the MCHC:  MCHC =	Corrected Hgb x 100

Corrected Hct

- 13.14   Special Reticulocyte Precautions

Specimens with verify retic flag must be Repeated with the Second Instrument, If 			the Flag Persist, Send the Specimen to Chantilly.

- 13.15	Alternative Procedures

None

# LIMITATIONS OF METHOD

14.1	CBC-Line Linearity

- CBC-Line Linearity is an assayed material used to establish the Analytical Measurement Range and verify the calibration of the LH750.  This product allows Quest Diagnostics to comply with CLIA 88 Regulations and CAP requirements.

- Analytical Measurement ranges should be performed at installation and when necessary to verify linearity.

| Recommended CVA, Linearity Kit | Supplier & Catalog Number |
| --- | --- |
| STRECK CBC-Line Linearity Kit STRECK Retic-Chex Linearity Kit (Follow manufacturers requirements for storage and stability) | STRECK Systems –Cat.# 291712 (LH750) STRECK Systems –Cat.# 285535 (LH750) |

Follow Manufacturers requirements for storage and stability.

- 14.2	Analytical Measurement Range (AMR)

| Parameter | Analytical Measurement Range |
| --- | --- |
|  | LH750 |
| WBC | 0-400 x 103 |
| RBC | 0-8.00 x 106 |
| HGB | 0-25 |
| MCV | 0-150 |
| Reticulocyte | 0.0 – 30.0 % |
| PLT | 0-3,000 x 103 |

- 14.3	Precision

| Inter-Run |  |  |  |  |
| --- | --- | --- | --- | --- |
| Analyte | Level | Mean | 1SD | CV% |
| WBC | LOW | 1.15 | 0.08 | 6.90 |
|  | MID | 20.01 | 0.18 | 0.90 |
|  | HIGH | 94.57 | 1.28 | 1.36 |
| PLATELET | LOW | 4.82 | 0.48 | 10.06 |
|  | HIGH | 647.29 | 20.69 | 3.20 |
| HEMOGLOBIN | LOW | 5.16 | 0.08 | 1.53 |
|  | HIGH | 16.76 | 0.13 | 0.76 |

| Intra-Run |  |  |  |
| --- | --- | --- | --- |
| Analyte | Mean | 1SD | CV% |
| WBC | 7.40 | 0.10 | 1.33 |
| RBC | 4.87 | 0.03 | 0.62 |
| HEMOGLOBIN | 15.21 | 0.09 | 0.62 |
| PLATELET | 235.86 | 8.15 | 3.46 |
| % NEUTS | 63.85 | 0.42 | 0.33 |
| % LYMPHS | 27.30 | 0.43 | 1.57 |

- 14.4	Interfering Substances

See 13.5

- 14.5	Clinical Sensitivity/Specificity/Predictive Values

Not applicable.

# SAFETY

You, the employee, have a direct responsibility to avoid injury and illness at work.  Nearly all-harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental, Health and Safety (EHS) Manual to learn the requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

Report all accidents and injuries to your Supervisor or the Environmental Health         and Safety Coordinator.

# RELATED DOCUMENTS

- Material Safety Data Sheets

- LH 750 Reference Manual

- S-Cal package insert

- Critical Alert Policy

- QC / QA Policy

- Procedure Validation Package

# REFERENCES

1)	Coulter Counter Model GENS Reference Manual (PN 4237221), November, 1996.  Coulter Counter Model STKS Reference Manual (PN42371828), March 1995.

2)	Coulter Counter Model LH750 Operator’s Guide (PN 4277249B), May 2002. Coulter Counter Model GENS Operator’s Guide (PN 4237222), November 1996. Coulter Counter Model STKS Operator’s Guide (PN437188B), March, 1995.

3)	Coulter Counter Model GENS, Special Procedures and Troubleshooting (PN 4237223), November 1996.  Coulter Counter Model STKS, Special Procedures and Troubleshooting (PN4237187B), March 1995.

4)	“Hematology Procedures for Abnormal Bloods” (PN 4206695A), April, 1999, Coulter Education Center, Miami Lakes, Fl

5)	Bull, B.S. and Korpman, R. “Interlaboratory quality control using patient’s data.” In Carell, I., Ed., Quality Control, Churchill Livingston, 1982.

6)	“Differential Leukocyte Counting”, CAP Conference, Aspen 1977, John A. Koepke, M.D., Ed., Published by College of American Pathologists.

7)	Color Atlas of Hematology CAP Hematology and Clinical Microscopy Resource Committee Distinction between megakaryocyte and giant platelets in the section “Megakaryocyte Maturation”.

8)	Hematology:  Principles and Procedures by Barbara A. Brown, 2nd Edition, 1976, pp 77-81.

9)	In-House (Atlanta Business Unit), Truth table data performed with STKS 2B software, March 1995.

10)	Gulah, Asselta, Chen.  Using a Vortex to Disaggreate Platelet Clumps.  Laboratory Medicine, Oct 1997 Volume 28.  665-557.

11)	A Color Atlas & Instruction Manual of Peripheral Blood Cell Morphology, O'Connor Bartera, 1984.

12)	Nathan, David G. & Oski, Frank A. “Hematology of Infancy and Childhood”, 3rd edition, 1987.

13)	Lanzkowsky, Philip, “Pediatric Hematology-Oncology, a Treatise for the Clinician”, 1980.

14)	Miller, Dennis R. Ed. “Blood Diseases of Infancy and Childhood”, 5th edition, 1984.

15)	Henry, J B, ed. “Clinical Diagnosis and Management by Laboratory Methods”, 20th edition, 2001.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
| 000 | 7/16/08 | Page 1 | Changed local to laboratory approval. Updated owner. | O Dualeh | R. Levy |
| 000 | 7/16/08 | 13 | Added FDA Status and Validated test Modifications | O Dualeh | R. Levy |
| 000 | 7/16/08 | 11.2 | Update section title to new terminology | O Dualeh | R. Levy |
| 001 | 8/31/09 | 6.5 | Added Patient QC for nRBCs | O Dualeh | R. Levy |
|  | 8/31/09 | 19.3 | Updated Addendum 3 | O Dualeh | R. Levy |
|  | 9/3/09 | 1 | Corrected test code for Retic count: Retc | O Dualeh | R. Levy |
| 002 | 10/13/09 | 21 | Updated Linearity kit and Removed Manual Retic. | O Dualeh |  |
| 003 | 11/16/09 | 8.5 | Updated Abnormal slide review policy | O Dualeh |  |
| 004 | 8/29/10 | Page 1 | Update Owner | A. Foster | R. levy |
|  |  | 5.3 | Added new requirements to calibration frequency |  |  |
|  |  | 6.1 | - Deleted patient control QC - Added Mode-to-mode QC check |  |  |
|  |  | 6.3 | - Deleted patient control from multimode sampling, and added commercial control and performance check - Deleted Patient control and added new reagent or diluent |  |  |
| 004 | 8/29/10 | 6.4 | Added Lead Tech as QC reviewer | A. Foster | R. levy |
|  |  | 6.5 | - Deleted patient control use - Added mode-to-mode Quality Control |  |  |
|  |  | 6.8 | Added Tech requirements for differential review and correlation |  |  |
|  |  | 8.3 | Formatted cassette loading chart to include a 5th step |  |  |
|  |  | 8.5 | - Added abnormal retic patterns or retic counts 20% to the pathology slide review - Deleted Lymphocyte count >70% in adults |  |  |
|  |  | 10.4 | Added Clinically Reportable Range for Reticulocyte percentage |  |  |
|  |  | 13.2 | Complete Format change of the WBC estimate table |  |  |
|  |  | 13.3 | Minor format and placement changes to the platelet estimate table |  |  |
|  |  | 13.8 | Added less than 29.0 to MCHC parameter |  |  |
|  |  | 16.0 | Added Procedure validation package to related documents |  |  |
| .005 | 3/1/11 | Page 1 | Removed Effective date from Laboratory Approval | A. Foster | R. Levy |
|  |  | 5.3 | Added re-calibration procedure for CBC parameters |  |  |
| .006 | 4/9/11 | 6.2 | Added Expiration date of Retic-C and 5C Controls | A.Foster | R. Levy |
|  |  | 5.3 | Removed re-calibration and added calibration verification procedure |  |  |
| .007 | 8/2/11 | 6.4 | Reference: PHQ.020 for QC review | A. Foster | R.Levy |
| .008 | 11/22/11 | Page 1 | Updated owner | W.McMillan | R. Levy |
|  |  | 8.5 | Removed criteria for Pathologist review for >20% reticulocyte count |  |  |
|  |  | 10.4 | Reformatted table. No change to ranges |  |  |
|  |  | Addendum 2 | Updated repeat criteria for Reticulocytes |  |  |
|  |  | Addendum 3 | Updated Required action for abnormal retic patterns. Removed requirement to send sample to Chantilly |  |  |
| .009 | 03/07/12 | Page 1 | Removed ‘annual’ from ‘review’ table. | W.McMillan | R. Levy |
|  |  | 3.2 | Changed formatting of table in section 3.2 |  |  |
|  |  | 5.3 | Added section 5.3 on calibration preparation |  |  |
|  |  | 5.4 | Added section 5.4 on calibration documentation |  |  |
|  |  | 6.1 | Reformatted section 6.1. Boxed and bolded ‘caution’ statement. |  |  |
|  |  | Addendum 4 | Added addendum 4 “Daily Quality Control for LH750” |  |  |
|  |  | Addendum 5 | Added addendum 5 “Calculation Formulas” |  |  |

# ADDENDA

| Addendum | Title |
| --- | --- |
| 1 | Reference Ranges |
| 2 | Repeat Criteria |
| 3 | LH 750 Decision Rules and Criteria |
| 4 | Daily Quality Control for LH750 |
| 5 | Calculation Formulas |

- ADDENDUM 1

- CBC AND DIFFERENTIAL REFERENCE RANGES

- Male reference ranges	                     Female reference ranges

| Parameter/Units of measurement |  |  |  |
| --- | --- | --- | --- |
| WBC/ 103/µL | 4.0-10.2 |  | 4.0-10.2 |
| RBC/ 106/µL | 4.5-6.0 |  | 4.0-5.0 |
| HGB/ g/dL | 13.0-16.8 |  | 11.5-16.0 |
| HCT/ % | 40-54 |  | 37-47 |
| MCV/ fL | 77-98 |  | 77-98 |
| MCH/ pg | 24.6-32.5 |  | 24.6-32.5 |
| MCHC/ g/dL | 32.2-35.5 |  | 32.2-35.5 |
| RDW/ % | 11.6-14.7 |  | 11.6-14.7 |
| PLT/ 103/µL | 140-400 |  | 140-400 |
| Neutrophils/ % | 50-70 |  | 50-70 |
| Lymphs % | 20-40 |  | 20-40 |
| Monocytes % | 4-13 |  | 4-13 |
| Eosinophils % | 0-6 |  | 0-6 |
| Basophils | 0-2 |  | 0-2 |
| Bands | 0-6 |  | 0-6 |
| Nucleated RBC % | 0 |  | 0 |
| Retic – Automated % | 0.6-2.7 |  | 0.6-2.7 |

- ADDENDUM 2

REPEAT CRITERIA

|  | KEY |
| --- | --- |
| RPT – | Repeat CBC on LH750 |
| SCAN – | Microscopically scan smear & manual differential if required |

| Parameter | Condition | Immediate Action | Follow-Up Required Action |
| --- | --- | --- | --- |
| WBC | 1.0 | RPT, SCAN | - Re-analyze, verify count within + 15% Add the comment that the result was checked. - Check sample for clots. If clotted, cancel the test and notify the ordering doctor or unit. - Refer to WBC <1.5 - If unable to evaluate 100 cells, do a 50-cell diff and multiply the results by 2. |
|  | 1.5 | RPT, SCAN | - Re-analyze, scan to verify count verify count within + 15%. Add the comment that the result was checked. Rule out erroneous decrease due to: - Presence of fibrin. Check microscopically and macroscopically by using applicator stick to rim specimen. If fibrin is present, add the comment and request for a redraw. - Excessive number of small WBCs below the 35 fL threshold |
|  | 30.0 | SCAN | Scan to verify count. Rule out erroneous increase due to: - 2-3+ presence of large/giant platelets. Add appropriate message code. - Presence of abnormal protein/cryoglobulin (blue streaks in smear). - Presence of NRBC. Correct WBC. Add appropriate message code. - Presence of unlysed particles above WBC threshold of 35 fL (crystals, lyse-resistant RBC). |
|  |  400.0 | RPT by Dilution, SCAN | - Re-analyze by dilution. Refer to AMR limits. Add comment RESULTS VERIFIED BY REPEAT ANALYSIS. - Refer to WBC >30.0 |
|  | WBC R or * flag with an NRBC flag or Cellular Interference | SCAN | - Remove the WBC count before releasing the hemogram - Scan to verify diff and WBC estimate. - Rule out erroneous results due to the presence of NRBC, PLT clumps or giant PLTs. |
| RBC |  8.00 | RPT by Dilution SCAN | - Re-analyze by dilution. Refer to AMR limits. Add comment REP = RESULTS CONFIRMED, TEST REPEATED |
|  | RBC Morphology Flag | SCAN | - Scan to verify morphology. Report morphology. |
| HGB | 6.0 | RPT, SCAN | - Re-analyze, verify count. Add comment. - Check for good H&H match. - Check sample for clots. |
|  | 20.0 25.0 (AMR) | RPT SCAN SCAN | - Re-analyze if greater than 20.0. Add comment. - If greater than 25.0 repeat by dilution. - Rule out hemoconcentration. (Pour off) - Check age of patient. |
| HGB (con’t) |  |  | Refer to supervisor or designee if this is not a pediatric newborn. Another quick check is to view the clot tubes on the patient for visibly high HCT level. |
| MCV | 50.0 | RPT SCAN | - Verify by repeat analysis. Add comment. - Verify value consistent with morphology review. See Action needed on ≤70.0. |
|  | 70.0 | SCAN | - Verify value consistent with morphology review. Denote any Target Cells, Sickle Cells, Schistocytes or Spherocytes. For 2+ or greater RBCs below threshold, evaluate accuracy of RBC count, consult supervisor if necessary. |
|  | 130.0 | SCAN | - Verify that value is consistent with morphology review. - Denote any rouleaux or RBC agglutinins, apply message codes, and consider holding quantitative values. If necessary, consult supervisor. - Pathologic conditions include macrocytic anemias such as pernicious anemia (oval macrocytes with hypersegmented neutrophils) and other megaloblastic anemia. - Check for presence of cold agglutinins or cryoglobulins. Usually see elevation of MCHC also. Warm specimen to 37C, 30 minutes and retest. - Apply message codes. |
|  | 150.0 | RPT SCAN | - Verify by repeat analysis. - Verify that the value is consistent with morphology review. See action needed on >130. |
| MCHC | 36.5 29.0 | RPT (warmed) SCAN | Refer to Section 13.8. |
| RDW | > 19.0 | Scan | - Verify value consistent with morphology review. - Review smear for RBC abnormalities. |
| Platelet | 30 | RPT, SCAN, check for clot | - Verify by repeat analysis. Add comment - Be suspicious if occasional fields on morphology review have 2-3 platelets/hpf. Check closely for fibrin, >2+ large/giant platelets, platelet satellitism or platelet clumps. Check tube for clot. |
|  | 100 | SCAN Check for clot | If clotted, cancel test and request for a redraw. |
| Platelet (cont) | 130 with Platelet Clumped suspect flag | SCAN, check for clot | - Remove the PLT count before the hemagram is released - Perform scan to rule out interferences caused by 2+ large or giant platelets, platelet clumps, platelet satellitism, fibrin, NRBCs, RBC fragments, or old blood/excessive degeneration, WBC fragments or clumps. - Re-result as in section 13.6 if significant platelet clumping is noted. |
|  |  3000 | RPT by dilution, SCAN | - Re-analyze by dilution. Correlate with morphology review. Add comment REP = RESULTS CONFIRMED, TEST REPEATED |
|  | Platelet Flags | SCAN | Scan diff to verify flag. Report noted observation. |
| Lymphocytes % | 1% | SCAN | Perform scan to verify % lymph count and to check immature or blast cells. |
|  | 70% >50% in presence of 1 of the following flags: blasts, variant lymph, or NRBC | SCAN | Perform scan to rule out immature, variant, atypical or blast cells. |
| Neutrophil | 1% | SCAN | Perform scan to verify % segmented neutrophil count and to check for presence of immature or blast cells. |
|  | 80% in the presence of 2 or more of the following flags: variant lymph, Ig2, blasts, or NRBC | SCAN | - Perform scan to rule out left shift, signs of inflammation or leukemic processes. - Perform manual diff if one or more of the following are noted: Band count appears >8%, any Meta, myelo, pro or blasts. If blasts present and appear myeloid check for Auer rods. Note any megaloblastic processes such as abnormal nuclear/cytoplasmic maturation, hypersegmentation of neutrophils, hypogranular neutrophils pyknotic nuclei. |
| % Monocytes | 1.0 15% | SCAN | Scan to verify monocytes. Rule out immature monocytes or precursors. |
| % Eosinophils | 15% | SCAN | Scan to verify eosinophils. Rule out presence of parasites. |
| % Basophils | 3.5% | SCAN | - Scan to verify presence. If suspect artificial increase, retrieve blood and check by repeat analysis, if still elevated, perform manual count. - If smear review reveals increase in basophils, check for leukemic picture. |
| Reticulocytes | >30% | RPT | - Repeat on different analyzer. - Scan smear to verify presence. - Report as “>30%”. - Note: Do NOT send sample to Chantilly for verification. |

- ADDENDUM 3

LH 750 DECISION RULES AND CRITERIA

| Condition (IF) | Action Required (THEN) |
| --- | --- |
| WBC aH | Make a slide and Scan Smear |
| No Match or No read | Check cassette for sample ID and test request |
| Basophilia % | Make a slide and Scan Smear |
| NE Blast or LY Blast or MO Blast | Perform Manual Differential |
| WBC ----- or RBC ----- or HGB ---- or PLT -- | Check for clots with a stick and rerun sample |
| WBC cL or WBC cH or PLT cL | Repeat Sample ,Call and document results |
| HCT cL or HCT cH or HGB cL or HGB cH | Repeat Sample ,Call and document results |
| Neutrophilia % or Lymphocytosis % or Monocytosis % or Eosinophilia % | Make a slide and Scan Smear |
| Imm.NE 2 or Variant LY or Low event # or Verify Diff | Perform Manual Differential |
| Imm. NE 1 and Variant LY | Perform Manual Differential |
| NE % --- or LY % --- or MO % --- or EOS % --- | Perform Manual Differential |
| MCHC aL > 36 | Check the Age of the Specimen and normal instrument operation |
| MCV ≤70.0 or MCV ≥110.0 or MCH ≤20.0 or RDW ≥20.0 | Perform RBC Morphology |
| Anemia | Perform RBC Morphology |
| Large Platelets | Make a slide and Scan Smear for large Platelets |
| MPV ----- | Check platelets and rerun |
| Part. ASP. | Check for clots, Volume of sample, and rerun. |
| Giant Platelets or Platelet Clumps or Thrombocytopenia or (PLT R ) | Remove PLT result before releasing the hemogram(CBC), Make a slide Perform PLT Estimate |
| Dimorphic RBC or Red cell Agglutination | Make a slide and Perform RBC Morphology |
| Abnormal Retic Pattern or Verify Retic or Reticulocytosis or Ret % aL | Confirm Result with the Second Instrument, check slide to verify presence. If slide review correlates with instrument result, release result. Document slide review on instrument printout. |
| Leukopenia or Leukocytosis | Make a slide and Scan Smear |
| NRBC % aH or NRBC % R | Remove WBC and NRBC Result before releasing the hemogram(CBC),, Make a slide and Scan Smear for NRBC |
| Cellular Interference or (WBC R ) | Remove WBC and PLT Result before releasing the hemogram(CBC), Make a slide and Perform WBC and PLT Estimate |
| Sys. Alarm | Repeat Sample and Verify Result |
| H & H Check Failed | Verify RBC Indices Check HCT 3x HGB If not Warm Sample for 30 Min |
| BA % or ≥3.0 or H&H Check Failed | Mixed Sample |

- ADDENDUM 4

DAILY QUALITY CONTROL FOR LH750

All controls must be run per shift at the designated time +/- 30 minutes.

A.	NIGHT SHIFT

0030      	Run a 5C Control on LH1 in closed mode before shutdown.  Rotate the level used each day.

0100       Shutdown LH1. Leave in shutdown for 60 minutes.

0200    	Startup LH1.  If Daily Checks pass, run the Latron primer and Latron control. (See LH help    screen for instructions to run Latron.)

Then run all three levels of 5C and Retic controls in closed mode.

File all startup printouts in the maintenance log.

0215    	Perform all above steps on LH2.

0430    	Run the 5 C Normal controls on both instruments in closed AND open modes.

Run Retic Level I control in open mode on both LH1 and LH2.

B.	DAY SHIFT

0830     	Run all three levels of 5C and Retic controls in closed mode on both LH1 and LH2.

1130     	Run 5 C Abnormal I on LH1 and LH2 in open AND closed mode.

Run Retic Level II control in open mode on LH1 and LH2.

C.	EVENING SHIFT

1600      Run all three levels of 5C and Retic controls in closed mode on both LH1 and LH2.

2100      Run 5C Abnormal II on LH1 and LH2 in open AND closed mode.

Run Retic Level III control in open mode on LH1 and LH2.

- It is critically important that each time any of the levels of the quality controls fail to follow the troubleshooting protocol.  When all levels of the quality controls are in range, document what was done to correct problem and assess patient impact.

- NEVER report patients when quality controls fail.  Call the supervisor if necessary.

- ADDENDUM 5

Calculation Formulas

MCV =  (Hct x 10) / RBC

MCH = (Hgb / RBC) x 10

MCHC = (Hgb / Hct) x 100

Absolute neutrophil count =  (neutrophil % /100) x WBC

Absolute lymphocyte count =  (lymphocyte % /100) x WBC

Absolute monocyte count =  (monocyte % /100) x WBC

Absolute eosinophil count =  (eosinophil % /100) x WBC

Absolute basophil count =  (basophil% / 100) x WBC