Technical SOP

| Title | The Gram Stain |  |
| --- | --- | --- |
| Prepared by | Ron Master, Mollie Kirchner, Kimberly Hicks, Rachel Davidson, Hospital BPT | Date: 5/4/2010 |
| Owner | Wendell R. McMillan II | Date: 07/03/2012 |
| Laboratory Approval | Effective Date: |  |
| Print Name and Title | Signature | Date |
| Refer to the electronic signature page for approval and approval dates. |  |  |
| 12-Month Review |  |  |
| Print Name and Title | Signature | Date |

TABLE OF CONTENTS

1.	Test Information	2

2.	Analytical Principle	3

3.	Specimen Requirements	3

4.	Reagents	5

5.	Calibrators/Standards	7

6.	Quality Control	7

7.	Equipment And Supplies	9

8.	Procedure	10

9.	Calculations	13

10.	Reporting Results And Repeat Criteria	13

11.	Expected Values	15

12.	Clinical Significance	16

13.	Procedure Notes	16

14.	Limitations Of Method	16

15.	Safety	18

16.	Related Documents	18

17.	References	18

18.	Revision History	18

19.	Addenda	19

# TEST INFORMATION

| Assay | Method/Instrument | Order Code |
| --- | --- | --- |
| Gram Stain | Manual | GRAM |

Synonyms/Abbreviations

Smear for bacteria, GS, Gram smear

Department

Microbiology

# ANALYTICAL PRINCIPLE

The Gram Stain is used to classify bacteria on the basis of their forms, sizes, cellular morphologies, and Gram reaction. It is also a critical test for the rapid presumptive diagnosis of infectious organisms and serves to assess the quality of clinical specimens.

Bacteria vary in their ability to retain crystal violet after treatment with an organic solvent (either alcohol or acetone).  Those organisms which retain crystal violet are considered gram positive; those which do not retain the crystal violet and are stained pink or red by the counter-stain are considered gram negative.  The difference in the staining reaction can be directly related to the cell wall structure of gram negative and gram positive organisms, although the actual mechanism is still unclear.

# SPECIMEN REQUIREMENTS

3.1	Patient Preparation

| Component | Special Notations |
| --- | --- |
| Fasting/Special Diets | Not Applicable |
| Specimen Collection and/or Timing | Not Applicable |
| Special Collection Procedures | Refer to Quest Diagnostics Incorporated Directory of Services, Specimen Collection section. |
| Other | Not Applicable |

3.2	Specimen Type & Handling

| Criteria |  |
| --- | --- |
| Type -Preferred | - Body Fluids and CSF: Specimens should be processed either by a Cytocentrifuge (preferred) or regular centrifuge and sediment used to prepare a smear. Using slides with etched rings may help locate the inoculated area. Smear should be prepared in a BSC. - Urine Specimens: Place one drop of well-mixed, un-centrifuged specimen on a slide with a sterile Pasteur pipette (do not spread the drop). Allow the drop to dry. - Sputum Specimen - Select the most purulent or most blood-tinged portion of the sample using sterile loops, rayon swabs, sticks or a pipette. THE USE of COTTON TIPPED SWABS FOR SMEAR PREPARATION IS NOT OPTIMAL. Smear should be prepared in a BSC. - Swab Specimens: - Note: If two swabs are submitted, use one to inoculate media and the other to prepare the smear. - If only one swab is received, inoculate the solid media first. (see section 8.1 step 2.) - Smear should be prepared in a BSC. |
| Type-Preferred cont. | B. Specimens from secondary sources: - Liquid medium (e.g., thioglycolate broth) Prepare smear using sterile pipette as for urine. (in 2. above). Prepare smear in a BSC. - Cultures growing on solid media: Mix enough organisms from an isolated colony with a drop of saline on a slide to obtain a barely visible turbidity. - Blood cultures: Prepare smear using a venting needle (to avoid manipulating a needle and syringe) and transfer 1 to 2 drops to the slide. Smear should be prepared in a BSC. |
| -Other Acceptable | - Tissues in transport media or in a small amount of sterile saline are acceptable. - Stool. Smear should be prepared in a BSC. |
| Collection Container | Specimens may be collected using various devices such as sterile swabs, scalpels, syringes or Pasteur pipettes. Inform clinicians not to submit syringe with needle attached. If submitting syringe, the client must remove needle, expel air and recap syringe. |
| Volume - Optimum | 2-3mL of fluid or 1-2 grams of tissue. |
| - Minimum | One properly prepared slide (air-dried and heat or methanol fixed) or any amount of material. |
| Transport Container & Temperature | - Air-dried smears should be submitted in a slide envelope, plastic or cardboard slide holder or culture swabs may be used. - Sterile screw cap cups or other sterile containers may be used for whole specimens. - Temperature same as below. |
| Stability & Storage Requirements | - Air dried smears are stable for at least 10 days at room temperature. - Specimens on swabs in transport media are stable for 2 days at room temperature. - For Gram stain purposes, respiratory specimens are stable for up to 24 hours at room temperature and 48 hrs refrigerated. Stability for culture is less. - Body Fluids - Stable for 48 hrs at room temperature - Port-a-cult – Up to 72 hours room temperature |
| Timing Considerations | None |
| Unacceptable Specimens & Actions to Take | Reject: - Broken slides or smears too thick to read. - Smears fixed with Cytology fixative. - Slides previously stained by cytology and cover slipped. - Specimens in DNA probe transports or PVA transports. - Dry swabs. |
| Compromising Physical Characteristics | Not Applicable |
| Other Considerations | Not Applicable |

# REAGENTS

Refer to the Material Safety Data Sheet (MSDS) supplied with the reagents for complete safety hazards.  Refer to the section in this procedure covering “SAFETY” for additional information. See www.med-chem.com for more reagent safety information

  - Reagent Summary

Manual Stain:

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| Gram Crystal Violet | Becton, Dickinson Cat# 212525 | 4 X 250 mL |
| Gram Iodine | Becton, Dickinson Cat# 212529 | 4 X 250 mL |
| Gram Decolorizer | Becton, Dickinson Cat# 212527 | 4 X 250 mL |
| Gram Safranin | Becton, Dickinson Cat# 212531 | 4 X 250 mL |

bioMérieux Stainer:

| Reagents | Supplier & Catalog Number | Quantity |
| --- | --- | --- |
| Reagent A: Safranin | bioMérieux 29583 | 500 mL |
| Reagent B: Iodine | bioMérieux 29586 | 500 mL |
| Reagent C: Crystal Violet | bioMérieux 29587 | 500 mL |
| Reagent Grade Methanol | Wescor SS-MeOH | 500 mL |
| Nozzle Cleaning Solution | bioMérieux 29588 | 5 L |

4.2	Reagent Preparation and Storage

NOTES:  Date and initial all reagents upon opening.  Each container must be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, (6) any special storage instructions; check for visible signs of degradation. Refer to the Material Safety Data Sheet (MSDS) for a complete description of hazards.  If a specific hazard is present, it will be noted in this procedure when the hazard is first encountered in a procedural step.

4.2.1 Manual Stain

| Reagent | Crystal Violet (1% aqueous solution) Store at room temperature |
| --- | --- |
| Container | Manufacturer’s packaging |
| Storage | Store at room temperature. |
| Stability | Stable until expiration noted on bottle. |
| Preparation | None |

| Reagent | Gram’s Iodine (Aqueous solution of iodine and potassium iodide) |
| --- | --- |
| Container | Manufacturer’s packaging |
| Storage | Store at room temperature. |
| Stability | Working solution stable for 3 months from preparation. |
| Preparation | Working solution prepared by adding entire 2.5 mL ampule of Gram Iodine 100X to 250 mL bottle of Gram Diluent. Mix thoroughly. Write date prepared and your initials on the bottle. Caution: Iodine is corrosive. Avoid contact. |

| Reagent | Gram’s Decolorizer (Mix of 50% acetone and 50% of 95% ethanol) |
| --- | --- |
| Container | Manufacturer’s packaging |
| Storage | Store at room temperature. Store decolorizer in a flame proof cabinet away from sparks or high heat. |
| Stability | Stable until expiration noted on bottle. |
| Preparation | None |

| Reagent | Gram’s Safranin (0.04% Safranin and water in reagent alcohol) |
| --- | --- |
| Container | Manufacturer’s packaging |
| Storage | Store at room temperature. |
| Stability | Stable until expiration noted on bottle. |
| Preparation | None |

4.2.2	bioMérieux Stainer

| Reagent | Crystal Violet |
| --- | --- |
| Container | 500 ml |
| Storage | Store at room temperature away from direct light. |
| Stability | Stable until expiration noted on bottle. |
| Preparation | Ready for use |

| Reagent | Gram’s Iodine |
| --- | --- |
| Container | 500 ml |
| Storage | Store at room temperature away from direct light. |
| Stability | Sensitive to light, heat, and oxygen. Stable until expiration noted on bottle. |
| Preparation | Ready for use. |

| Reagent | Decolorizer with Safranin |
| --- | --- |
| Container | 500 ml |
| Storage | Store at room temperature away from direct light. |
| Stability | Photosensitive, may slowly form peroxidase. Stable until expiration noted on bottle. |
| Preparation | Ready for use |

| Reagent | Reagent Grade Methanol |
| --- | --- |
| Container | 500 ml |
| Storage | Store at room temperature away from direct light. Store in a flame proof cabinet away from heat, flame, and sources of ignition. |
| Stability | Stable until expiration noted on bottle. |
| Preparation | Ready for use |

# CALIBRATORS/STANDARDS

5.1	Calibrators/Standards Used

Not applicable

# QUALITY CONTROL

6.1	Controls Used

- Gram stain reagents should be tested using known gram-positive and gram-negative organisms.

  - Gram Positive Control:   Staphylococcus aureus ATCC25923

  - Gram Negative Control:  Escherichia coli ATCC25922

Additional organisms such as a Bacillus species may be added in situations where problems with decolorization or interpretation are encountered.

Use of a clinical specimen such as a sputum smear may also be considered.

6.2	Control Preparation and Storage

NOTE:  Date and initial all controls upon opening.  Each container should be labeled with  (1) substance name,  (2) lot number,  (3) date of preparation,  (4) expiration date, (5) initials of tech, and (6) any special storage instructions; check for visible signs of degradation.

Refer to the control organism product insert sheet for preparation, storage and handling instructions. Make a lightly turbid suspension of both control organisms in 0.5 mL of saline. Pipette or apply several loops full of the suspension to the surface of a glass microscope slide and allow to air dry. Fixing with 95% methanol is preferred.  Alternatively, heat fix (allow slides to cool before staining).  Store smears in an appropriately labeled box at room temperature. Control slides are stable for 1 year from preparation.

| Preparation | Prepare a suspension of mixed Staphylococcus aureus ATCC 25923 and Escherichia coli ATCC 25922 in distilled water. Place a thin drop on the: surface of a slide and allow it to air dry. Fix the air-dried smear on the heat block no longer than 10 minutes and stain the smear in the same manner as the test slides. Examine microscopically. |
| --- | --- |
| Storage/Stability | Prepared slides can be made in advance. Store smears in an appropriately labeled box at room temperature. Control slides are stable for 1 year from preparation. |

6.3	Frequency

- Quality control testing is performed before first patient use with each lot number and/or shipment of any reagent. Additionally, quality control is performed daily.

- Refer to specific test procedures or department QC plan for QC frequency requirements.

- See the Corporate SOP entitled “QC Frequency for Batch, Random Access and STAT testing” for more details on minimum requirements.

  - Tolerance Limits

- Do not interpret tests unless the controls yield expected results as below.

| Control Organism | Expected Results |
| --- | --- |
| Staphylococcus species | Deep purple (gram positive) cocci must be observed. |
| Escherichia coli ATCC 25922 | Pink (gram negative) rods must be observed. |
| Bacillus species | Purple (gram positive), may be stain irregularly (purple and pink) |

- Corrective Action:

- All steps taken in response to QC failures must be documented, including: a description of the QC failure, the root cause of the problem, actions taken to correct the problem, how patient samples were handled, and the date and initials of the person recording the information.

- Supervisors may override rejection of partial or complete runs based on findings in the corrective-action investigation and based on established override criteria that have been approved by the Medical Director. The rationale for the override must be thoroughly documented.

- See the National QC SOP “QC Responsibilities” for more details.

6.5	Review Patient Data

Review patient results for unusual patterns, trends or distributions in patient results, such as an unusually high percentage of abnormal results.

6.6	Documentation

Refer to local policies and procedures for QC documentation and to Quest Diagnostics records management program for record retention requirements.

6.7	Quality Assurance Program

Refer to National and local policies and procedures for other quality assurance activities applicable to this procedure

# EQUIPMENT and SUPPLIES

7.1	Assay Platform

- Manual stain or an optional automated slide stainer may be used.

  - Equipment

- Heat block or Slide warmer 45-62C or Bacti Cinerator

- Drain rack

- Microscope

- Centrifuge

- Biosafety cabinet

- bioMérieux stainer

    - Optional materials, depending on specimen source.

- Sterile scalpel, forceps

- Cytocentrifuge

- Sterile tubes, screw cap

- Tissue grinder

- Vortex mixer

- Sterile saline, sterile broth

  - Supplies

- Glass slides, frosted ends desirable

- Slide labels

- Pencil

- Sterile Applicator Sticks or swabs.

- Pasteur pipettes, or wood applicator sticks (sterile)

- Sterile Disposable Inoculating  Loops

- Absorbent paper

- Immersion oil

- Lens paper

- Lens cleaner

# PROCEDURE

NOTE:   For all procedures involving specimens, buttoned lab coats, gloves, and face protection are required minimum personal protective equipment.  Report all accidents to your supervisor. The package insert for a new lot of kits must be reviewed for any changes before the kit is used.

| 8.1 | Specimen Prep and Manual Staining Procedure |
| --- | --- |
| 1. | A smear for Gram stain may be prepared from clinical specimens, broth cultures, or colonies growing on solid media. Appropriate Personal Protective Equipment (PPE) must be utilized to prepare smear. A biological safety cabinet must be utilized to prepare all smears. Label slide using pencil with the accession number, patient last name, source, and date. Refer to section 3.2 for additional PPE and BSC requirements. |
| 2. | For clinical specimens on swabs, roll the swab gently across a slide (cleaned with 85% alcohol) or use pre-cleaned slides to avoid destruction of cellular elements and disruption of bacterial arrangements, covering an area no more than the size of a quarter. Note: If only one swab is received, and a culture ordered as well, inoculate the solid media first, and then prepare the smear prior to placing the swab tip into liquid media. Alternately, vortex the swab in 0.5-1 ml of sterile saline and use the resulting suspension to inoculate media and prepare smear. Fix the slide with 95% methanol as this prevents the lysis of RBCs, avoids damage to the host cells, as well as bacteria and results in a clear background. Alternatively, heat fix at 45-62C until dry for a maximum of 10 minutes and allow the smear to cool before staining. |
| 3. | Apply crystal violet to prepared slide. Let sit for 30-60 seconds, and then rinse slide gently with running tap water. Caution: Excessive rinsing in this step could cause crystal violet to be washed from gram-positive cells. |
| 4. | Apply iodine solution for 30 -60 seconds, and then rinse slide gently with running tap water. |
| 5. | Decolorize with alcohol-acetone for up to 10-15 seconds or until run-off is clear while the slide is held at and angle. (Thickness or type of smear may have an impact on the decolorizer timing.) Rinse slide gently with running tap water. |
| 6. | Counter stain slide with Safranin for 30-60 seconds. Rinse slide with a gentle flow of tap water and air dry or blot on absorbent paper. |

| 8.2 | Repeat Staining |
| --- | --- |
| 1. | If after preliminary review, the smear does not appear adequately stained, re-staining may be performed. |
| 2. | Gently blot off any oil. |
| 3. | Flood smear with decolorizer |
| 4. | Rinse gently with tap water. |
| 5. | Repeat staining procedure steps 8.1.4 to 8.1.6 above |

| 8.3 | Automated Gram Stain Test Procedure/ bioMérieux Stainer |
| --- | --- |
| 1. | Air-dry smear. |
| 2. | Heat fix on heat block/slide warmer no longer than 10 minutes. |
| 3. | Check reagent level in each reagent bottle on the stainer to verify sufficient quantity for staining. Reagents must be placed according to the corresponding letter on the stainer: - Safranin Solution - Iodine - Crystal Violet - Deionized Water - Reagent grade Methanol |
| 4. | Load slides into carousel slots with smears facing clockwise. If loading fewer than 12 slides, balance the rotor by placing slides in opposing positions. Use blank slides as necessary to balance. Refer to the PREVI™ Color Gram User’s Manual for instructions on properly loading the stainer. |
| 5. | Replace carousel lid and place the carousel on the instrument hub. Close the lid. Caution: Do not lock carousel lid in place while carousel is sitting on the drive hub. |
| 6. | For instructions on changing settings on the stainer as needed, refer to the PREVI™ Color Gram User’s Manual. |
| 7. | Press RUN to begin cycle. |
| 8. | An audible signal sounds at the end of the staining cycle. Remove the rotor and unload slides for reading. |
| 9. | Examine microscopically under oil immersion. |

| 8.4 | Daily Maintenance |
| --- | --- |
| 1. | Follow the procedures as outlined in the PREVI™ Color Gram User’s Manual. |
| 2. | Check reagent levels and expiration dates. |
| 3. | Run the CLEAN cycle. |
| 4. | Perform a PATTERN test. |
| 5. | Run the CLEAN cycle alcohol purge stage. |
| 6. | Leave the instrument in STAND BY. |
| 7. | Spray and wipe the interior of the bowl and the nozzles with 70% ethanol. |
| 8. | Clean the instrument lid with 70% ethanol. |
| 9. | Record that the procedures have been completed by recording in the appropriate section of the Preventive Maintenance Chart. |

| 8.5 | Weekly Maintenance |
| --- | --- |
| 1. | Follow the procedures as outlined in the PREVI™ Color Gram User’s Manual. |
| 2. | Disassemble and clean all of the nozzles. |
| 3. | Perform the B line flushing procedure. |
| 4. | Perform the PATTERN test. |
| 5. | Perform the VOLUME test. |
| 6. | Record that the procedures have been completed by recording in the appropriate section of the Preventive Maintenance Chart. |

| 8.6 | Monthly Maintenance |
| --- | --- |
| 1. | Follow the procedures as outlined in the PREVI™ Color Gram User’s Manual. |
| 2. | Disassemble and clean all of the nozzles. |
| 3. | Perform the B line flushing procedure. |
| 4. | Perform the Pattern test. |
| 5. | Perform the Volume test. |
| 6. | Record that the procedures have been completed by recording in the appropriate section of the Preventive Maintenance Chart. |

# CALCULATIONS

Not applicable

# REPORTING RESULTS AND REPEAT CRITERIA

10.1	Interpretation of Data

| Result | Interpretation |
| --- | --- |
| Purple or deep blue staining | Gram positive |
| Pink or light red staining | Gram negative |
| Both purple and pink staining | Considered gram variable when seen in smears of culture isolates. This is often seen with Bacillus, Clostridium, anaerobic cocci, Moraxella and Gardnerella vaginalis. It is suggested that fresh subcultures be performed and the stain repeated. A lead tech or supervisor must review the Gram stain before reporting as gram variable. |

10.2   Clinical Specimens

Check that white cells, epithelial cells and proteinaceous material allow the crystal violet stain to wash out with decolorization. Smears that have been correctly prepared and stained should appear completely gram negative (pink).  Smears which appear purplish-pink have not been adequately decolorized and should be re-stained.  Thick areas of the smear may be stained purple, perhaps with precipitated crystal violet present.  Do not read in these areas.  After determining that the staining is adequate, proceed to look for bacteria.  Do not mistake thin crystal violet precipitation needles for gram-positive bacillus-shaped bacteria.

Examine the slide only in the central area where the specimen has been applied. Edges of slides may be contaminated with bacteria.

10.3   Smears of Isolates

Evaluate only in areas where the organisms are well dispersed as gram negative organisms may not adequately decolorize in thick areas.  Avoid evaluating morphology in areas where the inoculating loop has scraped the smear as this can cause distortion of organisms such as Neisseria gonorrhoeae, making them appear rod-shaped or bizarre.

  - Units of Measure

Determine the average number of cells and bacteria in 20 to 40 fields of the smear. Skip fields where there are no cells or bacteria, and do not average these fields in the count if there are fields where cells and/or bacteria are present.

If cells and organisms are observed, use the following guidelines for quantitation and message codes to report

Examine smears prepared from clinical specimens under low power for evidence of inflammation. Enumerate host cells in a Gram smear using low-power objective (10 x objective with 10x eyepiece) and report relative number accordingly. Use the following guidelines for quantitation and message codes to report:

| Average Quantity Observed 10X LPF | Report As |
| --- | --- |
| <1 per lpf | Rare |
| 1—9 per lpf | Few |
| 10—24 per lpf | Moderate |
| >25 per lpf | Many |

Count bacteria, yeast and/or hyphal elements under oil immersion (100x objective with 10x eyepiece) and report relative numbers from areas with cells. Use the following guidelines for quantitation:

| Enumeration of Bacteria and Yeast |  |
| --- | --- |
| Average Quantity Observed (100X) | Report As |
| <1 per oil immersion field | Rare |
| 1 to 9 per oil immersion field | Few |
| 10 – 24 per oil immersion field | Moderate |
| > 25 per oil immersion field | Many |

10.5	Clinically Reportable Range (CRR)

Semi-quantitative test, reported with host cells and organisms. See section 10.4 and 10.6.

  - Repeat Criteria and Resulting

| If the Gram stain is CSF, sterile body fluid or tissue and organisms or WBCs are present: | Initially 2 technologists must read the Gram stain. If the interpretation is not definitive, a lead technologist, supervisor, or video microscopy remote technologist must be used to determine the correct interpretation. |
| --- | --- |
| If the Gram stain is for other specimen types: | If the interpretation is not definitive, a second technologist must read the Gram stain. If interpretation is not definitive, a lead technologist, supervisor, or video microscopy remote technologist must be used to determine the correct interpretation. |

| IF the body site is … | THEN… |
| --- | --- |
| Tissues and Unspun Fluids | Report all organisms seen. Report the presence and relative amount of WBC’s and any organism(s) seen. |
| Normally Sterile Fluids, BAL, Bronchial Brushings | Report presence and relative amount of epithelial cells, WBC’s and any organism(s) seen. |

| IF the result is … | THEN… |
| --- | --- |
| Purple/blue bacteria | It is gram-positive rods or cocci. |
| Pink/red bacteria | It is gram-negative rods or cocci. |
| If the STAT gram stain result of a sterile body fluid is positive and there is no on-site supervisor or senior tech. to verify it | Repeat gram stain. Open up a new box of slides, use sterile pipette and observe aseptic technique. If result remains the same, go ahead and report it out. |
| The nuclei of WBCs stain purple and/or the nuclei of epithelial cells stain red. | Repeat the gram stain. |

# EXPECTED VALUES

11.1	Reference Ranges

- Gram-negative organisms stain pink.

- Gram-positive organisms stain deep violet

- The expected distribution of bacteria and cells is based on the body site and type of infection.

  - Priority One and Two Limits

- The extremely serious nature of central nervous system infections demand that CSF specimens be treated as an emergency specimen and should be worked on right away. A Gram stain should always be performed and the smear result reported to the clinician immediately.

- Refer to the current Critical Values Procedure.

11.3	Priority Three Limit(s)

Refer to the current Critical Values Procedure.

# CLINICAL SIGNIFICANCE

Clinical significance varies with source. The Gram stain is a critical test for the rapid, presumptive diagnosis of infectious agents and also serves to assess the quality of clinical specimens.

# PROCEDURE NOTES

13.1	FDA Status: bioMérieux  Automated Stainer:  Approved/Cleared

Validated Test Modifications:  None

  - Possible Sources of Error

- Some specimens may yield negative results unless they are centrifuged (Cytospin). Specimens of body fluids especially CSF require centrifugation.

- Thick specimens or specimens with a high mucous content are difficult to decolorize and may appear falsely gram positive.

- Use of glass slides that have not been pre-cleaned or degreased (Storing slides in a jar with 95% ethanol will ensure clean slides, drain excess alcohol before use).

- Do not over heat fix slide. Overheating leads to over-decolorization and will cause all cells to appear Gram-negative.

- Over-decolorization: Often caused by using a strong decolorizer or by leaving the decolorizer on the slide for too long.

- Prolonged washing between any of the steps can cause over-decolorization

  - Special Precautions

- Requests for Gram staining of material processed in Mycobacteriology must be  denied as the NaOH causes lysis of gram negative organisms

# LIMITATIONS OF METHOD

- Some organisms may give a Gram stain reaction which is opposite from what may be expected.  In particular, some Bacillus and Clostridium species stain gram negative as the cultures age.  Other species in these groups are frankly gram negative.

- Some organisms do not stain well with the Gram stain and can be overlooked.  This is particularly true of some anaerobes and small gram negative rods (Brucella, Campylobacter spp, Fusobacterium spp, Bacteroides spp, Bordetella spp, Franciscella spp, Haemophilus spp) in addition to Legionella spp.  Carefully observe the smear (particularly blood smears) so as not to overlook these organisms. Longer exposure to the counter stain can enhance faintly staining Campylobacter and Fusobacterium spp

- Staphylococci, streptococci and pneumococci in clinical preparations may decolorize and appear gram negative.  Always ask the question “Is this organism reasonable from this source and is this morphology correct?”

- Some streptococci (S. mutans) may look like rods when stained from solid media.  Gram stains from liquid media should show the more typical coccus morphology.

- Nutritionally-deficient streptococci and organisms recovered from patients receiving antibiotics may be pleomorphic.

- Clue cells are epithelial cells completely covered by tiny, gram variable bacilli and coccobacilli, suggestive of bacterial vaginosis.

- Gram stain-positive, culture-negative specimens may be the result of contamination of reagents and other supplies, presence of antimicrobial agents, or failure of organisms to grow under usual culture conditions (media, atmosphere, etc.).

- False Gram stain results may be related to inadequately collected specimens.

14.1	Analytical Measurement Range (AMR)

Not applicable

14.2	Precision

Not applicable

14.3	Interfering Substances

- Be wary of artifacts in Gram stains, particularly from clinical material.  Precipitated crystal violet may appear as gram positive cocci, but close examination will show that these “cocci” vary in size and appear in “fused” groups.  Smears from cultures of Neisseria may have areas where there appear to be gram negative rods.  These areas are usually next to a “scrape” in the smear and probably represent proteinaceous material.  If gram negative rods are present, they should be distributed throughout the smear.

- It has been noted that smears from eyes used for corneal transplants and from eye aspirates contain material which resemble gram positive coccobacilli. This fact should be kept in mind when evaluating Gram stains from eye specimens.

14.4	Clinical Sensitivity/Specificity/Predictive Values

Not applicable

# SAFETY

The employee has direct responsibility to avoid injury and illness at work.  Nearly all harmful exposures to infectious substances and chemicals, and other injuries, can be avoided with effective training and consistent safe work practices.

Become familiar with the Environmental Health and Safety (EHS) Manual to the learn requirements on working safely and protecting the environment from harm.  Although lab work typically focuses on the hazards of working with specimens and chemicals, we must also control other important hazards.

- Slips, trips, and falls cause many serious injuries.  Please ensure that spills are cleaned quickly (to avoid slippery floors) and that you can see and avoid obstacles in your path.

- Ergonomic injuries result from performing tasks with too much repetition, force, or awkward position.  Ergonomic injuries include strains and back injuries.  Learn about ergonomic hazards and how to prevent this type of injury.

- Scratches, lacerations, and needlesticks can result in serious health consequences.  Attempt to find ways to eliminate your risk when working with sharp materials.

Report all accidents and injuries immediately to your supervisor or the business unit Environmental Health and Safety Manager or Specialist.

# RELATED DOCUMENTS

N/A

# REFERENCES

- Betty A. Forbes, Daniel F. Sahm, Alice S. Weissfeld. 2007. Bailey and Scott’s Diagnostics Microbiology, 12th edition. Mosby, St. Louis, Mo.

- McClelland, R.  2001.  Gram's Stain: the Key to Microbiology.  Med Lab Observer.  April  pp. 20- 31.  Web site: http://www.microbiologyforum.org/PMFNews/PMFNews.12.02.0602.pdf

- Package Inserts and MSDS forms from the Becton Dickinson Web site.

- Miller, J. Michael. 1999. Specimen Management in Clinical Microbiology, Second edition, ASM press , Washington, DC

- Isenberg, H.D. ed. 1992. Clinical Micro Procedures Handbook. American Society for Microbiology, Washington, DC.

- PREVI™ Color Gram User’s Manual version B 03/2009. bioMérieux S.A. 69280 Marcy l’Etoile, France.

# REVISION HISTORY

| Version | Date | Section | Reason | Reviser | Approval |
| --- | --- | --- | --- | --- | --- |
|  | 6/6/11 |  | SOP QDHOS701 adapted as PHM.012 | A. Foster | R. Levy |
|  | 07/05/2012 |  | Supercedes PHM.012 due to change in Gram Stain platform | W.McMillan | R. Levy |

# ADDENDA

| Addendum | Title |
| --- | --- |
| 1 | Providence Codes |

ADDENDUM 1

| Message | Providence Hospital Mnemonic |
| --- | --- |
| * If no organisms are present, report “No organism seen” | NOS |
| WBC’s | WBCS |
| Epithelial cells | EPIT |
| Gram negative bacilli | GNR |
| Gram negative coccobacilli | GNCB |
| Gram negative diplococci | GNDC |
| Gram positive cocci | GPC |
| Gram positive coccobacilli | GPCB |
| Gram positive cocci in chains | GPCC |
| Gram positive cocci in clusters | GPCL |
| Gram Positive cocci in pairs | GPCP |
| Gram positive bacilli | GPB |
| Branching Gram positive bacilli | BGPB |
| Gram variable rods | GVR |
| Yeast | YST |
| Budding Yeast | BYST |