NVML Microbiology IQCP

Vitek 2 Commercial Antimicrobial Susceptibility Testing

Facility: NVML 750 W. High St, Lima, Ohio

Test System: Commercial Antimicrobial Susceptibility Testing (AST) Vitek2 System

Test System Primary SOPs include: Vitek 2 Antimicrobial Susceptibility Testing (AST) nvml.micro.vitek.001.susceptibility(AST) Vitek II nvml.micro.vitek.013.vitekII.doc

Historical Quality Review: CLIA ’88 requires testing of QC strains daily (or each day patient’s tests are performed) for AST. Previously CLIA inspector guidelines recognized use of CLSI standards M100 and M07 which indicate that weekly testing of QC strains is acceptable following documentation of satisfactory daily QC testing. This laboratory has been following the CLSI standards for over 25 years without any significant QC problems. It is rare to encounter an out-of-range result with a QC strain that indicates a test system problem. Nearly all testing errors or delays in reporting occur with individual patient isolates and these errors are unrelated to testing QC strains or a problem with testing reagents or equipment. Processes to mitigate patient reporting errors and delayed reports are addressed in this IQCP.

Information Used to Conduct Risk Assessment

Regulatory and Accreditation Requirements:

Our IQCP is in compliance with CMS and CAP ( ALL COMMON and MICROBIOLOGY) requirements. CAP – MIC 21460 MIC 21624 MIC 21820 MIC 21840 MIC 21910 MIC 21930 MIC 21940 MIC 21943 MIC 21944 MIC 21950

Method verification: Instrument received and test system verification completed in year 2012. Subsequent verifications performed when new drugs were added in 2013 (AST GN74). All verifications were deemed acceptable by the Microbiology Director. Documentation filed in the Microbiology cabinet.

Training of personnel: Completion of training documented in CAP Competency Program.

Competency Assessment: New employees 6 months after initial training and annually thereafter. Documentation filed in CAP Competency Program.

Proficiency Testing: Rotate personnel; all personnel review results. Proficiency testing records filed in CAP Survey Binders and CAP Competency Program.

Quality Control: CLIA ’88 and Accrediting Agency require testing of QC strains daily (or each day patient’s tests are performed) for AST. NVML has developed an IQCP to modify frequency of testing QC strains. Each new lot and/or shipment has the recommended control organisms ran to confirm viability of the AST cards. Twenty day QC was performed and acceptable reducing QC to weekly testing of the GP-AST67 and GN- AST74 susceptibility cards.

Test System Information:

Manufacturer: Package insert contains system performance data and describes testing principle and procedure, QC recommendations, and limitations. Package inserts are located in the Micro cabinet. Manufacturer alerts and bulletins are located in the V2 Binder at the workbench. Operator’s manual including troubleshooting guide is located at workbench. NOTE: No risks were identified when looking at the PI for this system. Both Limitations and Performance data sections were closely reviewed to identify possible risks associated with the system and none were found.

Scientific publications used during collection of information for RA: CLSI document M07-A10. 2015. Manual of Clinical Microbiology 11th Edition 2015. NOTE: Upon review of these references no risk factors for using this test system were identified for our laboratory.

Summary of in-house data from routine testing of QC strains: QC testing was performed according to Vitek 2 Antimicrobial Susceptibility Testing (AST) nvml.micro.vitek.001.susceptibility (AST). Review of QC records for the past 12 months that contained over 1300 results demonstrated: - 0.02% occurrence of random QC errors that corrected upon repeat testing.(Incorrect ATCC strain used) - 0.00% occurrence of potential system QC errors that required corrective action.

Summary of in-house data from routine instrument performance checks: Instrument checks were done according to Vitek II nvml.micro.vitek.013.vitekII.doc Review of instrument QC records for the past 12 months that contained approximately 52 routine checks of the Vitek 2 instrument and 2 reports following scheduled maintenance performed by the company’s service engineer revealed no instrument performance problems that would impact patient results.

Summary of corrected reports and physician complaints: Documentation located online under Quality Measures PI variances. Review of reporting errors identified prior to report release, corrected reports and physician complaints for the past 12 months revealed: - 11 corrected reports showed errors were due to one or more of the following: 1) reporting inappropriate antimicrobial agents for the species/body site (n=1) 2) erroneous MIC or interpretation due to application of inappropriate interpretive criteria (n=1) 3) failure to add the correct reporting comment (n=9) - 1 formal physician complaint revealed: 1) Multiple strains of resistant Pseudomonas on a patient. The physician wanted a better choice of oral antibiotics. It was explained that this organism has very few oral drugs available for treatment. Note: during this review of corrected reports and physician complaints, none of the errors could have been avoided by any changes in protocol for testing of QC strains including frequency of testing QC strains.

Risk Assessment and Determination of Risk Level

| Frequency of occurrence: Unlikely (once every 2-3 years) Occasional (once per year) Probable (once per month) Frequent (once a week) | Severity of harm to patient: Negligible (temporary discomfort) Minor (temporary injury; not requiring medical intervention) Serious (impairment requiring medical intervention) Critical (life threatening consequences) |  |
| --- | --- | --- |
| Risk Level: |  |  |
| Risk level for any Risk Factor that is “Not Acceptable” must be addressed in the IQCP. Risk level for any Risk Factor that is “Acceptable” may be included in the IQCP at the discretion of the Laboratory Director. Note: Patient response plays a significant role in addition to AST results in guiding antimicrobial therapy and provides a limited safeguard for preventing harm in patients for which erroneous AST results are reported or results are delayed. |  |  |

Risk Acceptability Matrix

| Probability of Harm | Negligible | Minor | Serious | Critical |
| --- | --- | --- | --- | --- |
| Frequent | Not Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Probable | Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Occasional | Acceptable | Acceptable | Acceptable | Not Acceptable |
| Unlikely | Acceptable | Acceptable | Acceptable | Acceptable |

Risk Acceptability Assignment

| Risk Factor (Possible Sources of Error) | Frequency of occurrence | Severity of harm to patient | Risk Level |
| --- | --- | --- | --- |
| Preanalytical |  |  |  |
| Specimen (Primary): |  |  |  |
| Patient identification | probable | minor | Not Acceptable |
| Collection/container/volume | frequent | negligible | Not Acceptable |
| Integrity | frequent | negligible | Not Acceptable |
| Transport | frequent | negligible | Not Acceptable |
| Storage | probable | negligible | Acceptable |
| Specimen (Organism): |  |  |  |
| Clinically relevant | probable | minor | Not Acceptable |
| Colony age/viability/sampling | frequent | minor | Not Acceptable |
| Media type | unlikely | minor | Acceptable |
| Pure isolate | frequent | serious | Not Acceptable |
| Inoculum suspension preparation | occasional | minor | Acceptable |

| Risk Factor (Possible Sources of Error) | Frequency of occurrence | Severity of harm to patient | Risk Level |
| --- | --- | --- | --- |
| Analytical |  |  |  |
| Testing Personnel: |  |  |  |
| Training | probable | serious | Not Acceptable |
| Competency | probable | serious | Not Acceptable |
| Experience | probable | serious | Not Acceptable |
| Proficiency Testing | unlikely | negligible | Acceptable |
| Staffing | occasional | minor | Acceptable |
| Reagents: |  |  |  |
| Shipping/receiving/storage | occasional | minor | Acceptable |
| Expiration dates | unlikely | minor | Acceptable |
| Preparation/use | probable | minor | Not Acceptable |
| QC strain storage/prep | occasional | negligible | Acceptable |
| Environment: |  |  |  |
| Temperature/airflow/humidity/ ventilation | unlikely | negligible | Acceptable |
| Utilities | occasional | minor | Acceptable |
| Space | unlikely | negligible | Acceptable |
| Noise/vibration | unlikely | negligible | Acceptable |
| Test System: |  |  |  |
| Mechanical/electronic stability of instrument/equipment/jam | occasional | negligible | Acceptable |
| Software/antimicrobial reporting rules | frequent | serious | Not Acceptable |
| Transmission of results to LIS | unlikely | serious | Acceptable |
| Postanalytical |  |  |  |
| Test Results: |  |  |  |
| Results reported within 5 days | probable | serious | Not Acceptable |
| Transmission of results to Electronic Health Record | occasional | serious | Acceptable |
| Review reported results | frequent | serious | Not Acceptable |
| Clinician feedback | probable | serious | Not Acceptable |

Risk Assessment

| Possible Sources of Error |  | How can identified sources of error be reduced? |
| --- | --- | --- |
| Risk Factor | Possible Error |  |
| Preanalytical |  |  |
| 1A: Specimen - Biological | - Improper specimen procurement/ handling/processing | - Adhere to procedures: nvml.037 Labeling and Handling of Laboratory Specimens, nvml.054 Outpatient Test Request, nvml.058 Verification of Patient Identity, nvml.073 Specimen Rejection Criteria and nvml.micro.plating.005.Rejection of Specimens that addresses patient identification and specimen collection, labeling, transport, storage and remedial actions to control improperly handled specimens or delayed specimens. - Errors in specimen collection and handling are reviewed with involved individuals and frequent errors are addressed with all staff involved with patient specimen collection and handling. (Rejection Log). During initial training and competency assessment, emphasize: - Proper specimen handling/processing is the most critical part of any test - Failure to streak correctly (no isolated colonies) and delayed incubation may result in delayed AST reports |
| Patient/specimen identification |  | See above (Specimen) |
| Collection/container/ volume |  | See above (Specimen) |
| Integrity |  | See above (Specimen) |
| Transport |  | See above (Specimen) |
| Storage |  | See above (Specimen) |
| 1B: Specimen - Organism |  |  |
| Clinically relevant | - Clinically irrelevant organisms tested - Additional species may be significant in select patient types (e.g., immunosuppressed) - Physicians may request testing of isolates that are not clinically relevant; requests may be inappropriate and results misleading | - Vitek 2 Antimicrobial Susceptibilty Testing (AST) nvml.micro.vitek.001.susceptibility(AST) describes selecting organisms to test for AST based on organism ID, specimen source and quantity . - Physicians can request additional testing in select patients; comment added to media report indicating physician/health care provider special request. Supervisor/director may discuss with requesting physician those requests that may be inappropriate. |
| Old or less viable | - Colonies on source plate > 1 day old | During initial training and competency assessment, emphasize: - Organism growth requirements (especially S. pneumoniae) |
| Media type | - Media for inoculum source other than that recommended is used - Panel fails to support growth of test organism | During initial training and competency assessment, emphasize: - Appropriate media for inoculum - Species that can be reliably tested by test system based on manufacturer’s recommendations |
| Pure isolate | - Mixed inoculum or contaminated panel | - Solicit regular feedback on streaking of primary plates (for isolated colonies) - Inoculate purity plate - Daily review of AST profiles for aberrant results possibly due to mix/contamination During initial training and competency assessment, emphasize: - Proper organism selection for inoculum preparation - Risks of selecting “young” colonies or poorly isolated colonies - Potential sources of contamination during testing process - Impact of delayed results (if retesting needed) |
| Inoculum suspension | - Overinoculation or underinoculation - Use of nonviable colonies | - Turbidity meter for inoculum standardization During initial training and competency assessment, emphasize: - Proper inoculum suspension preparation - Impact of overinoculation (false R) or underinoculation (false S) |
| Species appropriate | - Testing of species not indicated for test system | During initial training and competency assessment, emphasize: - Species that can be reliably tested by test system based on manufacturer’s recommendations |
| Analytical |  |  |
| 2: Testing Personnel | - Incompletely trained - Unaware of updated recommendations for AST/reporting | During initial training and competency assessment, emphasize: - Key aspects of AST to include those described in this IQCP - Supervisor annually review any changes in AST recommendations described by accrediting agencies or standards organizations |
| Training |  | See above (Testing Personnel) |
| Competency |  | See above (Testing Personnel) |
| Experience |  | - Supervisor review AST reports generated by new employees prior to release for the first two months of their employment |
| Proficiency Testing |  | - All staff read (and sign off) on PT sample critiques |
| Staffing | Inadequate to perform testing without errors | - Supervisor to annually review appropriate staffing needs for AST and schedule staff accordingly |
| 3: Reagents |  | During initial training and competency assessment, emphasize standard rules to always: - Take responsibility for reagents/supplies (all staff) - Maintain reagents at proper storage conditions - Check expiration dates - Perform required QC |
| Receiving/storage | - Incorrect ordering - Depleted reagent supply - Reagent integrity compromised | - Designated staff member(s) assigned to inventory (order/receipt) AST reagents to ensure inventory properly maintained and testing materials are handled appropriately on receipt |
| Expiration dates |  | See above (Reagents) |
| Preparation/use | - Use incorrect panel/card for select organism | - Use color codes on boxes of panels |
| QC strain storage/prep | - QC out of control due to improper QC strain maintenance | During initial training and competency assessment, emphasize: - Proper maintenance of QC strains (limited number of subcultures) - Potential sources of QC failures - QC troubleshooting - QC frequency - Role of QC strains versus other QA measures to ensure reliable reporting of patient results |
| 4: Environment | - Results not reported (ancillary equipment failure, e.g., incubator malfunction) | - Instrument installed at a location following manufacturer’s suggestions. During initial training and competency assessment, emphasize standard rules for: - Take responsibility for any possible instrument/ environmental problem (out of the ordinary observation)(all staff) - Equipment maintenance - Temperature recording (done automatically with continuous monitoring device) - Electrical supply |
| Temperature/airflow/humidity/ ventilation |  | See above (Environment) |
| Utilities |  | See above (Environment) |
| Space |  | N/A (sufficient space available) |
| Noise/vibration |  | See above (Environment) |
| 5: Test System |  | During initial training and competency assessment, emphasize standard rules for: - Take responsibility for any possible instrument/test system problem (out of the ordinary observation) |
| Mechanical/electronic/jam | Results not reported (e.g., instrument malfunction and/or aborted test) | - Perform preventive maintenance according to recommended schedule During initial training and competency assessment, emphasize: - How to avoid and resolve jams |
| Software/antimicrobial reporting rules | - Inappropriate drugs reported - MICs interpreted incorrectly - Erroneous results reported - Report comments missing or inappropriate for the culture | - Software rules address (and flag) most (but not all) potential errors to be checked by tech; sometimes note for tech follow up action printed on internal report - Software flags unusual results requiring supervisor review - Daily supervisor (or supervisor designee) review of reported results During initial training and competency assessment, emphasize: - Intrinsic resistance patterns of commonly encountered species - Results requiring follow up action (e.g., confirmation by repeat testing) - Results requiring consultation with supervisor/director |
| Transmission of results to LIS | - Incorrect transmission of results - Delay in transmission of results | - Daily supervisor (or supervisor designee) review of reported results - Annual check of test system- LIS computer interface - QA monitor for time to reporting AST results |
| Postanalytical |  |  |
| 6: Test Results |  | - Supervisor maintains summary of incorrect results released and reviews this summary with technical staff. - QA monitor for time to reporting AST results During initial training and competency assessment, emphasize: - Need for timely results to guide therapy and identify potential multidrug resistant organisms that might require patient isolation - Reporting preliminary results (timely reporting) |
| Results reported within 5 days | - Results delayed beyond that expected for organism type | See above (Test Results) |
| Transmission of results to Electronic Health Record | - Incorrect transmission of results - Delay in transmission of results | See above (Test Results) |
| Review reported results | - Inappropriate drugs reported - Erroneous results reported - MICs interpreted incorrectly - Report comments missing or inappropriate for the culture | See above (Test Results and Test System) Note: results are checked at multiple steps by tech and then by supervisor |
| Clinician feedback | - Complaints/suggestions regarding delayed results and potential erroneous results | See above (Test Results) - Incorporate suggestions into QA plan, as appropriate. |

Final QCP for Vitek 2 AST System

Based on our risk assessment and Quality Assessment, the QCP consists of following the instructions that are provided in detail in Vitek 2 Antimicrobial Susceptibility Testing (AST) nvml.micro.vitek.001.susceptibility(AST) for Performance of AST and are summarized here.

Testing of appropriate QC strains on each new lot/shipment of panels before or concurrently with placing these materials into use for testing patient’s isolates.

Testing of appropriate QC strains on each panel type weekly.

Testing of appropriate QC strains on each panel type after major system maintenance or software upgrade concurrently with placing the equipment back into service.

Testing of appropriate QC strains against any new antimicrobial agent added to the panel at least 15 times (over a minimum of 5 days) prior to resuming weekly QC testing of the panel; accomplished during performance of verification study.

Recording and evaluating QC results according to QC acceptability criteria as defined in Vitek 2 Antimicrobial Susceptibility Testing (AST) nvml.micro.vitek.001.susceptibility(AST) for Performance of AST. Any out-of-range result is immediately investigated and corrective action performed prior to releasing any patient results.

| Quality Assessment: Ongoing Monitoring for QCP Effectiveness (Performed by supervisor and/or section head) |  |  |
| --- | --- | --- |
| Reasons for QC failures, PT failures, and patient isolate reporting errors will be examined and addressed as needed in a new/updated risk assessment: 1) Has a new risk factor been identified? 2) Does this change the frequency of risk? 3) Does the risk factor change the potential severity of harm to patient? |  |  |
| Daily review of patient results for reporting errors and clinician complaints. Take corrective action and revise QCP as needed. |  |  |
| Monthly or weekly review of QC results by supervisor or section head. Take corrective action and revise QCP when unexpected QC failures indicate adjustment to the QC plan defined herein is needed. |  |  |
| Periodic review of length of time from specimen collection to AST result reporting to determine incidence of reports delayed beyond 5 days. Take corrective action and revise QCP when number of delayed reports exceeds acceptable limit as established by the laboratory director. |  |  |
| Regular review of Proficiency Testing results. Take corrective action and revise QCP if necessary when PT results are not acceptable. |  |  |
| Monthly review of all equipment maintenance/monitoring logs according to standard laboratory protocols. Take corrective action and revise QCP as needed. |  |  |
| Regular training and competency assessment according to standard laboratory protocols. Modify training and revise QCP as needed. |  |  |
| Continual participation in this institution’s quality program that addresses specimen handling and erroneous specimen labeling. Take corrective action and revise QCP as needed. |  |  |
| This QCP has been reviewed and is approved by the laboratory director (as named on the CLIA license). | Signature | Date |