IQCP for Streamlined QC of Commercial Identification System Vitek 2

Facility: New Vision Medical Laboratory 750 W. High St Lima, Ohio 45801

Test System: Commercial Identification Vitek 2

Test System Primary SOPs include: nvml.micro.vitek.012.NH nvml.micro.vitek.006.GN nvml.micro.vitek.007.GP nvml.micro.vitek.008.YST nvml.micro.vitek.009.ANC

Historical Quality Review: CLIA ’88 regulations require testing QC strains on each new lot/shipment of identification panels (that utilize two or more substrates) to verify positive and negative reactivity for each substrate and/or reagent. In 2008, CMS amended the regulations referencing the CLSI M50-A document, allowing laboratories to use quality control guidelines for reduced or “streamlined” QC as described in CLSI document M50-A “Quality Control for Commercial Microbial Identification Systems”, 2008.1 A streamlined QC scheme is permitted as long as the manufacturer has designated “key indicator strains” in their product material. The 2008 CMS announcement letter Ref: S&C-09-06 may be found on the CMS website as indicated below under References.2 Since recent CMS regulations have removed all references to CLSI documents, including the M50A, an IQCP is required to continue streamlined QC of commercial identification system Vitek 2. This laboratory has been following CLSI M50-A and manufacturer’s streamlined QC recommendations for over 6 years without any significant identification QC problems. Any problem related to commercial identification performance has involved: - random and infrequent unsatisfactory QC of identification substrate(s) which corrected upon repeat testing Processes to mitigate patient reporting errors based on use of unacceptable identification substrates are addressed in this IQCP.

Information Used to Conduct Risk Assessment

Regulatory and Accreditation Requirements:

Checklist from Accrediting Agency: CAP: MIC 21626

Method verification: Instrument received and test system verification completed in year 2011. Streamlined microbial identification was verified in year 2008 through 2011(documented QC performance for 3 consecutive lot numbers of gram positive and gram negative panels from 3 different shipments that spanned 3 consecutive seasons). Documentation filed in Microbiology cabinet.

Training of personnel: Completion of training documented in CAP Competency Program.

Competency Assessment: New employees 6 months after initial training and annually thereafter. Documentation filed in CAP Competency Program.

Proficiency Testing: Rotate personnel; all personnel review results. Proficiency testing records filed in CAP Survey Binder and CAP Competency Program.

Quality Control: CLIA ’88 regulations require testing of QC strains on each new lot/shipment of commercial identification systems (utilizing two or more substrates) to verify positive and negative reactivity for each substrate and/or reagent. NVML has developed an IQCP for streamlined QC of microbial identification.

Test System Information:

Manufacturer: - The Vitek ID cards have no package inserts. Pertinent information for their instructions and use are found in the Vitek 2 Technology paper found on line. A copy of this paper is stored in the Microbiology cabinet. This paper contains system performance data and describes testing principle and procedure, QC recommendations (including streamlined QC), and limitations. Manufacturer’s alerts and bulletins are located in V2 Binder. - Operator’s manual including troubleshooting guide is located on bench.

References used during collection of information for RA: 1 CLSI: Quality Control for Commercial Microbial Identification Systems; Approved Guideline. Document M50-A. CLSI, 940 West Valley Road, Suite 1400, Wayne, Pennsylvania 19087-1898 USA. 2008. 2 CMS letter, Streamlined QC Ref: S&C-09-06. 2008 http://www.cms.gov/Medicare/Provider-Enrollment-and-Certification/SurveyCertificationGenInfo/downloads/SCLetter09-06.pdf

Summary of in-house data from streamlined QC testing: Streamlined QC testing was performed according to: nvml.micro.vitek.012.NH nvml.micro.vitek.006.GN nvml.micro.vitek.007.GP nvml.micro.vitek.008.YST nvml.micro.vitek.009.ANC Review of QC records for the past 12 months that contained approximately 200 results demonstrated: - 0 % occurrence of random QC errors.

Summary of in-house data from routine instrument performance checks: Instrument checks were done according to nvml.micro.vitek.012.NH nvml.micro.vitek.006.GN nvml.micro.vitek.007.GP nvml.micro.vitek.008.YST nvml.micro.vitek.009.ANC Review of instrument QC records for the past 12 months that contained approximately 32 routine checks of the Vitek 2 instrument revealed no instrument performance problems that would impact patient results.

Summary of corrected reports and physician complaints: Documentation located PI Variance Report on line. Review of corrected reports for the last 12 months indicated: 0 corrected reports with erroneous identification were performed on the Vitek 2 system. There were no physician complaints as a result of unsatisfactory microbial identification.

Risk Assessment and Determination of Risk Level

| Frequency of occurrence: Unlikely (once every 2-3 years) Occasional (once per year) Probable (once per month) Frequent (once a week) | Severity of harm to patient: Negligible (temporary discomfort) Minor (temporary injury; not requiring medical intervention) Serious (impairment requiring medical intervention) Critical (life threatening consequences) |  |
| --- | --- | --- |
| Risk Level: |  |  |
| Risk level for any Risk Factor that is “Not Acceptable” must be addressed in the IQCP. Risk level for any Risk Factor that is “Acceptable” may be included in the IQCP at the discretion of the Laboratory Director. |  |  |

Risk Acceptability Matrix

| Freq./Severity of Harm | Negligible | Minor | Serious | Critical |
| --- | --- | --- | --- | --- |
| Frequent | Not Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Probable | Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Occasional | Acceptable | Acceptable | Acceptable | Not Acceptable |
| Unlikely | Acceptable | Acceptable | Acceptable | Acceptable |

Risk Acceptability Assignment

| Risk Factor (Possible Sources of Error) | Frequency of occurrence | Severity of harm to patient | Risk Level |
| --- | --- | --- | --- |
| Preanalytical |  |  |  |
| Specimen (Primary): |  |  |  |
| Patient identification | probable | minor | Not Acceptable |
| Collection/container/volume | frequent | negligible | Not Acceptable |
| Integrity | frequent | negligible | Not Acceptable |
| Transport | frequent | negligible | Not Acceptable |
| Storage | probable | negligible | Acceptable |
| Specimen (Organism): |  |  |  |
| Clinically relevant | occasional | negligible | Acceptable |
| Colony age/viability/sampling | occasional | minor | Acceptable |
| Media type | unlikely | minor | Acceptable |
| Pure isolate | unlikely | minor | Acceptable |
| Inoculum suspension preparation | occasional | minor | Acceptable |
| Analytical |  |  |  |
| Testing Personnel: |  |  |  |
| Training | probable | serious | Not Acceptable |
| Competency | probable | serious | Not Acceptable |
| Experience | probable | serious | Not Acceptable |
| Proficiency Testing | unlikely | negligible | Acceptable |
| Staffing | occasional | minor | Acceptable |
| Reagents: |  |  |  |
| Shipping/receiving/storage | occasional | minor | Acceptable |
| Expiration dates | unlikely | minor | Acceptable |
| Preparation/use | occasional | minor | Acceptable |
| QC strain storage/prep | occasional | negligible | Acceptable |
| Environment: |  |  |  |
| Temperature/airflow/humidity/ ventilation | unlikely | negligible | Acceptable |
| Utilities | occasional | minor | Acceptable |
| Space | unlikely | negligible | Acceptable |
| Noise/vibration | unlikely | negligible | Acceptable |
| Test System: |  |  |  |
| Mechanical/electronic stability of instrument/equipment/jam | occasional | negligible | Acceptable |
| Software | unlikely | negligible | Acceptable |
| Transmission of results to LIS | unlikely | serious | Acceptable |
| Postanalytical |  |  |  |
| Test Results: |  |  |  |
| Results reported within 5 days | occasional | minor | Acceptable |
| Transmission of results to Electronic Health Record | occasional | serious | Acceptable |
| Review reported results | frequent | serious | Not Acceptable |
| Clinician feedback | unlikely | serious | Acceptable |

Risk Assessment

| Possible Sources of Error |  | How can identified sources of error be reduced? |
| --- | --- | --- |
| Risk Factor | Possible Error |  |
| Preanalytical |  |  |
| 1A: Specimen - Biological | - Improper specimen procurement/ handling/processing | - Adhere to procedures nvml.037 Labeling and Handling of Laboratory Specimens, nvml.054 Outpatient Test Request, nvml.058 Verification of Patient Identity, nvml.073 Specimen Rejection Criteria and nvml.micro.plating.005.Rejection of Specimens that addresses patient identification and specimen collection, labeling, transport, storage and remedial actions to control improperly handled specimens or delayed specimens. - Review representative specimen processing errors with all staff involved with patient specimens. During initial training and competency assessment, emphasize: - Proper specimen handling/processing is the most critical part of any test - Failure to inoculate/streak correctly (no isolated colonies) and delayed incubation may result in delayed microbiology reports |
| Patient/specimen identification |  | See above (Specimen) |
| Collection/container/ volume |  | See above (Specimen) |
| Integrity |  | See above (Specimen) |
| Transport |  | See above (Specimen) |
| Storage |  | See above (Specimen) |
| 1B: Specimen - Organism |  |  |
| Clinically relevant | - Clinically irrelevant organisms tested | During initial training and competency assessment, emphasize differentiation of organisms included in normal flora group vs. potential pathogens requiring identification |
| Colony age/viability/sampling | - Organism non-viable (or beyond maximum age for identification testing as described by the manufacturer) | During initial training and competency assessment, emphasize: - Lengths of time various organisms generally remain viable on various media - Proper age of colony for identification testing |
| Media | - Media for inoculum source other than that recommended is used | During initial training and competency assessment, emphasize: - Appropriate media for inoculum preparation - Species that can be reliably identified by the test system based on manufacturer’s recommendations |
| Pure isolate | - Mixed inoculum | During initial training and competency assessment, emphasize: - Proper selection of colonies for inoculum preparation - Hazards of picking from plates not incubated for sufficient time (“young” colonies) or poorly isolated colonies - Potential sources of contamination during testing process - Use of purity plates - Impact of delayed results if retesting needed |
| Inoculum suspension preparation | - Over- or under-inoculation - Use of nonviable colonies - Inoculum suspension sets too long before inoculating the identification panel | During initial training and competency assessment, emphasize: - Proper use of turbidity meter for inoculum standardization - Proper inoculum suspension preparation |
| Analytical |  |  |
| 2: Testing Personnel | - Incompletely trained - Unaware of updated protocols | During initial training and competency assessment, emphasize: - Key aspects of microbial identification - Supervisor annual review of any changes in microbial identification recommendations by accrediting agencies or manufacturer |
| Training |  | See above (Testing Personnel) |
| Competency |  | See above (Testing Personnel) |
| Experience |  | See above (Testing Personnel) |
| Proficiency Testing |  | - All staff read (and sign off) on PT sample critiques |
| Staffing | Inadequate to perform testing without errors | - Supervisor to annually review appropriate staffing needs for identification testing and schedule staff accordingly |
| 3: Reagents |  | During initial training and competency assessment, emphasize standard rules to always: - Take responsibility for reagents/supplies (all staff) - Maintain reagents at proper storage conditions - Check expiration dates - Perform required QC |
| Receiving/storage | - Incorrect ordering - Depleted reagent supply - Reagent integrity compromised | - Designated staff member(s) assigned to inventory (order/receipt) microbial identification materials to ensure supply is properly maintained and testing materials are handled appropriately on receipt |
| Expiration dates |  | See above (Reagents) |
| Preparation/use | - Use incorrect panel for select organism | - Use color codes on boxes of panels |
| QC strain storage/prep | - QC out of control due to improper QC strain maintenance | During initial training and competency assessment, emphasize: - Proper maintenance of QC strains (limited number of subcultures) - Potential sources of QC failures - QC troubleshooting - QC frequency Role of QC strains versus other QA measures to ensure reliable reporting of patient results |
| 4: Environment | - Results not reported (ancillary equipment failure, e.g., incubator malfunction) | During initial training and competency assessment, emphasize standard rules for: - Take responsibility for any possible instrument/ environmental problem (out of the ordinary observation)(all staff) - Equipment maintenance - Temperature recording (done automatically with continuous monitoring device) - Electrical supply |
| Temperature/airflow/humidity/ ventilation |  | See above (Environment) |
| Utilities |  | See above (Environment) |
| Space |  | N/A (sufficient space available) |
| Noise/vibration |  | See above (Environment) |
| 5: Test System |  | During initial training and competency assessment, emphasize standard rules for: - Take responsibility for any possible instrument/test system problem (out of the ordinary observation) (staff trained on test system) |
| Mechanical/electronic/jam | Results not reported (e.g., instrument malfunction and/or aborted test) | - Perform preventive maintenance according to recommended schedule During initial training and competency assessment, emphasize: - How to avoid jams |
| Software | - Erroneous results reported | - Daily supervisor (or supervisor designee) review of reported results - Software flags unusual results requiring supervisor review (e.g. organism identification is not consistent with associated AST results) During initial training and competency assessment, emphasize: - Results requiring follow up action (e.g., confirmation by repeat testing) - Atypical results requiring consultation with supervisor |
| Transmission of results to LIS | - Incorrect transmission of results - Delay in transmission of results | - Daily supervisor (or supervisor designee) review of reported results - Annual check of test system- LIS computer interface |
| Postanalytical |  |  |
| 6: Test Results |  | - Supervisor maintains summary of incorrect results released and reviews this summary . - During initial training and competency assessment, emphasize timely reporting of both preliminary results and final reports |
| Results reported within 5 days | - Results delayed beyond that expected for organism type | See above (Test Results) |
| Transmission of results to Electronic Health Record | - Incorrect transmission of results - Delay in transmission of results | See above (Test Results) |
| Review reported results | - Erroneous results reported | See above (Test Results) |
| Clinician feedback | - Complaints/suggestions regarding delayed or potential erroneous results | See above (Test Results) - Incorporate suggestions into QA plan, as appropriate. |

Final QCP for Streamlined QC of Commercial Identification System XYZ

Based on our risk assessment and Quality Assessment, the QCP consists of following the instructions that are provided in nvml.micro.vitek.012.NH nvml.micro.vitek.006.GN nvml.micro.vitek.007.GP nvml.micro.vitek.008.YST nvml.micro.vitek.009.ANC

Testing of appropriate QC strains for streamlined QC on each new lot/shipment of panels before or concurrently with placing these materials into use for testing patient’s isolates.

Testing of appropriate QC strains on each panel type when appropriate after major system maintenance or software upgrade

Recording and evaluating QC results according to QC acceptability criteria as defined in nvml.micro.vitek.012.NH nvml.micro.vitek.006.GN nvml.micro.vitek.007.GP nvml.micro.vitek.008.YST nvml.micro.vitek.009.ANC Any abnormal result is immediately investigated.

| Quality Assessment: Ongoing Monitoring for QCP Effectiveness (Performed by supervisor and/or section head) |  |  |
| --- | --- | --- |
| Reasons for QC failures, PT failures, and patient isolate reporting errors will be examined and addressed as needed in a new/updated risk assessment: 1) Has a new risk factor been identified? 2) Does this change the frequency of risk? 3) Does the risk factor change the potential severity of harm to patient? |  |  |
| Daily review of patient results for reporting errors and clinician complaints. Take corrective action and revise QCP as needed. |  |  |
| Weekly or monthly review of QC results by supervisor or section head. Take corrective action and revise QCP when unexpected QC failures indicate adjustment to the QC plan defined herein is needed. |  |  |
| Periodic review of length of time from specimen collection to microbial identification result reporting to determine incidence of reports delayed beyond 5 days. Take corrective action and revise QCP when number of delayed reports exceeds acceptable limit. |  |  |
| Regular review of Proficiency Testing results. Take corrective action and revise QCP if necessary when PT results are not acceptable. |  |  |
| Monthly review of all equipment maintenance/monitoring logs according to standard laboratory protocols. Take corrective action and revise QCP as needed. |  |  |
| Regular training and competency assessment according to standard laboratory protocols. Modify training and revise QCP as needed. |  |  |
| Continual participation in this institution’s quality program that addresses specimen handling and erroneous specimen labeling. Take corrective action and revise QCP as needed. |  |  |
| This QCP has been reviewed and is approved by the laboratory director (as named on the CLIA license). | Signature | Date |