New Vision Medical Laboratories

Microbiology IQCP Disk Diffusion and Etest Susceptibility

Facility: New Vision Medical Laboratories, 750 West High St, Lima, OH

Test System: Disk Diffusion and Etest Antimicrobial Susceptibility Testing

Test System Primary SOPs include: nvml.micro.vitek.003.KBdisk nvml.micro.vitek.002.E-Test

Historical Quality Review: CLIA ’88 requires testing of QC strains daily (or each day patient’s tests are performed) for AST. Previously, CLIA inspectors’ guidelines recognized use of CLSI standards M100 and M02 which indicate weekly testing of QC strains is acceptable following documentation of satisfactory daily QC testing. This laboratory has been following the CLSI standards for over 25 years without any significant QC problems. It is rare to encounter an out of range result with a QC strain that indicates a test system problem. Nearly all testing errors or delays in reporting occur with individual patient isolates and these errors are unrelated to testing QC strains or a problem with testing reagents. Processes to mitigate patient reporting errors and delayed reports are addressed in this IQCP.

Information Used to Conduct Risk Assessment

Regulatory and Accreditation Requirements:

Checklist from Accreditation Agency: CAP: MIC 21820 MIC 21840 MIC 21910 MIC 21930 MIC 21940 MIC 21943 MIC 21944 MIC 21946 MIC 21950

Method verification: Test system was deemed acceptable based on performance and completion of 20 day QC for the following antibiotics: Vancomycin Etest (2007), Pen Etest and Cefuroxime KB disk (2014). Weekly QC now routinely performed on these designated antibiotics.

Training of personnel: Completion of training documented in CAP Competency Assessment

Competency Assessment: New employees 6 months after initial training and annually thereafter. Documentation filed in CAP Competency Assessment.

Proficiency Testing: Rotate personnel; all personnel review results. Proficiency testing records filed in CAP yearly Proficiency Surveys and yearly competencies found on CAP Competency Program.

Quality Control: CLIA ’88 and Accrediting Agency require testing of QC strains daily (or each day patient’s tests are performed) for AST. Alternatively, NVML performed 20 day QC for cefuroxime, vancomycin and penicillin AST. The indicated QC strains by CLSI M-100 documents and control ranges were appropriately applied. Upon successful completion of 20 day QC, microbiology converted to weekly QC testing on these ASTs. All verifications were deemed acceptable by the proper designee. This documentation is filed in the Microbiology cabinet.

Test System Information:

Disk and E test Manufacturer Package insert contains testing principle and procedure, QC recommendations, and limitations. Package inserts are located in Microbiology cabinet. Manufacturer alerts and bulletins are located in the Microbiology cabinet and the Vitek 2 Binder.

AST Manufacturer: Package inserts for all disk diffusion susceptibility testing media contain testing principle and procedure, QC recommendations, and limitations. Package inserts are located in the Microbiology cabinet. Manufacturer alerts and bulletins are located in KB AST Binder.

Scientific publications used during collection of information for RA: CLSI document M02-A12. 2015.

Summary of in-house data from routine testing of QC strains: QC testing was performed according to procedures nvml.micro.vitek.003.KBdisk and nvml.micro.vitek.002.E-Test. Review of QC records for the past year contained approximately 160 results demonstrated: - 0.0% occurrence of random QC errors. - 0.00% occurrence of system QC errors.

Summary of corrected reports and physician complaints: Documentation located Quality Assurance Folder on the G drive under PI variances Binder. Review of reporting errors identified prior to report release, corrected reports and physician complaints and significantly delayed reports (> 5 days after specimen collection) for the past 12 months revealed: - 2 corrected reports due to one or more of the following: 1) erroneous interpretation due to application of inappropriate interpretive criteria (n=2) - 0 formal physician complaints. - 0 reports were not finalized within 5 days of specimen collection. Note: during this review of corrected reports and physician complaints, none of the errors could have been avoided by any changes in protocol for testing of QC strains including frequency of testing QC strains.

Risk Assessment and Determination of Risk Level

| Frequency of occurrence: Unlikely (once every 2-3 years) Occasional (once per year) Probable (once per month) Frequent (once a week) | Severity of harm to patient: Negligible (temporary discomfort) Minor (temporary injury; not requiring medical intervention) Serious (impairment requiring medical intervention) Critical (life threatening consequences) |
| --- | --- |
| Risk Level: |  |
| Risk level for any Risk Factor that is “Not Acceptable” must be addressed in the IQCP. Risk level for any Risk Factor that is “Acceptable” may be included in the IQCP at the discretion of the Laboratory Director. Note: Patient response plays a significant role in addition to AST results in guiding antimicrobial therapy and provides a limited safeguard for preventing harm in patients for which erroneous AST results are reported or results are delayed. |  |

Risk Acceptability Matrix

| Probability of Harm | Negligible | Minor | Serious | Critical |
| --- | --- | --- | --- | --- |
| Frequent | Not Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Probable | Acceptable | Not Acceptable | Not Acceptable | Not Acceptable |
| Occasional | Acceptable | Acceptable | Acceptable | Not Acceptable |
| Unlikely | Acceptable | Acceptable | Acceptable | Acceptable |

Risk Acceptability Assignment

| Risk Factor (Possible Sources of Error) | Frequency of occurrence | Severity of harm to patient | Risk Level |
| --- | --- | --- | --- |
| Preanalytical |  |  |  |
| Specimen (Primary): |  |  |  |
| Patient identification | probable | minor | Not Acceptable |
| Collection/container/volume | frequent | negligible | Not Acceptable |
| Integrity | frequent | negligible | Not Acceptable |
| Transport | frequent | negligible | Not Acceptable |
| Storage | probable | negligible | Acceptable |
| Specimen (Organism): |  |  |  |
| Clinically relevant | probable | minor | Not Acceptable |
| Colony age/viability/sampling | frequent | minor | Not Acceptable |
| Media type | unlikely | minor | Acceptable |
| Pure isolate | frequent | serious | Not Acceptable |
| Inoculum suspension preparation | occasional | minor | Acceptable |
| Risk Factor | Frequency of occurrence | Severity of harm to patient (maximum) | Risk Level |
| Analytical |  |  |  |
| Testing Personnel: |  |  |  |
| Training | probable | serious | Not Acceptable |
| Competency | probable | serious | Not Acceptable |
| Experience | probable | serious | Not Acceptable |
| Proficiency Testing | unlikely | negligible | Acceptable |
| Staffing | occasional | minor | Acceptable |
| Reagents: |  |  |  |
| Shipping/receiving/storage and media sterility check | occasional | minor | Acceptable |
| Expiration dates | occasional | minor | Acceptable |
| Preparation/use | probable | minor | Not Acceptable |
| QC strain storage/prep | occasional | negligible | Acceptable |
| Environment: |  |  |  |
| Incubator temperature/ humidity/atmosphere | occasional | minor | Acceptable |
| Utilities | occasional | minor | Acceptable |
| Test System: |  |  |  |
| Inoculation of agar plate | occasional | serious | Acceptable |
| Application of disks | occasional | serious | Acceptable |
| Incubation time/atmosphere | occasional | serious | Acceptable |
| Measurement of zones of inhibition | probable | serious | Not Acceptable |
| Interpretation of zone sizes | probable | serious | Not Acceptable |
| Antimicrobial reporting rules | probable | serious | Not Acceptable |
| Transcription errors during LIS entry | occasional | critical | Not Acceptable |
| Risk Factor | Frequency of occurrence | Severity of harm to patient (maximum) | Risk Level |
| Post Analytical |  |  |  |
| Test Results: |  |  |  |
| Results reported within 5 days | probable | serious | Not Acceptable |
| Transmission of results to Electronic Health Record | occasional | serious | Acceptable |
| Review reported results | frequent | serious | Not Acceptable |
| Clinician feedback | probable | serious | Not Acceptable |

Risk Assessment

| Possible Sources of Error |  | How can identified sources of error be reduced? |
| --- | --- | --- |
| Risk Factor | Possible Error |  |
| Preanalytical |  |  |
| 1A: Specimen - Biological | - Improper specimen procurement/ handling/processing | - Adhere to procedures in: nvml.037 Labeling and Handling of Laboratory Specimens, nvml.054 Outpatient Test Request, nvml.058 Verification of Patient Identity, nvml.073 Specimen Rejection Criteria and nvml.micro.plating.005.Rejection of Specimens that addresses patient identification and specimen collection, labeling, transport, storage and remedial actions to control improperly handled specimens or delayed specimens. - Annually review representative specimen processing errors with all staff involved with patient specimens. During initial training and competency assessment, emphasize: - Proper specimen handling/processing is the most critical part of any test - Failure to streak correctly (no isolated colonies) and delayed incubation may result in delayed AST reports |
| Patient/specimen identification |  | See above (Specimen) |
| Collection/container/ volume |  | See above (Specimen) |
| Integrity |  | See above (Specimen) |
| Transport |  | See above (Specimen) |
| Storage |  | See above (Specimen) |
| 1B: Specimen - Organism |  |  |
| Clinically relevant | - Clinically irrelevant organisms tested - Additional species may be significant in select patient types (e.g., immunosuppressed) - Physicians may request testing of isolates that are not clinically relevant; requests may be inappropriate and results misleading | Procedures: nvml.micro.vitek.003.KBdisk and nvml.micro.vitek.002.E-Test describes selecting organisms to test for AST based on organism ID, specimen source and quantity Physicians can request additional testing in select patients; comment added is added to media folder indicating physician/healthcare provider initiating special request. Supervisor/director may discuss with requesting healthcare provider those requests that may be inappropriate. |
| Old or less viable | - Colonies on source plate > 1 day old | During initial training and competency assessment, emphasize: - Organism growth requirements (especially S. pneumoniae) |
| Pure isolate | - Mixed inoculum or contaminated AST plate | - Solicit regular feedback on streaking of primary plates (for isolated colonies) - Daily review of AST profiles for aberrant results possibly due to mix/contamination During initial training and competency assessment, emphasize: - Proper organism selection for inoculum preparation - Risks of selecting “young” colonies or poorly isolated colonies - Potential sources of contamination during testing process - Check for contaminating colonies growing inside the zone (s) of inhibition or superimposed on lawn of growth. - Impact of delayed results (if retesting needed) |
| Media Type | - Media for inoculum source other than that recommended is used - Disk diffusion AST plate fails to support growth of the test organism | During initial training and competency assessment, emphasize: - Appropriate media for inoculum - Species that can be reliably tested by disk diffusion test system based on CLSI recommendations |
| Inoculum suspension | - Overinoculation or underinoculation - Use of nonviable colonies | During initial training and competency assessment, emphasize: - Proper inoculum suspension preparation - Use of 0.5 McFarland turbidity standard (or photometric device) for inoculum standardization - Impact of overinoculation (false R) or underinoculation (false S) |
| Species appropriate | - Testing of species not indicated for test system | During initial training and competency assessment, emphasize: - Species that can be reliably tested by test system based on CLSI recommendations |
| Analytical |  |  |
| 2: Testing Personnel | - Incompletely trained - Unaware of updated recommendations for AST/reporting | During initial training and competency assessment, emphasize: - Key aspects of disk diffusion AST to include those described in this IQCP - Supervisor annually review any changes in AST recommendations described by accrediting agencies or standards organizations |
| Training |  | See above (Testing Personnel) |
| Competency |  | See above (Testing Personnel) |
| Experience |  | - Supervisor review AST reports generated by new employees prior to release for the first two months of their employment |
| Proficiency Testing |  | - All staff read (and sign off) on PT sample critiques |
| Staffing | Inadequate to perform testing without errors | - Supervisor to annually review appropriate staffing needs for AST and schedule staff accordingly |
| 3: Reagents |  | During initial training and competency assessment, emphasize standard rules to always: - Take responsibility for reagents/supplies (all staff) - Maintain AST media and disks at proper storage conditions (disks in dessicated environment) - Check expiration dates - Perform required QC |
| Shipping/receiving/storage and AST media sterility check | - Incorrect ordering - Depleted disk or media supply - Disk integrity compromised - Contaminated AST media (sterility check not performed prior to lot/shipment use) | - Designated staff member(s) assigned to inventory (order/receipt) AST disk and media to ensure inventory properly maintained and testing materials are handled appropriately on receipt - Staff member inoculating AST plates examines plates prior to inoculation for contamination |
| Preparation/use | - Disks and/or media removed from refrigerator/freezer and not warmed to room temperature (disk container must not be opened until it reaches room temperature) - AST Media surface wet | During initial training and competency assessment, emphasize: - Remove appropriately stored disks from freezer or refrigerator with sufficient time to equilibrate to room temperature. - Examining plates, and if necessary leave lids ajar to remove excess moisture |
| QC strain storage/prep | - QC out of control due to improper QC strain maintenance | During initial training and competency assessment, emphasize: - Proper maintenance of QC strains (limited number of subcultures) - Potential sources of QC failures - QC troubleshooting - QC frequency - Role of QC strains versus other QA measures to ensure reliable reporting of patient results |
| 4: Environment | - Results not reported (equipment failure, e.g., incubator malfunction) | During initial training and competency assessment, emphasize standard rules for: - Take responsibility for any possible equipment/ environmental problem (out of the ordinary observation)(all staff) - Equipment maintenance - Temperature recording (done automatically with continuous monitoring device) and CO2 incubator atmosphere recording - Electrical supply |
| Incubator temperature/ humidity/atmosphere |  | See above (Environment) |
| Utilities |  | See above (Environment) |
| 5: Test System |  | During initial training and competency assessment, emphasize standard rules for: - Adherence to the procedure in SOP 5.1.9 for correct inoculation of AST media, application of disks, and incubation time and atmosphere - Take responsibility for any possible test system problem (out of the ordinary observation) |
| Inoculation of agar plate | - Improper streaking of plate leading to lawn of growth that is not confluent | - See above (Test System) |
| Application of disks | - Disks relocated after placing on plate - Disks placed too close together so that zones are overlapping | - See above (Test System) |
| Incubation time/atmosphere | - AST plates not incubated under the appropriate atmospheric conditions based on the organism - AST plates not incubated for the recommended time range based on the organism | - See above (Test System) |
| Measurement of zones of inhibition | - When measuring zones of inhibition: - Failure to recognize inadequate growth or contamination - Under interpreting zone sizes may lead to smaller recorded zone sizes and potential false resistance - Over interpreting zone sizes may lead to larger recorded zone sizes and potential false susceptibility | During initial training and competency assessment, emphasize: - Proper measurement of zone sizes using a measuring caliper or metric ruler - Adequate lawn of growth and risk of false susceptible results if inadequate - Recognition of contaminated plate |
| Interpretation of zone sizes | - Incorrect interpretation of zone of inhibition leading to erroneous results | During initial training and competency assessment, emphasize: - Follow M02-A12 tables for correct interpretive criteria based on the organism and zone size |
| Antimicrobial reporting rules | - Inappropriate drugs reported - Report comments missing or inappropriate for the culture - Failure to follow laboratory reporting rules to confirm unusual susceptibility or resistance based on M100-S25 Appendix A and B (see SOP # 5.1.4) | During initial training and competency assessment, emphasize: - LIS reporting rules address (and flag) most, but not all, potential errors - Daily supervisor (or designee) review of all reported results - Laboratory reporting rules based on unusual susceptibility or resistance for an organism/antibiotic combination - Results requiring follow up action (e.g. confirmation by repeat testing) - Results requiring consultation with supervisor/director |
| Transcription errors during LIS entry | - Incorrect zone size values and/or interpretations entered during computer entry | During initial training and competency assessment, emphasize: - Accuracy when entering manual results - Supervisory review of manually entered results prior to LIS release |
| Postanalytical |  |  |
| 6: Test Results |  | - Supervisor maintains summary of incorrect results released and reviews this summary . - QA monitor for time to reporting AST results During initial training and competency assessment, emphasize: - Need for timely results to guide therapy and identify potential multidrug resistant organisms that might require patient isolation - Reporting preliminary results (timely reporting) |
| Results reported within 5 days | - Results delayed beyond that expected for organism type | See above (Test Results) |
| Transmission of results to Electronic Health Record | - Incorrect transmission of results - Delay in transmission of results | See above (Test Results) |
| Review reported results | - Inappropriate drugs reported - Erroneous results reported - Zone sizes interpreted incorrectly - Report comments missing or inappropriate for the culture | See above (Test Results and Test System) |
| Clinician feedback | - Complaints/suggestions regarding delayed results and potential erroneous results | See above (Test Results) - Incorporate suggestions into QA plan, as appropriate. |

Final QCP for Disk Diffusion AST System

Based on our risk assessment and Quality Assessment, the QCP consists of following the instructions that are provided in explicit detail in nvml.micro.vitek.003.KBdisk and nvml.micro.vitek.002.E-Test procedures.

Testing of appropriate ATCC QC strains on each new lot/shipment of disks and appropriate AST media, before or concurrently with placing these materials into use for testing patient’s isolates.

Testing of appropriate ATCC QC strains on each AST media type weekly.

Testing of appropriate ATCC QC strains against any new antimicrobial agent at least 15 times (over a minimum of 5 days) prior to initiation of weekly QC testing (for that antimicrobial agent); accomplished during performance of verification study.

Daily monitoring and recording of freezer, refrigerator and room temperatures. Daily monitoring and recording of CO2 incubator conditions. Any out of range result is investigated and reported to the supervisor.

Recording and evaluating QC results according to QC acceptability criteria as defined in nvml.micro.vitek.003.KBdisk and nvml.micro.vitek.002.E-Test procedures for performance of disk diffusion and Etest AST. Any out of range result is immediately investigated and corrective action performed prior to releasing any patient results.

| Quality Assessment: Ongoing Monitoring for QCP Effectiveness (Performed by supervisor and/or section head) |  |  |
| --- | --- | --- |
| Reasons for QC failures, PT failures, and patient isolate reporting errors will be examined and addressed as needed in a new/updated risk assessment: 1) Has a new risk factor been identified? 2) Does this change the frequency of risk? 3) Does the risk factor change the potential severity of harm to patient? |  |  |
| Daily review of patient results for reporting errors and clinician complaints. Take corrective action and revise QCP as needed. |  |  |
| Weekly or monthly review of QC results by supervisor or section head. Take corrective action and revise QCP when unexpected QC failures indicate adjustment to the QC plan defined herein is needed. |  |  |
| Periodic review of length of time from specimen collection to AST result reporting to determine incidence of reports delayed beyond 5 days. Take corrective action and revise QCP when number of delayed reports exceeds acceptable limit as established by the laboratory director. |  |  |
| Regular review of Proficiency Testing results. Take corrective action and revise QCP if necessary when PT results are not acceptable. |  |  |
| Monthly review of all equipment maintenance/monitoring logs according to standard laboratory protocols. Take corrective action and revise QCP as needed. |  |  |
| Regular training and competency assessment according to standard laboratory protocols. Modify training and revise QCP as needed. |  |  |
| Continual participation in this institution’s quality program that addresses specimen handling and erroneous specimen labeling. Take corrective action and revise QCP as needed. |  |  |
| This QCP has been reviewed and is approved by the laboratory director (as named on the CLIA license). | Signature | Date |