Sure-Vue hCG

1. PRINCIPLE:

Human chorionic gonadotropin (hCG) is a glycoprotein hormone produced by the developing placenta shortly after fertilization. In normal pregnancy, hCG can be detected in both urine and serum as early as 7 to 10 days after conception (1-4). hCG levels continue to rise very rapidly, frequently exceeding 100mIU/mL by the first missed menstrual period (2-4), and peaking in the 100,000-200,000mIU/mL range about 10-12 weeks into pregnancy. The appearance of hCG in both urine and serum soon after conception, and its subsequent rapid rise in concentration during early gestational growth, make it an excellent marker for the early detection of pregnancy.

2. CLINICAL SIGNIFICANCE:

Sure-Vue Serum/Urine hCG-STAT is a visual and rapid test for the qualitative determination of human chorionic gonadotropin (hCG) in serum or urine to aid in the early detection of pregnancy. The detection of hCG in serum and urine is valuable in the presumptive diagnosis of pregnancy. The developing placenta secretes this glycoprotein hormone after fertilization. The hCG hormone doubles approximately every 2.2 days during the first trimester. Detectable levels start at 5 mIU/mL during the first week of gestation and rise to 100,000 mIU/mL at two to three months. A slower rise may be associated with high-risk abortions.

3. SPECIMEN: A. COLLECTION AND PROCESSING:

Urine – A urine specimen must be collected in a clean and dry container. A first morning urine specimen is preferred since it generally contains the highest concentration of hCG; however, urine specimens collected at any time of the day may be used. Urine specimens exhibiting visible precipitates should be centrifuged, filtered, or allowed to settle to obtain a clear specimen for testing. Serum-Blood should be collected aseptically into a clean tube without anticoagulants. Separate the serum from blood as soon as possible to avoid hemolysis. Use clear non-hemolyzed specimens when possible.

B. REJECTION:

|  | Gross hematuria may prevent an accurate reading of test results by masking the positive line. |  |
| --- | --- | --- |
| C. STORAGE AND PRESERVATION: |  |  |

Store as packaged in the sealed pouch at 2-30°C. The test device is stable through the expiration date printed on the sealed pouch. The test device must remain in the sealed pouch until use. DO NOT FREEZE. Do not use beyond the expiration date.

4. REAGENTS, STANDARDS, AND CONTROLS: A. PREPARATION:

Test device containing monoclonal mouse-hCG colored conjugate and hCG antibody coated on a membrane.

B. CONTROL PROCEDURE:

INTERNAL CONTROLS: The appearance of a Control Line in the C region of the device is a positive procedural control. Correct procedural technique, specimen flow and test device performance is confirmed when a colored line appears in the “C” (control) window of the membrane. If the colored line fails to appear in the “C” (control) area, the test result is invalid. A clear background is an internal negative procedural control. The background color should be white to light pink and should not interfere with the reading of the test result. If a more intensely red background color appears, it may interfere with the ability to read the test result, therefore the rest should be repeated. EXTERNAL CONTROLS: Negative and positive controls will be performed when each kit box is opened. Results will be recorded into the laboratory computer system A test device will be used for each control. See procedure below.

5. EQUIPMENT: A. INSTRUMENTS: NA

6. PROCEDURE: A. PERFORMANCE:

|  | - Allow the test device, serum or urine specimen and/or controls to equilibrate to room temperature (15-30°C) prior to testing. - Bring the pouch to room temperature before opening it. Remove the test device from the sealed pouch and use it as soon as possible. - Place the test device on a clean and level surface. Hold the dropper vertically and transfer 3 full drops of serum or urine (approx. 100μL) to the specimen well of the test device, and then start the timer. Avoid trapping air bubbles in the specimen well. - Wait for the red line(s) to appear. Read the result at 3 minutes when testing a urine specimen, or at 5 minutes when testing a serum specimen. It is important that the background is clear before the result is read. Note: A low hCG concentration might result in a weak line appearing in the test region (T) after an extended period of time; therefore, do not interpret the result after 10 minutes. |  |
| --- | --- | --- |
| C. INTERPRETATION OF RESULTS: |  |  |

|  | POSITIVE:* Two distinct red lines appear. One line should be in the control region (C) andanother line should be in the test region (T). *NOTE: The intensity of the red color in the test line region (T) will vary depending on theconcentration of hCG present in the specimen. However, neither the quantitative value nor the rate of increase in hCG can be determined by this qualitative test. NEGATIVE: One red line appears in the control region (C). No apparent red or pink line appears in the test region (T). INVALID: Control line fails to appear. Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure. Review the procedure and repeat the test with a new test device. If the problem persists, discontinue using the test kit immediately and call 1-800-637-3717 for Technical Assistance. |  |
| --- | --- | --- |
| B. FAILURE/REMEDIAL ACTION: |  |  |

If QC results are unacceptable (either internal or external); repeat test with new test device. If still unacceptable, contact a technical specialist or call the manufacturer for technical guidance. Do not report patients.

8. EXPECTED RESULTS: A. REPORTABLE RANGE:

Negative results are expected in healthy non-pregnant women and healthy men. Healthy pregnant women have hCG present in their urine and serum specimens. The amount of hCG will vary greatly with gestational age and between individuals. Sure-Vue® Serum/Urine hCG-STAT has a sensitivity of 10mIU/mL in serum and 20mIU/mL in urine and is capable of detecting pregnancy as early as 1 day after the first missed menses.

9. REPORTING RESULTS: A. NORMAL VALUES:

Results will be recorded in the patient chart.

10. PROCEDURAL NOTES: A. BACKUP FOR INOPERABLE SYSTEM: NA

B. REFERAL OF SPECIMENS: NA

C. SUBMISSION/HANDLING OF REFERRAL SPECIMENS: NA

11. LIMITATIONS AND INTERFERRING SUBSTANCES:

1. Very dilute urine specimens, as indicated by a low specific gravity, may not contain

representative levels of hCG. If pregnancy is still suspected, a first morning urine specimen

should be collected 48 hours later and tested.

2. False negative results may occur when the levels of hCG are below the sensitivity level of the test. When pregnancy is still suspected, a first morning serum or urine specimen should be collected 48 hours later and tested.

3. Very low levels of hCG (less than 50mIU/mL) are present in urine and serum specimen

shortly after implantation. However, because a significant number of first trimester

pregnancies terminate for natural reasons (5), a test result that is weakly positive should be

confirmed by retesting with a first morning serum or urine specimen collected 48 hours later.

4. A number of conditions other than pregnancy, including trophoblastic disease and certain

non-trophoblastic neoplasms including testicular tumors, prostate cancer, breast cancer, and

lung cancer, cause elevated levels of hCG (6-7). Therefore, the presence of hCG in serum or

urine specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.

5. As with any assay employing mouse antibodies, the possibility exists for interference by

human anti-mouse antibodies (HAMA) in the specimen. Specimens from patients who have

received preparations of monoclonal antibodies for diagnosis or therapy may contain HAMA.

Such specimens may cause false positive or false negative results.

6. This test provides a presumptive diagnosis for pregnancy. A confirmed pregnancy diagnosis should only be made by a physician after all clinical and laboratory findings have been evaluated.

12. METHOD VALIDATION: October 2017

13. REFERENCES:

Sure-Vue Serum/Urine hCG-STAT product insert, revised 05/57/2014