| New Vision Medical Laboratories Joint Township District Memorial Hospital (St. Marys, Ohio) Technical Procedure |  | Procedure Number: nvml.jtdmh.HEMO-133 POLICY INITIATED: 04/2013 |
| --- | --- | --- |

SUBJECT: BLOOD SMEAR EXAMINATION

SCOPE:

This policy applies to all laboratory testing personnel.

POLICY:

This policy provides guidelines for the testing personnel at JTDMH-New Vision Laboratory. It includes how to make a blood smear, how to perform a blood film scan, and some guidelines to use during a manual WBC differential analysis.

SPECIMEN REQUIREMENTS: 1 drop K3EDTA whole blood

A well-stained slide is necessary for accurate interpretation of cellular morphology.

The best staining results are obtained from freshly made slides that have been prepared within 2 hours of blood collection. Heparinized blood is not acceptable. For additional specimen collection and handling information see:

Sysmex XN-550 Operator Guide

Monitoring Abnormal Results HEMO-117

Slide Review Criteria HEMO-121

Automated Slide Stainer Operation HEMO-206

REAGENTS, SUPPLIES, & EQUIPMENT:

Diff-Safe Blood Dispensers	Microscope

Microscope slides	Immersion Oil

For reagents and supplies and equipment see:

Sysmex XN Operators Guide

Automated Slide Stainer Operation Procedure HEMO-206

Diff-Safe Blood Dispenser HEMO- 229

CONTROLS: N/A

CALIBRATION REQUIREMENTS: N/A

CALCULATIONS: N/A

PROCEDURE:

A. SLIDE PREPARATION AND LABELING:

1. Ensure that no dust, oil, or film is on the slide and that the frosted label side is up.

2. Using the Diff-Safe blood dispensing system, place a small drop of blood l - l l/2 cm. (3/4“) from the unfrosted end of slide (Diff –Safe HEMO-229).

3. Using another slide as a spreader, pull spreader back to touch the blood drop. Allow the entire spot to spread across spreader by capillary action.

4. With the spreader tilted at a 30- 40-degree angle, push forward with evenly distributed pressure toward the frosted end of the glass slide. Do not push blood into the frosted area.

5. Identify slides (i.e. first and last name, LIS number) in the frosted section of the glass slide. May use slide labels for identification.

6. Allow slide to air dry completely.

7. Stain slide using the Automated Slide Stainer using Modified Wright-Giemsa stain (HEMO-206)

B. SMEAR EXAM:

Examination of a blood smear is a multistep process.

1. Perform slide quality review:

Each exam begins with a scan of the slide at 50x or lower to assess the overall quality of the smear. Assessment includes staining quality, feathered-edge quality, presence of fibrin, cell distribution, artifacts, crenation etc. Document review for the Slide Quality Check on the Hematology/Coag-Monthly/Daily logsheet, once daily.

2. Perform a scan:

Using the 50X objective find an area of the smear in which the red cells are evenly distributed and barely touching one another (two or three cells may overlap).

2.1 WBC Evaluation (50x):

a) Count estimation - Scan 8-10 fields. The average number of white blood cells per field is determined. The average number of WBC per high power field (50x) is multiplied by 2000 to get an approximation of the total WBC count/uL (mm3). Compare with the automated WBC count as an additional quality control measure.

b) Evaluate morphology of the white cells. Any suspect WBC should be examined under 100x objective.

c) An estimated differential should be performed to verify WBC distribution. This may require a scan of more than 8-10 fields depending on the WBC count.

2.2 RBC morphology (50x)

a) Performed in conjunction with the instrument RBC indices to describe cells as normal or abnormal in size, shape, and color.

b) See “Guidelines for Grading of Morphological Abnormalities” below.

c) Any suspect RBC should be examined under 100X objective to rule out any inclusion bodies or stippling.

2.3 Platelet evaluation (100x)

a) Platelet count estimation: Under 100X oil objective, in an area where approximately 200-250 RBC are seen, count the platelets in 10 fields.

b) The average number of platelets is multiplied by 20,000 to provide an estimate of the total number of platelets present in the sample.

c) Determine if the count is increased or decreased.

d) Compare with the automated platelet count.

e) If the platelet count is <100,000, then appropriate actions must be taken to verify the platelet count and may include the following:

- Verify that the sample does not contain clots

- Rerun the patient sample

- Check for platelet clumping or fibrin strands on the smear

- Perform a patient history search

- Check the patient’s diagnosis

NOTE: When the platelet count is <100,000 on a Negative CBC on the INITIAL finding, then perform a SCAN to confirm the electronic platelet count and rule out platelet clumping. Subsequent findings do not need scanned.

f) If platelet clumps are observed on the smear, then follow the flowchart below:

Platelet Clumps Observed    on Peripheral Smear

Is the instrument platelet count

LESS than 100,000?

NO

Verify the instrument count, with the reportable canned comment @PLTF - Platelet clumping observed. Platelet count may be falsely decreased in number.

YES

Do not verify platelet count.

Attempt to resolve clumping issue.

Can the patient be recollected?

NO

Verify the instrument count, with the reportable canned comment @PLTU - Platelet clumping observed. Platelet count may be falsely decreased in number. Unable to recollect.

YES

Collect LAV and BLUE (when ruling out EDTA sensitivity). Preferably wrap in heel warmer. Send to lab ASAP.

Analyze both specimens, make slides, and compare.

Did the issue get resolved?

YES

NO

If the BLUE tube is used for reporting, then:

1) the WBC and PLT counts must be multiplied          by 1.1 to correct for the Na Citrate tube                     dilution factor before reporting results.

2) add an internal result comment that the platelet          count was obtained using BLUE tube due to             EDTA sensitivity.

Verify the instrument count.

If this is a new patient with EDTA Sensitivity, then add an Internal Patient Stay comment.

Verify the instrument count, with the reportable canned comment @PLTA - Platelet clumping observed. Platelet count may be falsely decreased in number. Attempt to resolve platelet clumping issue was unsuccessful.

Add an Internal Patient Stay comment, suggesting to draw LAV and BLUE on future CBC draws.

g) Evaluate platelet morphology and size abnormalities.

NOTE: The tech may order a QPATH for pathologist review, if needed.

Guidelines for the Grading of Morphological Abnormalities

(Based on Adult Male Reference Ranges)

| ## MICROCYTOSIS 1+ = MCV 70-79 DNR 2+ = MCV 60-69 DNR 3+ = MCV <60 (ages >5days) OR ≤ 90 (ages ≤5days) | ## MACROCYTOSIS 1+ = MCV 100 –110 DNR 2+ = MCV 111-125 DNR 3+ = MCV >125 (ages >5 days) OR ≥130 (ages ≤5 days) |
| --- | --- |
| ### HYPOCHROMI A Also evaluate MCHC <25 (3+) 1+ =MCH 22-26 DNR 2+ =MCH 18-21 DNR 3+ = MCH <18 | ### POLYCHROMATOPHILIA-100x 1+ = 1–3 / field 2+ = 4-12 / field 3+ = >12 / field |
| ### ANISOCYTOSIS (100x) Vary from normal size 1+ = RDWCV & RDWSD DNR 2+ = RDWCV & RDWSD DNR 3+ = RDWCV >20 and/or RDWSD >65 | ### POIKILOCYTOSIS (100X) Vary from normal shape (Evaluate combination of RDW and MCH) 1+ = 2-6/field DNR 2+ = 7-10/field DNR 3+ = 11-20/field |
| WBC MORPHOLOGY: Variant lymphocytes, toxic granulation, hypersegmentation: 1+ = 10-24 % of total WBC 2+ = 25-50% of total WBC 3+ = >50% of total WBC | RBC MORPHOLOGY (100x): Stippling, elliptocytes, echinocytes, acanthocytes, stomatocytes: 1+ = 5-10/field few 2+ = 11-25/field moderate 3+ = >25/field heavy Schistocytes, sickle cells, spherocytes, teardrops, target cells: 1+ = 0-2/field few 2+ = 3-5/field moderate 3+ = >5/field heavy Smudge cells, Howell-Jolly bodies: Present |

C. PERFORMANCE OF A MANUAL DIFFERENTIAL (100X)

1. Performed as indicated by slide review criteria or when specifically ordered.

2. Manual differential analysis includes slide quality review, film scan and WBC differential count.

3. Performed in the same area of the smear as the platelet estimate, where approximately 200-250 RBC per oil field (100x) are seen.

4. Use the 100X oil immersion objective to count 100 WBC. Each WBC is differentiated based on cell type.

5. Performed using the "battlement" track technique. (See Figure below)

Push Wedge Film

Feather Edge     	 Examination Area           Origin

Lateral Edge

6. For WBC <2,000:

6.1 There is no criterion for counting less than 100 cells on a differential.

6.2 A buffy coat preparation is useful to scan for abnormal cell types however, it is not recommended for differential counting.

6.3 It may be necessary to prepare additional slides to allow counting to occur in the proper part of the slide. (200-250 RBCs/100x field)

7. For blasts seen or suspected OR if an inpatient has a history of blasts during their current hospital stay, do NOT verify the differential results. Pathologist must review before reporting. Refer to policy HEMO-121 (Slide Review Criteria).

8. The experienced technologist can use a 50x oil immersion objective to perform the differential analysis. However, all abnormal findings must be verified under the 100x objective.

9. It is imperative that the LIS is marked as scan or differential perform prior to validation. (See Procedure Notes #2 for explanation.)

PROCEDURE NOTES:

1. Slides are prepared and stained for all samples with a POSITIVE (abnormal) result, abnormal or suspect flag (IP messages) from the hematology instrument. Slide retention is 2 weeks.

2. It is imperative that the LIS is marked as scan or differential perform prior to validation. Failure to do so will result in a decrease/increased differential count of 100 cells. (Marking it as scan when a manual differential has been done will result in a greater than 100 cells differential, because of the IG% and IG# being added to your count. Changing the scan to diff performed after you already did the diff manually will cause the total segs and absolute segs to recalculated and thus edit the report.)

3. Smear Preparation-acceptable slide characteristics:

3.1 Smear should be two-thirds length of slide or l - l l/2 inches in length, but not into the frosted portion of the slide.

3.2 Should not be too thick or thin (red cells should not be on top of each other).

3.3 Should be even edged but feathered and evenly distributed and smear should have a metallic sheen.

3.4 Should have no streaks or holes in smear.

3.5 Sources of error: Wrong angle, drop too large, dirty slide, or slide allowed to dry on a surface that was not level.

4. Staining of blood smears: An optimally stained smear has the following characteristics:

4.1 The red blood cells (RBC) should be pink to salmon in color.

4.2 Nuclei are dark blue to purple.

4.3 Cytoplasmic granules are lilac.

4.4 Cytoplasmic granules of basophils are dark blue to black.

4.5 Cytoplasmic granules of eosinophils are red to orange.

4.6 The area between the cells should be clean and free of precipitated stain.

REFERENCE:

1. Sysmex XN-L Series Information for Use Manuals (North American Edition), Sysmex Corporation, Kobe, Japan.

2. Sysmex XN-L Series Automated Hematology Systems Flagging Interpretation Guide.

3. Davidson and Henry, "Clinical Diagnosis," l5th Edition, PP.l0l-l03.

4. J.H. Carr and B.F. Rodak  CLINICAL HEMATOLOGY ATLAS, W.B.Sanders Co.,1999.

5. NCCLS Document H20-A, Vol. 12 No. 1, 1992.

6. Graig A. Lehmann; Saunders Manual of Clinical Laboratory Science; W.B. Saunders Company, 1998.

Policy Review:  Laura Walterbusch, MT

Hematology Lead Technologist

Date:  5/22/2024

Policy Approved:  Patrick Feasel, MD

Date:  5/31/2024