Returning Blood Products Issued to Infusion Center for Remote Storage

| ##### Purpose: | The purpose of this procedure is to provide instructions for the return of blood products stored in the Infusion Center Remote Storage devices. |
| --- | --- |

| ##### Policy: | - Unused blood products must be returned to the Transfusion Service when the Infusion Center closes at the end of day. - Blood products cannot be stored in the Infusion Center Remote Storage devices overnight. - Units returned from the Infusion Center must be accompanied by the corresponding Storage Log & copy of temperature log for review to determine whether the units are acceptable for reissue. - Units with unacceptable or incomplete storage documentation will be discarded. - Only Transfusion Service trained MLT/CLS may determine whether or not a unit is acceptable for reissue. |
| --- | --- |

| ##### Procedure: ##### IVT | Return all units that have not been transfused prior to close of day. Blood products are NOT to be stored in remote storage devices when clinic is closed. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Store blood products in the appropriate monitored storage device until transportation is available to immediately transport blood to Transfusion Service (Blood Bank). |
| 2. | Remove blood product to be returned from storage device. If:; Then:. RBC or FFP; - Place unit at bottom of IVC Transport cooler with ice bag from morning run on top Note: Bag from morning run should be left in sealed Transport Cooler throughout day.Return Ice bag with units or box at end of each day.. Platelet; - Open platelet box - Remove foam plug and 1st layer of gel insulator packs. - Place unit (s) to be returned on top of remaining gel insulator layer - Place layer of gel insulator packs on top of products - Replace foam plug - Replace the Styrofoam lid - Seal box with packing tape.. |

| 3. | Record time that the blood product is removed from storage device in the “Units returned to Lab Blood Bank” column of the appropriate Storage Log. |
| --- | --- |
| 4. | Make copies of the appropriate Temperature & Blood Storage Logs and send to the Transfusion Service (Blood Bank) with the blood product. |
| 5. | Transport blood product immediately to the Lab Transfusion Service (Blood Bank). |

| ##### Procedure: ##### LAB |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Step | Action |  |  |  |  |
|  |  | 1. | Record the date and time that the unit was returned at the bottom of the transfusion record. |  |  |  |  |
|  |  | 2. | Review the copy of the Infusion Center Blood Refrigerator or Platelet Temperature Log |  |  |  |  |
|  |  |  |  | If: | Then: |  |  |
|  |  |  |  | No Temperature Log submitted | - Call IVC to request copy - Place units on the Quarantine shelf with a note indicating why products are on shelf until copy of log is received. |  |  |
|  |  |  |  | Refrigerator Temperature on log is within 1-6C | - Proceed to step 3. |  |  |
|  |  |  |  | Refrigerator Temperature on log is <1C or >6C | - Record “Not OK”, time and intials at the bottom of the Transfusion Record form. - Check the Not OK box on Infusion Center Blood Storage Log sign and date reviewed line - Update status of unit in LIS as DS, use appropriate RE code - Discard unit in Biohazard Trash |  |  |
|  |  |  |  | Platelet Temperature on log is within 20-24C | - Proceed to step 3. |  |  |
|  |  |  |  | Platelet temperature on log is <20C or >24C | - Record “Not OK”, time and intials at the bottom of the Transfusion Record form. - Check the Not OK box on Infusion Center Blood Storage Log sign and date reviewed line - Update status of unit in LIS as DS, use appropriate RE code - Discard unit in Biohazard Trash |  |  |

| 3. | Perform visual inspection on unit. |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | If: | Then: |  |
|  |  | Pass | - Proceed to step 4. |  |
|  |  | Fail | - Record “Not OK”, time and initials at the bottom of the Transfusion Record form. - Check the Not OK box on Infusion Center Blood Storage Log sign and date reviewed line - Update status of unit in LIS as DS, use appropriate ETC or free text reason for failure - Discard unit in Biohazard Trash |  |
| 4. | Check the unit to determine whether or not the hermetic seals are intact. |  |  |  |
|  |  | If: | Then: |  |
|  |  | The seals are intact | Proceed to step 5 |  |
|  |  | One of the seals is broken | - Record “Not OK”, time and initials at the bottom of the Transfusion Record form. - Check the Not OK box on Infusion Center Blood Storage Log sign and date reviewed line - Update the status of the unit to DS in the LIS, use appropriate ETC or free text reason for failure - Discard unit in red biohazard trash. |  |

Continued on next page

| 5. | Perform a temperature check on the unit. See “Using the TEMPCHECK Model TC-12 Rapid Response Thermometer” procedure. |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | If: | Then: |  |
|  |  | The temperature is within specifications for the type of component | - Unit is suitable for reissue. - Check the OK box on Infusion Center Blood Storage Log. Sign/date reviewed line - Update the status of the unit to INV in the LIS. Enter Temperature of unit in the Temperature field. - Discard the lower portion of the cardboard copy of the transfusion record in the Issued bin. - Remove the unit tag from the unit. - Discard the Transfusion Record form, if applicable, and unit tag in Shred-it trash - If the unit is to remain allocated to patient, reprint unit tag and tag unit - Return unit to the appropriate blood storage device |  |
|  |  | The temperature is outside of the specifications for the type of component | - Unit is NOT suitable for reissue - Check the Not OK box on Infusion Center Blood Storage Log. Sign/date reviewed line - Update status of unit in LIS as DS. Enter RE code and temperature of the unit in the Temperature field - Discard the lower portion of the cardboard copy of the transfusion record in the Issued bin. - Remove the Transfusion Record from the unit. - Discard original and cardboard copies of the Transfusion Record form in Shred-it trash - Discard unit in biohazard trash - Process a replacement unit, if indicated. - Write QAR. |  |

Continued on next page

| 6. | Save completed logs in QC binder for monthly supervisory review. |
| --- | --- |

| ##### Related Documents | - Infusion Center Remote Blood Storage - Using the TEMPCHECK Model TC-12 Rapid Response Thermometer - Blood Status Update |
| --- | --- |