#### Verifying Thermometer Calibration

| ##### Purpose | The purpose of this procedure is to provide instructions on the frequency and method used to ensure that thermometers are accurate. |
| --- | --- |

| ##### Policy | - Certified non-mercury thermometers are purchased whenever possible. - Thermometers will be verified upon receipt and annually for liquid thermometers and monthly for electronic thermometers or at the interval recommended by the manufacturer. - Certificates provided by the manufacturer will be accepted as verification of accuracy upon receipt. - Certified thermometers must be verified after the expiration date of certificate prior to use. - Non-certified thermometers will be verified against a standard thermometer available from the National Standards and Technology (NBS Certified Thermometer) prior to being placed into service. - Thermometers will be verified at a temperature within the temperature range for which they will be used, whenever possible. - The NBS certified “reference” thermometer will have at least 0.1C gradations. - The NBS certified thermometer will be recertified annually or at the interval specified by the manufacturer, for accuracy by an outside source. - Records of thermometer verification are kept in a binder in the General Laboratory. |
| --- | --- |

| ##### Procedure: Calibration |  |
| --- | --- |

| Step | Action |  |  |  |
| --- | --- | --- | --- | --- |
| 1. | Locate the Thermometer Inventory Log sheet in the Thermometer Binder located in Gen Lab. |  |  |  |
| 2. | Retrieve all thermometers to be tested from their respective locations. |  |  |  |
| 3. | Examine thermometer for column separation. |  |  |  |
|  |  | If: | Then: |  |
|  |  | Non separation in column | Proceed to step 5 |  |
|  |  | Column is separated | Heat or cool thermometer to reunite column. |  |
| 4. |  |  |  |  |
|  |  | If: | Then: |  |
|  |  | Column is reunited | Proceed to step 5 |  |
|  |  | Column still separated | Remove thermometer from service and notify a supervisor |  |

| 5. | Using the chart below, select the appropriate temperature range of the solution to be used for testing. |  |  |  |
| --- | --- | --- | --- | --- |
|  |  | Refrigerator | 1-6C range |  |
|  |  | Freezer | -20 to 0C |  |
|  |  | Platelet Incubator | 20-24C |  |
|  |  | Plasma Thawer, Heat blocks | 35-37C |  |
| 6. | Prepare a water solution at the desired temperature in a container that will accommodate the NBS thermometer and the thermometer to be tested. |  |  |  |
| 7. | Place the NBS thermometer and the thermometer being tested in the solution until equilibrium is established. (Minimum of 5 minutes) |  |  |  |
| 8. | Record the serial number of the thermometer being tested, date and temperatures of the NBS thermometer and the thermometer being tested on the Thermometer Calibration Log sheet. |  |  |  |
| 9. | Calculate the difference between the thermometer being tested and the NBS thermometer. |  |  |  |
| 10. | Compare the difference to the tolerance limits listed at the top of the worksheet. |  |  |  |
|  |  | If: | Then: |  |
|  |  | Within tolerance level | Proceed to step 12 |  |
|  |  | Outside of tolerance level | Repeat steps 6-9 then proceed to step 11. |  |
| 11. |  |  |  |  |
|  |  | If: | Then: |  |
|  |  | Within tolerance level | Proceed to step 12 |  |
|  |  | Outside of tolerance level | Remove thermometer from service and notify a supervisor |  |
| 12. | Write thermometer ID, date & initials on small label. |  |  |  |
| 13. | Affix label to top end of thermometer so that the temperature range needed is not obscured. |  |  |  |
| 14. | File Thermometer Verification log sheet in Thermometer binder and return binder to storage shelf. |  |  |  |

| ##### Attachment | - Thermometer Verification Log |
| --- | --- |