#### Disconnecting and Reconnecting IPU to Host Computer on the Sysmex XN Hematology Analyzer

|  |  |  |
| --- | --- | --- |
| Background | This procedure describes how to disconnect and reconnect the IPU from the Host Computer during LIS downtime to ensure continuous operation of the Sysmex analyzer |  |

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Procedure A | Follow the steps below to disconnect IPU from Host Computer   |  |  | | --- | --- | | Step | Action | | 1 | On IPU Main menu, select Log off icon | | 2 | Click Yes when prompted to log off | | 3 | Enter username info: admin | | 4 | Enter password info: m116m | | 5 | Click OK | | 6 | On IPU Main Menu, click IPU setting icon | | 7 | In IPU setting display screen , under Connect option, select Host Computer | | 8 | In Host computer connection setting display screen, click “Host Computer Connection” box to remove check mark | | 9 | Click OK to save changes and disconnect IPU from Host Computer  **Note: Host computer status indicator light turns green to gray** | |

Continued on next page

Disconnecting and Reconnecting IPU to Host Computer on the Sysmex XN Hematology Analyzer, Continued

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Procedure B | Follow the steps below to reconnect IPU to Host Computer   |  |  | | --- | --- | | Step | Action | | 1 | On IPU Main menu, select Log off icon | | 2 | Click Yes when prompted to log off | | 3 | Enter username info: admin | | 4 | Enter password info: m116m | | 5 | Click OK | | 6 | On IPU Main Menu, click IPU setting icon | | 7 | In IPU setting display screen, under Connect option, select Host Computer | | 8 | In Host computer connection setting display screen, click empty “Host Computer Connection” box to insert check mark | | 9 | Click OK to save changes and reconnect IPU to Host Computer  **Note: Host computer status indicator light turns gray to green** | |

|  |  |
| --- | --- |
| References | * Sysmex XN series Operator’s manual, March 2017 |

*End*