#### Locating, Reprinting, and Retransmitting sample data on the Sysmex XN Hematology analyzer

|  |  |
| --- | --- |
| Background | This procedure describes how to locate sample data in the Sysmex XN IPU, reprint to the graphic printer or retransmit to the Host Computer |

|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Procedure A | Follow the steps below to locate sample data:   |  |  |  |  |  | | --- | --- | --- | --- | --- | | Step | Action | | | | | 1 | On IPU Main Menu, click Sample Explorer button | | | | | 2 | On the tool bar, click Find button | | | | | 3 | When “Find” screen displays, in sample number field, enter the sample number | | | | | 4 | Click Next or Previous button to locate sample data | | | | | 5 | Once blue highlight bar is on sample data, click close button | | | | |  |  | | | | |  | **If** | **Then** |  | | Retransmit results to Host Computer | * Sample data selected using blue highlight bar * At the tool bar, click Output * Click Host computer | | Reprint results to graphic printer | * Sample data selected using blue highlight bar * At the tool bar, click Output * Click report (GP) | |  | | | | |

|  |  |
| --- | --- |
| Reference | * Sysmex XN series Operator’s manual, March 2017 |

*End*