#### Performing Instrument to Instrument Correlation on the Sysmex XN Hematology Analyzer

| ##### Purpose | This procedure describes how to perform correlation studies between two Sysmex XN hematology analyzers |
| --- | --- |

| ##### Specimen requirements | - Specimen is EDTA whole blood with minimum volume of 3 ml - Hemogram results of specimens selected for correlation study should fall within laboratory’s normal ranges - Specimens should be kept at room temperature during the analysis - Specimens should be run on all analyzers within 2 hrs of analysis on the reference analyzer - Specimens must not be placed on mechanical blood mixers |
| --- | --- |

| ##### Procedure | Step; Action. 1; Identify one instrument as a reference analyzer in which the other analyzer ( test analyzer) will be compared to.. 2; Select 10 specimens and run all 10 specimens on the reference analyzer.. 3; Record results for WBC, RBC, HGB, HCT, MCV, PLT, %RET and calculate the average results for each parameter obtained on the reference analyzer.. 4; Run the same 10 specimens on the test analyzer.. 5; Record results for WBC, RBC, HGB, HCT, MCV, PLT, %RET and calculate the average results for each parameter obtained on the reference analyzer ( ensure specimens are run on the test analyzer within 2 hrs of running on the reference analyzer).. 6; Compare each averaged parameter result on the test analyzer to the same averaged result on the reference analyzer and ensure the results are within manufacturer’s acceptable limits. If; Then. Correlation study within acceptable limits; Proceed to step 7. Correlation study outside acceptable limits; - Call Sysmex Tech support for further instructions. |
| --- | --- |

Continued on next page

Performing Instrument to Instrument Correlation on the Sysmex XN Hematology Analyzer, Continued

| ##### Procedure, cont. | Step; Action. 7; Correlation is complete. File all correlation documents accordingly.. |
| --- | --- |

| ##### Intrument to Instrument Acceptable Limits | Paramter; Limits. WBC; +/- 7.5 %. RBC; +/- 3.0 %. HGB; +/- 3.5 %. HCT; +/- 3.0 %. MCV; +/- 3.0 %. PLT; +/- 12.5 %. %RET; +/- 30.0%. |
| --- | --- |

| ##### References | - Sysmex product Notification Sysmex XN series Instrument to Instrument Correlation Document # 62-1457 4/2018 |
| --- | --- |

End