#### Fluid Cell Counts:Pleural, Peritoneal, Pericardial, Synovial

| ##### Purpose | This procedure describes how to perform cell counts on body fluids. |
| --- | --- |

| ##### Policy | - Diluting fluids will be checked each day of use for non specimen background particulates and documented on the provided log - The cell count is correlated with the number of cells in the stained smear to ensure accuracy of the chamber count - 2% Acetic Acid is used to assist in distinguishing RBC’s from other cell types - Phase Microscopy is used as an option to assist in distinguishing RBC’s from other cell types |
| --- | --- |

| ##### Specimen Requirements | - Fluids should be submitted in a lavender or heparinized tube. - If it is not, remove an aliquot, and place on an EDTA or green top tube, place on rocker to mix. - If received “clotted”, call physician to inquire if still want the cell count. If physician approves, enter the result comment “specimen clotted, count may be inaccurate”. - Check the specimen macroscopically for fibrin strands (may indicate cell clumping ) and debris, note them on worksheet & enter a comment in LIS indicating that results may be affected. Remove the strands before diluting &/or counting. - If cell clumps are observed in counting chamber, enter a comment in LIS stating that results may be inaccurate due to cell clumping. - Labeled per protocol |
| --- | --- |

| ##### Diluting Fluids/reagents | - Normal Saline - Stored at room temperature - See Mfr. exp. date - 2% Glacial Acetic Acid (Used to lyse RBC’s) - To 98 mL water, add 2 mL concentrated glacial acetic acid. - Place into a properly labeled container, with preparation & expiration dates. - Stable at room temperature, record a 1 year expiration date 3. Hyaluronidase - Store frozen |
| --- | --- |

Continued on next page

Fluid Cell Counts:Pleural, Peritoneal, Pericardial, Synovial, Continued

| ##### Procedure | Step; Action. 1; Generate a FLUID WORKSHEET: - RVFLDM – pleural, peritoneal, pericardial - RVSFM – synovial. 2; Record fluid total volume.. 3; Record color and appearance of fluid.. 4; Check the specimen macroscopically for fibrin strands( may indicate cell clumping) and debris. Notate on worksheet and enter a comment in LIS stating results may be inaccurate due to cell clumping when confirmed in counting chamber.. 5; Determine the appropriate dilution when necessary which may require a dilution for RBC & another dilution for WBC. Properly label each 10 X 75 tube used for each dilution: - Use 2% Glacial Acetic acid to lyse RBC’s and count WBC’s - Use Normal Saline to dilute and count RBC’s - 1:2 = 100 L Diluent + 100 L patient sample - 1:11 = 100 L Diluent + 10 L patient sample - 1:101 = 1.0 mL Diluent + 10 L patient sample. 6; Plate both sides of the chamber using a microhematocrit capillary tube: - If RBC & WBC have to be counted separately, (different dilutions & or using glacial to lyse the RBC), use 2 counting chambers. They can be plated at same time.. 7; Cell counts are performed in duplicate by counting both sides of the hemocytometer. |
| --- | --- |

Continued on next page

Fluid Cell Counts:Pleural, Peritoneal, Pericardial, Synovial, Continued

| ##### Procedure,cont | Step; Action. 8; In general, count the WBC’s in 5 large squares on both sides (count the 4 “corners” & the “middle” on each side for a total of 10 large squares) - Counts should agree within 15% of each other. If not then repeat. 9; In general, count the RBC’s using the large middle square ( count the 5 small RBC squares on each side for a total of 10 RBC squares) - Counts should agree within 15% of each other. If not then repeat - Use of Phase Microscopy can assist in distinguishing red blood cells from other cell types. 10; Record the following information on the fluid worksheet for both RBC and WBC - Counts for both sides of the counting chamber - Dilution factor and Diluting fluid used - Number of squares counted - Volume used. 11; Enter the information above into the LIS to obtain the calculated RBC and WBC counts. Record the calculated counts on the worksheet:. If; Then. WBC manual calculation; Cells counted X Dilution factor Squares counted X Volume. RBC manual calculation; Cells counted X Dilution factor Squares counted X Volume. 12; Make the body fluid smear using the Cytospin.. 13; Stain the body fluid smear using the Wright Giemsa stain.. |
| --- | --- |

Continued on next page

Fluid Cell Counts:Pleural, Peritoneal, Pericardial, Synovial, Continued

| ##### Procedure,cont | Step; Action. 14; Perform the fluid differential: - If the counting chamber WBC count is ≤ 5 and body fluid smear shows insufficient cells to perform 100 cell diff then type HIDE for the diff paramters. Append ETC NODIF in the fluid comment field - The fluid differential is used as a procedural control by correlating the chamber count to the numbers of cells in the stained smear - If chamber count correlates with number of cells in stained smear then check the “QC Run Accepted” prompt on the fluid worksheet - See chart below for procedural control guideline. If; Then. No cells seen in the chamber count; 0 – 2 cells per 20X (LPF). 5 cells in chamber count; > 8 cells per 20X (LPF). 20 cells in chamber count; > 7 cells per 50X ( oil). 15; Differentiate the cells as Lympocytes, Polymorphonuclear, Eosinophils and Histiocytes/Monocytes. 16; Quantitate mesothelial cells as few, moderate, or many and result in the Fluid Comment Field.. 17; If “suspect tumor” cells are seen, result this as “Atypical cells seen” in the Fluid comment field. Then refer the slide to a pathologist for review. (See Smear, Pathologist Review” SOP).. |
| --- | --- |

Continued on next page

Fluid Cell Counts:Pleural, Peritoneal, Pericardial, Synovial, Continued

| ##### Reference Ranges | - Synovial Fluids (Adults) = 0 - 200 WBC’s / mm3 - Pleural & Pericardial Fluid: - Total leukocyte counts in Pleural & Pericardial Fluid are of limited value. In general, WBC’s < 1000/L are associated with transudates, while WBC’s > 1000/L are associated with exudates Total leukocyte counts in Peritoneal Fluid are of limited value. In general, a total leukocyte count >300/L is considered abnormal |
| --- | --- |

| ##### Related Documents | - Diluting Fluid Quality Control SOP - Evaluating Patient Flagged Results SOP Appendix A: Path Review Criteria - Cytospin SOP |
| --- | --- |

| ##### References | - Technical Hematology, Arthur Simmons, 2nd Edition, J.B. Lippincott Company, pp 339-341, 1976 - Kjeldsberg, C. R. Knight, J.A., Body Fluids, Laboratory Examination of Cerebrospinal, Synovial and Erops Fluids: A Textbook Atlas, ASCP, Chicago, 1982, reprint 1984 - Color Atlas of Body Fluids, CAP 2006 |
| --- | --- |

End