#### ESR, Automated Maintenance

| ##### Purpose | This procedure describes how to perform daily maintenance on the ESR Auto plus |
| --- | --- |

| ##### Policy | - The daily maintenance is comprised of running the test rack which is used to verify accuracy throughout the reportable range - The test rack is run each day of patient testing and prior to the QC run |
| --- | --- |

| ##### Procedure | Step; Action. 1; At the Standby mode, press and hold “X” until you hear a beep to access the Main Menu. 2; Enter menu selection number 5 ( Run Test Rack).. 3; When prompted, insert the test rack. 4; After the reading is complete, results will print out. If; Then. Test rack results failed; - Press “√” to retry - Re register test rack (see operator’s manual) - Call tech support at (800) 843-0912. Test rack results passed; Continue to step 5. 5; Staple the result printout to the ESR worksheet.. 6; File the worksheet appropriately when test is completed.. |
| --- | --- |

| ##### Related Documents | - ESR Auto plus maintenance record logsheet |
| --- | --- |

| ##### References | - ESR Auto Plus operator’s manual by Streck laboratories |
| --- | --- |

End