Washing Red Blood Cells for Testing Using the Helmer UltraCW II Cellwasher

| ##### Purpose: | To provide instruction for washing red cells to remove the unbound proteins before the addition of Anti-Human Globulin for direct and indirect antiglobulin tests using a Helmer UltraCW II cellwasher. |
| --- | --- |

| ##### Equipment/Reagents/Supplies: | Equipment | Reagents | Supplies |
| --- | --- | --- | --- |
|  | - Helmer UltraCW II | - 0.9% saline | - Test tubes - Disposable pipettes - Segment sampler - Test tube rack |

| ##### Specimen Requirement: | Specific to individual procedure(s) being performed. |
| --- | --- |

| ##### Procedure: | - Follow the steps below to wash red cells using Helmer UltraCW II cellwasher. |
| --- | --- |

| Step | Action |
| --- | --- |
|  | Place the labeled tubes with red blood cells to be washed in the cellwasher carriage holder, ensuring balance. |
|  | Select the Program Menu icon. |
|  | Select either Wash x3 or x4 cycle. |
|  | Select green arrow icon to move forward and review wash steps. |
|  | Select the green check mark at the end of the program review to accept. |
|  | Press Start Program icon. |
|  | When the cellwasher beeps indicating the end of the cycle, press the Lid Open icon on the screen to unlock the lid. |
|  | Remove the tubes from the cellwasher. |
|  | A dry cell button will remain. |
|  | Proceed with testing. |

| ##### Limitations: | - Loss or hemolysis of the red blood cells from improper washing may result in false results. - Failure to adequately wash can cause false-negative reactions due to neutralization of the AHG reagent by unbound globulins - False negative results may occur if the washing process is delayed or interrupted. |
| --- | --- |

| ##### References | - Helmer UltraCW II Operating Manual |
| --- | --- |

| ##### Related Procedures: | - Direct Antiglobulin Test - Antibody Screen Tube Procedures - Antibody Identification Tube Procedures - Antibody Titration Procedure - Antigen Screening Procedure, Indirect - AHG Crossmatch, Tube Procedure |
| --- | --- |