Using the Helmer UltraCW II Cellwasher

| ##### Purpose: | To provide instruction for using the Helmer UltraCW II cellwasher as a serological centrifuge or to wash red cells to remove the unbound proteins before the addition of Anti-Human Globulin for direct and indirect antiglobulin tests. |
| --- | --- |

| ##### Equipment/Reagents/Supplies: | Equipment | Reagents | Supplies |
| --- | --- | --- | --- |
|  | - Helmer UltraCW II | - 0.9% saline | - Test tubes - Disposable pipettes - Segment sampler - Test tube rack |

| ##### Specimen Requirement: | Specific to individual procedure(s) being performed. |
| --- | --- |

| ##### Procedure: | - Follow the steps below to use the Helmer UltraCW II cellwasher as a serological centrifuge. |
| --- | --- |

| Step | Action |
| --- | --- |
|  | Press the open lid icon. |
|  | Load labeled test tubes in the carriage holder. Note: Be sure to balance sample tubes. |
|  | Close cellwasher lid. |
|  | Press Menu icon, if program menu is not open. |
|  | Select spin program that corresponds to time posted for testing on cellwasher. |
|  | Press the √ button. |
|  | Press Start program. |
|  | When centrifuge stops, press the √ button. |
|  | Press open lid icon. |
|  | Remove test tubes form cellwasher. |

_________________________________________________________________

Continued on Next page

| ##### Procedure: | - Follow the steps below to wash red cells using Helmer UltraCW II cellwasher. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Press the open lid icon. |
| 2. | Load labeled test tubes in the carriage holder. Note: Be sure to balance sample tubes. |
| 3. | Close cellwasher lid. |
| 4. | Press Menu icon, if program menu is not open. |
| 5. | Select either Wash x3 or x4 cycle. |
| 6. | Select green → to move forward and review wash steps. |
| 7. | Select the green √ mark at the end of the program review to accept. |
| 8. | Press Start Program icon. |
| 9. | When the cellwasher beeps indicating the end of the cycle, press the Lid Open icon on the screen to unlock the lid. |
| 10. | Remove the tubes from the cellwasher. |
| 11. | A dry cell button will remain. |
| 12. | Proceed with testing. |

| ##### Limitations: | - Loss or hemolysis of the red blood cells from improper washing may result in false results. - Failure to adequately wash can cause false-negative reactions due to neutralization of the AHG reagent by unbound globulins - False negative results may occur if the washing process is delayed or interrupted. |
| --- | --- |

_______________________________________________________________

| ##### Appendices | - Appendix A Error Codes – 36038/D |
| --- | --- |

| ##### References | - Automatic Cell Washing System UltraCW II Quick Reference Guide |
| --- | --- |

| ##### Related Procedures: | - Performing ABO and Rh Test - Direct Antiglobulin Test - Antibody Screen Tube Procedures - Antibody Identification Tube Procedures - Antibody Titration Procedure - Antigen Screening Procedure, Indirect - AHG Crossmatch, Tube Procedure |
| --- | --- |