## Cell Washer and Centrifuge Maintenance

| ##### Purpose | The purpose of this procedure is to ensure cell washers and centrifuges are properly used and maintained in safe and efficient operating condition. |
| --- | --- |

| ##### Policy | - Cell washers and centrifuges will be inspected and cleaned daily. - Cell washer saline volume will be checked each time the saline cube is changed or at a minimum of once per month. - Interior and exterior spills and/or tube breakage will be cleaned up as they occur. - Cell washers will be flushed with cleaning solution weekly. - Routine maintenance will be performed on Day shift - Centrifuge RPM will be checked semi-annually by eQuip |
| --- | --- |

| ##### Equipment/ ##### Supplies |  |
| --- | --- |

| Equipment | Reagents | Supplies |
| --- | --- | --- |
| - Helmer Cell Washer - Serologic centrifuges - Graduated Cylinder | - 10% Bleach Solution - Bleach wipes - Blood Bank Saline - Mild detergent | Centrifuge tubes Dry cloth |

| ##### Procedure: Daily Maintenance | Cell washer must be inspected and cleaned daily and whenever a tube breaks in the centrifuge or cell washer. |
| --- | --- |

|  |  | Step | Action |  |
| --- | --- | --- | --- | --- |
|  |  | 1. | Wipe down centrifuges with bleach wipes or 10% bleach solution daily. |  |
|  |  | 2. | Load the rotor of the cell washer with tubes, leaving every other position on the rotor empty. |  |
|  |  | 3. | Close the lid. |  |
|  |  | 4. | Prepare a container of approximately 1L of deionized or distilled water. |  |
|  |  | 5. | Remove the Saline tube from the saline container and place it into the container with the deionized or distilled water. |  |
|  |  | 6. | Select and run the “flush 1” program |  |
|  |  | 7. | Remove the saline tube from the deionized or distilled water container, dry the exterior of the tubing and place it into the saline solution container. |  |
|  |  | 8. | Select and run the “flush 1” program. |  |
|  |  | 9. | Open the lid of the cell washer. |  |
|  |  | 10. | Using damp cloth wipe down the bowl of the cell washer. |  |
|  |  | 11. | Using a dry cloth wipe down the entire inside of the lid, including the drainage system and painted surfaces of the cell washer. |  |
|  |  | 12. | Document each task performed by initialing the appropriate space on the Transfusion Service Maintenance Log. |  |
| ##### Procedure: Weekly Maintenance of Cell washer | Perform steps 1-3 more whenever ports are clogged, preventing saline solution from properly dispensing into tubes. |  |  |  |

| Step | Action |
| --- | --- |
| 1. | Remove the rotor from the cell washer. |
| 2. | Soak rotor is clean warm water or run warm water directly into the top of the rotor for several minutes. Note: Ensure water is flowing freely out of all of the ports. |
| 3. | If a port is blocked, gently slide the lid release tool in and out of the fill port several times from the outside toward the center of the rotor. Note: If rotor will be used immediately for testing, run “flush 1” program before performing testing. |
| 4. | Fill the rotor of the cell washer with tubes, leaving every other position on the rotor empty. |
| 5. | Install the rotor and close the lid. |
| 6. | - Prepare a container with 400ml of 10% bleach cleaning solution. |
| 7. | Prepare a container of approximately 1L of deionized or distilled water. |
| 8. | Remove the Saline tube from the saline container and place it into the container with the bleach cleaning solution. |
| 9. | Select and run the “flush 1” program and wait 5 minutes. |
| 10. | Remove the saline tube from the container, and place it into the container with deionized or distilled water. |
| 11. | Select and run the “flush 1” program, twice |
| 12. | Open the lid and dry the centrifuging area. |
| 13. | Remove the saline tube from the deionized or distilled water container, dry the exterior of the tubing and place it into the saline solution container. |
| 14. | Select and run the “flush 1” program, twice. |
| 15 | Document each task performed by initialing the appropriate space in the Weekly maintenance section of the Transfusion Service Maintenance Log. |

| ##### Procedure: Monthly cell washer maintenance | The Cell washer must be primed and the saline fill volume checked whenever the saline supply or tubing is changed, or at a minimum of monthly as part of preventive maintenance. This process may only be performed by a CLS. |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Replace saline cube as needed. - Document Saline Cube replacement on the corresponding date line in the Cellwasher Saline Cube section of the Transfusion Service Maintenance log. - Write opened expiration date in appropriate space on log and on container. (Saline open date expiration is 1 month after cube opened.) |
| 2. | - To prime or check the saline volume: - Open the Cell washer lid - From the Start screen, touch the login button - Enter password 46060 and select the green √ to confirm - Select the settings button - On the systems setting screen, select the Tools button to open the Service Menu - On the service Menu screen, select the ruler icon to open the Calibration screen - Holding a graduated cylinder under the saline dispenser in the lid, repeatedly press the “Fill button” until no bubbles appear. - Empty graduated cylinder - Touch the target calibration volume (36 ml) - Hold a clean, dry graduated cylinder below the spout on the lid and press the Start button. |
| 3. | Once the liquid has stopped flowing from the spout, measure the liquid collected in the graduated cylinder. If:; Then:. Total volume is within 34.2-37.8 ml; Cell washer is acceptable for use.. Total volume is not within 34.2-37.8 ml, first check.; Repeat volume check steps above.. Total volume is not within 34.2-37.8 ml on repeat volume check.; - Use the “+”, “++” , “-“ and “—“ on the right side of screen to adjust the number as needed. - Select the green √to finish calibration - Use red arrow to return to the Home screen.. |
| 4. | Power Off and on device to return the device to locked mode. |

| ##### Procedure: Splash Guard | Remove and replace Splashguard as needed |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Remove the splash guard cap by pulling upward. |
| 2. | Remove the splash guard by folding back the seal to expose the edge of the splash guard, then gently lift the plash guard. Do this a section at a time until the entire guard is clear of the gasket, then lift from the bowl. |
| 3. | To reinstall the splash guard, place the splash guard in the bowl so that the drain hole in the guard is directly above the drain in the bowl. |
| 4. | Working one section at a time around the bowl, fold back the gasket and press the splash guard downward so the edge of the guard rests on the lip of he bowl. Note: The edge of the guard should slightly overlap the lip of the bowl. |
| 5. | With the labeled side up, place the splash guard cap into the bowl on top of the splash guard. |

| ##### Procedure: Annual Cell Washer Maintenance |  |
| --- | --- |

| Step | Action |
| --- | --- |
| 1. | Replace supply and drain tubing |
| 2. | Replace tube holder inserts for 10 x 75 mm tube, if used. |

| ##### Attachments | - Transfusion Service Maintenance Log |
| --- | --- |

| ##### Reference | - Operation Manual for Helmer Ultra CWII - Service Manual for Helmer Ultra CWII |
| --- | --- |