Procedure Change/Revision (IMM 36-C)

Procedure:  Quality Control Program (Doc# IMM 174)

Change/Revision: The 1-3S rule has been changed.

Control results >3 SD are no longer being repeated in duplicate. Repeat the out of range control level only once. See below for updated actions.

    - Rule 1-3S:

13s refers to a control rule that is commonly used with a Levey-Jennings chart when the control limits are set as the mean plus 3s and the mean minus 3s. A run is rejected when a single control measurement exceeds the mean plus 3s or the mean minus 3s control limit.

Action to be taken:

Rerun the control level

  - If the result is < 2 SD the QC is accepted.

  - If the result is >2 SD, calibrate.

  - If after calibration, any level is > 2SD consult a supervisor.

Date of Revision:  10/11/12

Director Signature/Date:

Technologist Signature/Date: SEE MTS