I.	SCOPE:

The Hematology Lab minimally utilizes deionized water from the taps in the laboratory. 	Water used for reagent hydration, control dilutions, and stain preparation is either sterile 	distilled/deionized water or Nerl water. The “tap” deionized water is used in the lab 	primarily to make 10% bleach solutions and rinsing salt deposits on the instruments.

The Hematology Lab monitors the deionized water taps for microbial growth every	6 months, after major distiller repair and if a perceived problem exists with water.  The

areas monitored are:, routine coags (1), urines (1), special coags (1)  sp hem (1), and

platelet function (2). Water testing is performed by the Microbiology Laboratory on Tuesday mornings (samples need to be in micro lab before 10am)

II.	PROCEDURE:

A.	Remove tubing from distilled water tap nozzle

B.        Run deionized water tap for 5 minutes

C.	Collect 100ml of deionized water into Nalgene 150ml analytical unit

(Storeroom #111188 size 0.45 mic)

D.	Cap unit and label with name of tap where water collected

E.         Order water testing on Soft and print barcode label (procedure attached)

(Alternately can be found on My computer, L Drive, Microbiology, water testing,

water testing for Hematology)

F.         Place label vertically on bottom portion of water filter

G.	Deliver filter apparatus to Micro lab for analysis

H.        Review results:

- Acceptable- file initial results and file

- Unacceptable- distilled water from unacceptable tap cannot

be used until problem resolved and tap retested and is satisfactory.

III.	SOURCE INFORMATION:

A.	Urinalysis:

Sterile distilled water is used to prepare the lyophilized controls.

B.	Coagulation:

NERL Type I water is purchased from a manufacturer and is certified.

C.	Bone Marrow:

NERL Type I water is used to prepare the iron stain for bone marrow.

D.	Main Hematology:

Deionized water used to dilute bleach and as rinse to remove salt crystals

E.	Department wide:

The source - deionized water, Type I, is analyzed by the Clinical Microbiology

IV.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 10-30-2003.

H-2	This procedure was revised by N. Ortoli-Drew on 9-2005.

H-3      This procedure was revised by N. Ortoli-Drew 1-2011.

H-4      This procedure was revised by N. Ortoli-Drew 5-2011.