- PRINCIPLE:

The STA-Liatest vWF kit, (Cat # 0518), is intended for the immuno-turbidimetric

determination of vonWillebrand factor (VWF) by the measurement of light

produced by a suspension of microlatex particles coated with specific antibodies.

When a beam of monochromatic light is allowed to transverse a suspension of

microlatex particles to which specific antibodies have been attached and if the

wavelength of the light is much greater that the diameter of the latex particles,

the light is only slightly absorbed.  In the presence of the antigen being tested for,

the antibody-coated latex particles agglutinate to form aggregates of diameters

greater than the wavelength of the light and the latter is more absorbed.  This increase

of the absorption is proportional to the antigen level present in the sample.

- II.	SPECIMEN:

A.	Citrated blood 9:1 (blood to anticoagulant) 3.2% or 3.8% sodium citrate.

- Follow NCCLS guidelines H3-A3 and H21-A2.  No other anticoagulant

is acceptable.

B.   	Centrifugation:	10 minutes at 3000 g

C.   	Plasma storage:	3 month at -70 C.  Frozen plasmas should be thawed at

37C  for 5 minutes.

D.  	Unacceptable specimens: Samples that are QNS or clotted;  hemolyzed

may yield incorrect results and should be commented.

- III.	REAGENTS:

- Centrifuge

Pipettes

Pipette tips

STA/STA Compact

Cuvette roll (contains 1000 cuvettes)

- REAGENTS

A.	STA vWF Calibrator (Cat # 0520): Freeze-dried citrated human plasma.

Reconstitute each vial with 1.0 ml reagent grade water.  Let set 30 minutes at 				room  temperature.  Swirl gently.  Reconstituted stability on the STA/STA Compact is

4 hours.

B	STA  Liatest Control N+P (Cat # 0526): Control N is citrated normal human plasma and  	Control P is citrated abnormal human plasma, both freeze-dried.  Reconstitute each vial 	with 1.0 ml reagent grade water.  Let set 30 minutes at room temperature.  Swirl gently.  	Reconstituted stability on the STA/STA Compact is 8 hours.

C.  	Reagent 1: Glycine Buffer: Ready to use.  Request the reagent drawer to open

through the MAIN MENU under LOADING and bar code the buffer.  Place

the buffer into the reagent drawer, RD1.  Allow the vial to stand on the STA/STA

Compact for 15 minutes before use.  Stability on the STA/ STA Compact is 15

days.

D.  	Reagent 2: Latex: Suspension of microlatex particles coated with rabbit

anti-human vWF antibodies, then stabilized with bovine albumin.  Pour the entire

contents of a vial of Reagent 3 into a vial of Reagent 2 of the same kit.  Ensure

that all drops of Reagent 3 are transferred.  Swirl gently.  Let set 5 minutes at

room temperature.  Request the reagent drawer to open through the MAIN MENU

under LOADING and bar code the reagent.  Place the reagent into the reagent

drawer, RD2.  Allow the vial to stand on the STA/STA Compact for 15 minutes

before use.  Stability on the STA/STA Compact is 15 days.

E.  	Reagent 3: Latex Diluent: Ready for use.

- IV.	CALIBRATION (if needed):

A. 	STA vWF Calibrator: After the reconstitution period, request the sample

drawer to open (STA) or the reagent drawer to open (STA Compact) through the

MAIN MENU under LOADING and bar code the calibrator.  Place the calibrator

into the appropriate drawer.  The assay value for vWF is transferred to the

STA/STA Compact when the calibrator is loaded and prompted for the bar code

information.

B	All standard dilutions are automatically prepared by the STA/STA Compact by diluting 	the STA vWF Calibrator with Owren-Koller Buffer according to the 	        	parameters entered in the Test Set-up.

C.   	To order the calibration: Through the MAIN MENU select CALIB/CONTROL.

Select CALIBRATION and enter.  Place the cursor on vWF and enter.  Select

vWF by pressing F1 and then F10 to run the calibration.

D.   	STA vWF Control N and P are used by the STA/STA Compact to automatically

validate the calibration.  (Refer to the QUALITY CONTROL section).

E.   	To examine the calibration curve: Through the MAIN MENU under

CALIB/CONTROL select CALIBRATION and enter.  Place the cursor on

vWF and enter.  The calibration curve will appear on the STA/STA Compact

screen.

F. 	To verify the calibration curve is valid: Check that the correlation coefficient

(r value) of the curve is 0.90 or greater.  Compare the actual vs the interpolated

values of the curve.  If a data point difference is greater than 10% for the actual

vs the interpolated value, the point can be deleted from the curve data.  When this

is done, the STA/STA Compact will automatically re-draw the curve, re-calculate

the r value and re-run the controls.  If any controls are outside the expected range,

the control must be re-run.  If the value is still outside the expected range, a new

calibration must be run along with fresh controls.

G. 	To print the calibration curve: While examining the calibration curve on the

STA/STA Compact screen, press ESC for options.  Select PRINT.  The

STA/STA Compact cannot print a calibration curve while the STA/STA Compact

is running.

- V.	QUALITY CONTROL:

A.	STA Liatest Control N+P: After the reconstitution period, request the

sample drawer to open (STA) or the reagent drawer to open (STA Compact)

through the MAIN MENU under LOADING and bar code the controls.  Place

the controls into the appropriate drawer.

B.  	QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC when

the patients’ samples have been loaded.  The QC can also be ordered manually

before loading patients’ samples.

C.  	All control ranges are monitored automatically by the STA/STA Compact.  If any

controls are outside the ± 2 SD range, the STA/STA Compact will audibly and

visually alarm the operator.  Otherwise, the results can be found in the AUTO

CONTROL file (STA only) and the individual QC files.  Control results are

automatically filed in the STA/STA Compact QC file.  All results for a 24 hour

period will be converted to a mean value at midnight.  This mean is used in the

statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- VI.	PROCEDURE:

A.  	Refer to the START-UP procedure for the STA/STA Compact before running 		       	patient samples on the STA/STA Compact at the start of each shift.

B. 	Run the calibration if needed.

C.  	Request quality control if using a stored curve: Through the MAIN MENU

under CAL/CONTROL select QUALITY CONTROL and enter.  Place the

cursor on vWF.  Select vWF by pressing F1 and then F10 to run the QC.

D.  	Load patients’ samples: Access the sample drawer through the MAIN MENU

under LOADING.  After the drawer opens, identify the type of specimen, such

as micro-sample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.  		In MANUAL MODE, the operator must order the “VWF 1/8” test from the menu 			for each samples, and then press F10 to save. If this dilution yields a result which is 			less than 60% vWF antigen, the test must be repeated on the “VWF 1/2” test.

F.	In AUTO MODE, the STA/STA Compact will automatically order the test if it is selected 	in the AUTO MODE profile.

G.	In TELELOADING MODE, the STA/STA Compact will query the host computer and 	download the test as well as assign the status, i.e. stat.

H.	As soon as the sample drawer closes, the TEST PANEL screen will appear.

If  there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

I.  	All dilutions of the calibrator, controls and patents’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

dilution provided that the option has been entered in the Test Set-up.

J.  	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

K. 		For results in question that need operator intervention, cursor to the identification

number on the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left hand corner of the screen,

i.e. re-run test.

- VII.	CALCULATIONS:

A.	The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to % antigen.

B.	The STA/STA Compact uses a dilution of 1:8, sample to buffer.  Following pre-set 			parameters, the STA/STA Compact will automatically dilute the sample lower if the 			result falls below the lowest data point on the calibration curve.  The STA/STA Compact 			will also automatically dilute the sample at a higher dilution if the result falls above the 			highest data point on the calibration curve.

VIII.	REFERENCE RANGE:

_50____ - __150___% Each laboratory should determine its own reference range.

- IX.	REPORTING RESULTS:

The results for vWF are reported out to the nearest whole number %.  (Example:

65%.)

- XI.	LIMITATIONS OF THE PROCEDURE:

A.  	Discard any suspect sample.

B.   	Cloudy plasmas may lead to an under-estimation of the vWF level.

C.	The presence of rheumatoid factor may lead to an over-estimation of the vWF level.

D.  	The extremely rare presence of anti-bovine and/or anti-rabbit antibodies

in certain patients may lead to an over-estimation of the vWF level.

E.	The detection threshold is 2%; assay is linear to 105% (results which are >105% are run on a different dilution).

F. 	No prozone effect has been observed for vWF levels up to 800%.

G.    	The system is insensitive to hemoglobin, bilirubin and heparins (unfractionated

and low molecular weight) up to 5 g/l., 150 mg/l and 2 IU/ml  respectively.

H.   	von Willebrand disease (vWD) is the most common inherited bleeding disorder.

Clinically, it is often characterized by mucocutaneous hemorrhages.  The three

principle types of vWD are:

- type 1  corresponds to a quantitative deficiency of  vWF

- type 2  refers to a qualitative deficiency of  vWF

- type 3  is characterized by a total absence of  vWF

- XII.	REFERENCES:

1.  	STA-LIATEST vWF, Immuno-Turbidimetric Assay of van Willebrand Factor

(vWF) by STA Analyzer. Package insert 23327-Revised November 1997.

- 2.   	STA-vWF Calibrator, Calibration Plasma for the Immuno-Turbidimetric

Assay of vWF on STA Analyzer. Package insert 23405-Revised April 1997.

3.   	STA-Liatest Control N+P, Control Plasmas for the Immuno-Turbidimetric

Assays -Revised December 1997.

- STA/STA Compact Operators Manual.

XIII.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-22-2004.

H-2	This procedure was reviewed by N Kriz on 2/1/2010.

H-3	This procedure was reviewed by N Kriz on 2/3/2011.

H-4	This procedure was reviewed by N Kriz on 1/2/2012.

H-5	This procedure was reviewed by N Kriz on 1/28/2013.