I.	PROCEDURE:

A.	Once per week, coagulation supervisor or specialist should check the Easy Q			program on the BCS instrument to ensure that action comments were made if QC 			results 	were out of range. Results that are out of range should be deleted from the 			statistical calculations, and a comment entered as to why the deletion was done.

Tech making such comments should place initials after the comment.

B.	At the end of each month, after a final inspection/deletion of QC data, the file for 		each parameter (PT, PTT, fibrinogen, low fibrinogen, and d-dimer 1 and 2 should be printed out and initialed by supervisor.

C.	Each file should then be closed; a new file will be opened the next time the 			controls are run.

D.	For each file, calculate 1 SD as follows:

CV x mean     =    1SD

100

E.	Supervisor determines the 1SD for each file and compares the reference and actual means for each file

F.	Submit statistical data to Siemens Quality Assurance Program via their 				website (see attached instructions). When reports are compiled, file them in the 			notebook by month.

1.	Go to bookmark Siemens Hemostasis QAP Online

2.	Click “Lablink Online Login”

3.	User ID  11167

4.	Password 51605560

5.	Click “Login”

6.	Follow instructions on “Lablink Online Quick Instructions” (see following 			pages)

7.	Print reports and file in notebooks over supervisor’s desk.

G.	If any parameters show exception messages, some appropriate remedial action 			must be taken and recorded on exception sheet (for example: if the lab mean +/- 2 			S.D. range differs significantly from the test peer range, note whether 				recalibration was done, new lot # reagent started, etc, as fits the situation). Be sure 		to monitor the next month’s data carefully to assure that the problem has not 			recurred.

H.	If no explanation for the flag is readily apparent, recheck 1SD calculation and data entry. If no mistake found, contact Siemens technical support for additional help in addressing the issue. Document any follow-up action and continue to monitor the monthly statistical data for all parameters.

II. 	HISTORY:

H-1  	This procedure was written by N. Kriz on 1-12-2007.

H-2	This procedure was reviewed by N Kriz on 3/6/2010.

H-3	This procedure was revised by N Kriz on 2/4/2011.

H-4	This procedure was reviewed by N Kriz on 1/2/2012.

H-5	This procedure was reviewed by N Kriz on 1/28/2013.