- I.	PRINCIPLE:

Antithrombin (AT) exerts a powerful and immediate inhibitory action on thrombin when 	heparin is present.  The STA - Stachrom AT III kit, (Cat # 0596), test procedure consists 	of two steps:

A.	Plasma is incubated with a known excess of thrombin in the presence of 	heparin.

B.	The residual thrombin is quantitated by its amidolytic action

on the synthetic chromogenic substrate CBS 61.50 (pNA released

at 405 nm).  Since the quantity of thrombin that is neutralized in the

first reaction step is proportional to the AT  level present in the

plasma being tested, it follows that the residual thrombin in the

second reaction step is inversely proportional to the AT level

of the tested plasma.

- Since the first report (1965) of a hereditary deficiency of AT and its

consequences,  AT has been considered an important parameter in spontaneous

thromboembolic disorders.  The hereditary deficiency, either quantitative or qualitative,

is less common than the acquired deficiency.  An acquired AT deficiency has been

described in DIC, septic shock, nephrotic syndrome, liver disease and in L-

asparaginase treatments.

- II.	SPECIMEN:

A.  	 Citrated blood 9:1 (blood to anticoagulant) 3.2 % sodium citrate.

Follow NCCLS guideline H21-A2 and H3-A3.  No other anticoagulant

is acceptable.

B.   	 Centrifugation:	20 minutes for 3000 g

C.   	 Plasma Storage: 	  8 hours at 20 C

48 hours at 2-8 C

3 month at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

- Unacceptable Specimens:  Samples that are QNS or clotted

- III.	REAGENTS AND EQUIPMENT:

- Centrifuge

Pipettes

Pipette tips

Plastic test tubes

STA/STA Compact

Cuvette roll (contains 1000 cuvettes)

- IV.	REAGENTS:

A.   	STA Unicalibrator: Freeze-dried citrated human plasma.  This plasma is intended

for use as a calibration plasma for functional assays.  Reconstitute each vial with

1.0 ml reagent grade water.  Let set 30 minutes at room temperature.  Swirl vial

gently.  Reconstituted stability on the STA/STA Compact is 4 hours.

B.  	STA- System N and P: Control N is citrated normal human plasma and

Control P is citrated abnormal human plasma, both freeze-dried.  Reconstitute 			each vial with 1.0 ml reagent grade water.  Let set 30 minutes at room 				temperature.

Swirl gently.  Reconstituted stability on the STA/STA Compact is 8 hours.

C.   	Owren-Koller Buffer: Ready to use.  Request the sample drawer to open through

the MAIN MENU under LOADING and bar code the buffer.  Place the buffer

into the sample drawer.  Stability on the STA/STA Compact is 120 hours.

D.	Reagent 1: Thrombin.  Approximately 36 NIH units of bovine thrombin per vial, 	 	freeze-	dried.  Take one vial of Reagent 3 and shake it well.  Then, pour its entire 	 	contents into a vial of Reagent 1 of the same kit.  Allow the reconstituted reagent 	to stand for 60 minutes at room temperature.  Swirl gently.

E.	Request the reagent drawer to open through the MAIN MENU under LOADING 	and bar code the reagent.  Place the reagent into the reagent drawer, RD 1.  Let 	the reagent sit 	on the STA/STA Compact for 10 minutes before use.  	Reconstituted stability on the 	STA/STA Compact is 7 days.

F.	Reagent 2: Chromogenic substrate CBS 61.50.  Reconstitute each vial with 3.0 ml 	of reagent grade water.  Allow the reconstituted reagent to stand for 60 minutes at 	 	room temperature. Swirl gently.  Request the reagent drawer to open through the	 	MAIN 	MENU under LOADING and bar code the reagent.  Place the reagent into 	the reagent drawer, RD 2.  Let the reagent sit on the STA/STA Compact for 10 	minutes before use. Reconstituted stability on the STA/STA Compact is 7 days.

G.	Reagent 3: Thrombin Solvent.  Ready for use.  Shake vial well before use to 	reconstitute Reagent 1.

- V.	CALIBRATION (if needed) or every 6 months:

A.	STA Unicalibrator: After the reconstitution period, request the sample drawer 	     	(STA) or the reagent drawer (STA Compact) to open through the MAIN MENU 	under LOADING and bar code the calibrator.  Place the calibrator into the 	appropriate drawer. The assay value for AT is transferred to the STA/STA 	Compact when the calibrator is loaded and prompted for the bar code information.

B.	The standards are automatically prepared by the STA/STA Compact by diluting 	the STA- Unicalibrator with Owren-Koller Buffer according to the parameters 	entered in the Test Set-up.

C.	 To order calibration: Through the MAIN MENU select CALIB/CONTROL.

Select CALIBRATION and enter.  Select AT III by pressing F1 and then

F10 to run the calibration.

D.	STA-System N and P are used by the STA/STA Compact to validate the 	calibration. (Refer to the QUALITY CONTROL section).

E.	To examine the calibration curve: Through the MAIN MENU under 		 	CALIB/CONTROL select CALIBRATION and enter.  Place the cursor on ATIII 	and enter. The calibration curve will appear on the STA/STA Compact screen.

F.	Verify calibration curve is valid:  Check that the correlation coefficient (r value)

Of the curve is 0.90 or greater.  Compare the actual vs. the interpolated values

of the curve.  If a data point difference is greater than 10% for the actual vs.

the interpolated value, the point can be deleted from the curve data.  When this

is done, the STA/STA Compact will automatically re-draw the curve, re-calculate

the r value and re-run the controls.  If any controls are outside the expected range,

the control must be re-run.  If the value is still outside the expected range, a new

calibration curve must be run along with fresh controls.

G.	To print the calibration curve: While examining the calibration curve on the 	STA/STA Compact screen, press ESC for options.  Select PRINT.  The

STA/STA Compact cannot print a calibration curve while the instrument is 			running.

- VI.	QUALITY CONTROL:

A.	STA-System Control N and P: After the reconstitution period, request the sample 			drawer to open (STA) or the reagent drawer to open (STA Compact) through the 			MAIN MENU under LOADING and bar code the controls.  Place the controls 			into the sample drawer.

B.	QC will automatically be run when a new calibration curve has been requested.  If 		using a stored curve, the STA/STA Compact will automatically run QC if the 			patients’ samples have been loaded.  QC can also be ordered manually before 			loading patients’ samples.

C.  	All control ranges are monitored automatically by the STA/STA Compact.  If 			any controls are outside the + 2 SD range, the STA/STA COMPACT will audibly 			and visually alarm the operator.  Otherwise, the results can be found in the AUTO 	            CONTROL file (STA only) and the individual QC files.  Control results are

automatically filed in the STA/STA Compact QC file.  All results for a 24 hours

period will be converted to a mean value at midnight.  This mean is used in the

statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- VI.	PROCEDURE:

A.	Refer to the START-UP procedure for the STA/STA Compact before running 		      	patient samples on the STA/STA Compact at the start of each shift.

B.   	 Run the calibration only if needed.

C.   	 Request quality control if using a stored curve: Through the MAIN MENU

under CALIB/CONTROL select QUALITY CONTROL and enter.  Select

AT  by pressing F1 and then F10 to run the QC.

D.   	Load patients’ samples: Access the samples drawer(s) through the MAIN MENU

under LOADING.  After the draw opens, identify the type of specimen, such as

micro-sample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.	In MANUAL MODE, the operator must order the test from the menu, and then 	press F10 to save.

F.   	In AUTO MODE, the STA/STA Compact will automatically order the test if it is

selected in the AUTO MODE profile. In TELELOADING MODE, the STA/STA

Compact  will query the host computer and download the test as well as assign 			the status, i.e. stat.

G.  	As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

H.  	All dilutions of the calibrator, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in

the Test Set-up.  If the patients’ results fall outside of the assay range, the

STA/STA Compact automatically re-tests the sample in question at an 				appropriate dilution provided that the option has been entered in the Test Set-up.

I.  	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

J.   	For results in question that need operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left hand corner of the screen,

i.e. re-run test.

- VII.	CALCULATIONS:

A.    	The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to % activity. Linearity for this assay is 0-140% (results  >140% 	automatically redilute.

## VIII.	REFERENCE RANGE:

80-125 %     AT activity

- IX.	REPORTING RESULTS:

Enter AT results on Resulting Worklist; The results for the AT are reported out to the nearest whole number %.   (Example:  104%).

- X.	NOTES:

When the STA/STA Compact re-dilutes a patient sample at a more appropriate 	       dilution, (as pre-determined in Test Set-up) the results in the TEST PANEL screen

will appear in blue and have already been corrected by the STA/STA Compact for

the dilutional difference.  Order test “ATIII 40” as patients are more likely to be

normal;  if result is low, run test “ATIII”.

- XI.	LIMITATIONS OF THE PROCEDURE:

A    	The test procedure is not affected by therapeutic doses of heparin.  Therefore, it is 	suitable for testing of plasmas collected from patients receiving heparin therapy. 	Note that the AT result may be low “normally” on a heparinized patient, since 	ATIII is used in the heparinization process.

B.  	The use of bovine thrombin makes the interference by heparin cofactor II

insignificant.  However in very rare cases, when the AT level is low and

the heparin cofactor II level high, an overestimation of the AT  level is

possible.

C.   	The presence of heparin ensures that the AT exerts its inhibitory action

fully and instantly.  Thus the action of the progressive antithrombins

(i.e. alpha2-macroglobulin) are insignificant.

D.  	The presence of aprotinin ensures that the activity of plasmin is blocked

if it is present.

XII.	REFERENCE:

1.    	STA-Stachrom AT III, Colorimetric Assay of Antithrombin III by STA.

Package insert 24603-Revised December 1997.

2.    	STA Unicalibrator, Calibration Plasma for Functional Assays of Coagulation

Parameters on the STA.  Package insert 26344 - Revised May 1996.

3.   	STA System N + P, Control Plasmas for Assays of Coagulation Parameters

on STA.  Package insert 26384 - Revised May 1997.

- STA/STA Compact Operators Manual.

XIII.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-16-04.

H-2	This procedure was reviewed by N Kriz on 10/12/2010.

H-3	This procedure was reviewed by N Kriz on 2/22/2011.

H-4	This procedure was reviewed by N Kriz on 2/6/2012.

H-5	This procedure was reviewed by N Kriz on 2/13/2013.