- I.	PRINCIPLE:

The Rotachrom Heparin (Cat #0612), is intended for the quantitative determination

of the plasma levels of low molecular weight heparin (LMWH) by the measurement of 	their anti-Xa activity in a competitive system using a synthetic chromogenic substrate.

The normal function of a molecule of factor Xa, as soon as it appears in plasma,

is to cleave its natural substrate, prothrombin, to generate thrombin, the enzyme

responsible for the fibrin clot formation.  In the presence of heparin, competition

occurs between this mechanism and the inhibitory mechanism exerted by the

heparin-antithrombin III complex, this inhibition being largely responsible for the

action of heparin.  The proposed method is a one-step reaction based on a similar

principle: as soon as factor Xa is added to the plasma-substrate mixture, two reactions

take place simultaneously:

-  hydrolysis of the substrate by factor Xa

-  inhibition of factor Xa by the heparin-antithrombin III complex

After the necessary period of time for the competitive reaction to reach equilibrium,

the quantity of pNA that is released is inversely proportional to the concentration

of heparin present in the plasma.  Low molecular weight heparins (LMWH) are used for 	the prevention and treatment of thromboembolic diseases.  The quantitative 	determination of plasma LMWH levels is a useful tool for monitoring treatment efficacy.

The heparin-antithrombin III complex is made up from the heparin and the antithrombin III (ATIII) peculiar to the patient.

- II.	SPECIMEN:

A.  	Citrated blood 9:1 (blood to anticoagulant) 3.2% or 3.8% sodium citrate.

Follow NCCLS guidelines H21-A2 and H3-A3.  No other anticoagulant

is acceptable.

B.  	Centrifugation:	20 minutes at 3500 g - room temperature

C	Plasma storage:	2 hours for 20 C

3 months at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

D.   	Unacceptable specimens: Samples that are QNS or clotted; hemolyzed may yield incorrect results and should be commented.

- III.	REAGENTS AND EQUIPMENT:

- Centrifuge

Pipettes

Pipette tips

STA/STA Compact

Cuvette roll (contains 1,000 cuvettes)

IV.	REAGENTS:

A.  	STA Calibrator HBPM/LMWH: A set of calibrator plasmas intended to be used for low molecular weight heparin (LMWH) assays using the anti-Xa principle.

Reagent 1: STA-Calibrator HBPM/LMWH 0: plasma free of LMWH, freeze-			dried.

Reagent 2: STA-Calibrator HBPM/LMWH 9: plasma containing a well-defined

quantity of heparin (see enclosed Assay Value Insert), freeze-dried.

Reagent 3: STA-Calibrator HBPM/LMWH 18: plasma containing a well-defined

quantity of heparin (see enclosed Assay Value Insert), freeze-dried.

Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at

room temperature.  Do not invert vials.  Swirl gently.  Reconstituted stability

on the STA/STA Compact is 4 hours.

B.	STA- Quality HBPM/LMWH Control: Two plasmas containing different levels of low molecular weight heparin (LMWH) intended to be used for quality control of LMWH assays using the anti-Xa principle.

Reagent 1: STA-Quality HBPM/LMWH Control  8: plasma containing a well-			defined quantity of LMWH (see enclosed Assay Value Insert), freeze-dried.

Reagent 2: STA-Quality HBPM/LMWH Control  14: plasma containing a well-			defined quantity of LMWH (see enclosed Assay Value Insert), freeze-dried.

Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at

room temperature.  Do not invert vials.  Swirl gently.  Reconstituted stability

on the STA/STA Compact is 4 hours.

C.	Reagent 1: Substrate: lyophilized chromogenic substrate CBS 52.44, 	approximately 4.2 umoles per vial.  Reconstitute each vial with 4.0 ml reagent 	grade water. Let sit 60 minutes at room temperature.  Do not invert vials.  Swirl 	gently; install an STA-reducer and replace the perforated plastic cap on the vial.  	Request the reagent drawer to open through the MAIN MENU under LOADING 	and bar code the reagent.  Place the reagent into the reagent drawer, RD1.  	Reconstituted stability on the STA/STA Compact is 7 days.

D.   	Reagent 2: F.Xa: bovine factor Xa, approximately 1.1 IU per vial.  Reconstitute

each vial with the entire contents of one Reagent 3 vial from same kit.  Let sit 60 	minutes at  room temperature. Do not invert vials.  Swirl gently; install an STA-	reducer and  replace the perforated plastic cap on the vial. Request the reagent 	drawer to open through the MAIN MENU under LOADING and bar code the 	reagent.  Place the reagent into the       reagent drawer, RD2.  Reconstituted 	stability on the STA/STA Compact is 7 days.

E.  	Reagent 3: F.Xa Solvent: Ready to use.  Transfer the entire contents of Reagent 3

into a Reagent 2 vial.

- V.	CALIBRATION (if needed) or minimum of every 6 months:

A.	STA-Hepanorm HBPM/LMWH: After the reconstitution period, request the 	sample drawer to open (STA) or the reagent drawer to open (STA Compact) 	through the MAIN       MENU under LOADING and bar code each vial of 	calibrator.  Place each vial of 	       calibrator into the appropriate drawer after it 	has been bar coded.  The assay value

for each Heparin Level is transferred to the STA/STA Compact when each

calibrator is loaded and prompted for the bar code information.

B.	All standard dilutions are automatically prepared by the STA/STA Compact by 	diluting the STA-Calibrator HBPM/LMWH calibrators with Owren-Koller Buffer 	according to the parameters entered in the Test Set-up.

C.  	To order calibration: Through the MAIN MENU select CAL/CONTROL.

Select CALIBRATION and enter.  Place the cursor on HBPM 2UI and enter.

Select HBPM 2UI by pressing F1 and then F10 to run the calibration.

D. 	STA-HBPM/LMWH Controls are used by the STA/STA Compact to 				automatically validate the calibration.  (Refer to the QUALITY CONTROL 			section).

E.  	 To examine the calibration curve: Through the MAIN MENU under

CAL/CONTROL select CALIBRATION and enter.  Place the cursor on

HBPM 2UI and enter.  The calibration curve will appear on the STA/STA

Compact screen.

F.  	 To verify the calibration curve is valid: Check that the correlation coefficient

(r value) of the curve is 0.90 or greater.  If any controls are outside the expected

range, the control must be re-run.  If the value is still outside the expected range,

a new calibration must be run along with fresh controls.

G.	To print the calibration curve: While examining the calibration curve on the 	 	STA/STA Compact screen, press ESC for options.  Select PRINT.  The 	STA/STA Compact cannot print a calibration curve while the STA/STA Compact 	is running.

- VI.	QUALITY CONTROL:

A	STA-HBPM/LMWH Control: After the reconstitution period, request the 				sample drawer to open (STA) or the reagent drawer to open (STA Compact) 			through the MAIN MENU under LOADING and bar code the controls.  Place 			the controls into the appropriate drawer.

B	QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC when

the patients’ samples have been loaded.  The QC can also be ordered manually 		before loading patients’ samples.

C.	All control ranges are monitored automatically by the STA/STA Compact.  If any 	       	controls are outside the ± 2 SD range, the STA/STA Compact will audibly and 	       	visually alarm the operator.  Otherwise, the results can be found in the AUTO 	       	CONTROL file and the individual QC files.  Control results are automatically 	filed in the STA/STA Compact QC file.  All results for a 24 hour period will be 	converted to a mean value at midnight.  This mean value is used in the statistical 	data and is plotted on the Levy-Jennings chart as a daily mean.

- VII.	PROCEDURE:

A.	Refer to the START-UP procedure for the STA/STA Compact before running 	patient samples on the STA/STA Compact at the start of each shift.

B.	Run the calibration if needed.

C.   	Request quality control if using a stored curve: Through the MAIN MENU

under CAL/CONTROL select QUALITY CONTROL and enter.  Place

the cursor on HBPM 2UI.  Select HBPM 2UI by pressing F1 and then  F10 to run 		the QC.

D.   	Load patients’ samples: Access the sample drawer through the MAIN MENU

under LOADING.  After the drawer opens, identify the type of specimen, such

as microsample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.   	In MANUAL MODE, the operator must order the test from the menu, and then

press F10 to save.

F.  	 In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	 	      	 selected in the AUTO MODE profile.

G.  	In TELELOADING MODE, the STA/STA Compact will query the host 				computer and download the test as well as assign the status, i.e. stat.

H.  	As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

I.   	All dilutions of the calibrators, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an 				appropriate dilution provided that the option has been entered in the Test Set-up.

J.   	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

K.  	For results in question, that need operator intervention, cursor to the 				identification number in the TEST PANEL screen and enter.  This will bring up 			the FILE PROCESSING screen.  Follow the options in the left-hand corner of the 			screen, i.e. re-run test.

- VIII.	CALCULATIONS:

A.   The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to IU/ml.

B.   The STA/STA Compact uses a dilution of 1:2, sample to buffer.

IX.	REFERENCE RANGE:

The therapeutic range for Lovenox (LMWH) determined on a specimen drawn 3-5 hours post dose, is 0.5-1.0 U/mL. Therapeutic levels needed may change due to clinical considerations specific to the patient’s condition.

Detection threshold for the LMWH assay is 0.10 U/mL.

- X.	REPORTING RESULTS:

The results for the Heparin Assay are reported out to the nearest 0.1 IU/ml. Post results from the Compact into Soft, with reference to Soft manual if necessary.

(Example:  0.5 IU/ml).

For a result below the detection threshold report “0” U/mL.

- XI.	NOTES:

A.   	The assay’s linearity extends to 2.00 IU/ml.  In order to determine a heparin level 			on a sample with a higher level, the operator must:

-  mix the patient sample 1:1 with Normal Plasma (as the only source of  ATIII,                                                                                                                                                                                     which is required for this assay)

-     re-test the heparin assay on this mixture

- multiply the result  X 2 to correct for the dilutional difference.

- XII.	LIMITATIONS OF THE PROCEDURE:

A.	The lab must know which heparin is being administered.  In the case of 	intermittent  administration of heparin, the time interval between sample 	collection and the previous, or next injection should be specified.

B.	Any release of platelet factor 4 (PF4), which is a potent heparin inhibitor, will 	lead to an under-estimation of the heparin level in the plasma being  tested.  	Careful and adequate centrifugation is essential: the higher the level of residual 	platelets, the greater the risk of PF4 release.

C.	Icteric plasma (bilirubin >6.6 mg/dL) may render result lower than it should be, 	due to intense yellow coloration; the result , therefore, may be  invalid for 	monitoring purposes. Lipemic plasma may be ultracentrifuged prior to use on the 	instrument. Hemolyzed plasma may render results inaccurate and may need to be 	repeated on a fresh sample.

XIII.	REFERENCES:

1.   STA- Rotachrom HBPM/LMWH Colorimetric Assay of Low Molecular Weight 		      Heparin

(LMWH) by STA Analyzers.  Package insert 26377-Revised March 1997.

2.   STA-Hepanorm HBPM/LMWH, Calibration Plasmas for Assays of Low Molecular

Heparin (LMWH) using anti-Xa Method on STA. Package insert 26302-Revised

February 1997.

3    STA-HBPM/LMWH Control, Control Plasmas for Assays of Low Molecular Weight

(LMWH) using Anti-Xa Method on STA.  Package insert 26308-Revised April

1997.

4.    STA/STA Compact Operators Manual.  Revised 3/30/98.

XIV.  	HISTORY:

H-1	This procedure was written by N. Kriz on 12-23-2004.

H-2	This procedure was reviewed by N Kriz on 10/12/2010.

H-3	This procedure was reviewed by N Kriz on 2/22/2011.

H-4	This procedure was reviewed by N Kriz on 2/6/2012.

H-5	This procedure was reviewed by N Kriz on 2/13/2013.