VerifyNow P2Y12 Test Procedure

## I.	Principle:

- The VerifyNow® P2Y12 Test is designed to measure platelet P2Y12 receptor blockade.  Substances known to specifically block the P2Y12 receptor include the thienopyridines class of drugs, including clopidogrel and prasugrel.

- The test is based upon the ability of activated platelets to bind fibrinogen.  Fibrinogen-coated microparticles aggregate in whole blood in proportion to the number of expressed platelet GP IIb/IIIa receptors.  The rate of microbead aggregation is more rapid and reproducible if platelets are activated, therefore the reagents adenosine-5-diphosphate and prostaglandin E1 (ADP/PGE1) are incorporated into the assay channel to induce platelet activation without fibrin formation.  The reagent is formulated to specifically measure P2Y12-mediated platelet aggregation.  Light transmittance increases as activated platelets bind and aggregate with fibrinogen-coated beads.  The instrument measures this change in optical signal and reports results in P2Y12 Reaction Units (PRU).

- In a like manner, a second activator, iso-TRAP (Thrombin Receptor Activating Peptide) and fibrinogen-coated microparticles are incorporated into a second channel of the Test Device. Aggregation is induced via thrombin receptors; this estimates total possible platelet aggregation regardless of the presence of a P2Y12 receptor blocker. The instrument measures the change in transmittance in this channel, calculates the baseline platelet function for the sample, and reports the percent inhibition result for the sample.

II.	Materials Required:

- VerifyNow Instrument

- VerifyNow P2Y12 Test Device (reagent device), box of 25 Test Devices (PN: 85054) or box of 10 Test Devices (PN: 85054-10). Store at room temperature

- Greiner Bio-One Vacuette 2.0 mL partial fill collection tubes containing 3.2% sodium citrate, Greiner Catalog Number 454321/454322.

- Phlebotomy supplies, including needle of 21 gauge or larger

- Electronic Quality Control Device provided with instrument

- Wet Quality Control (WQC), P/N 85047, Box of 6 Diluent Tubes and 6 Pellets. QC materials are stable at room temperature until the expiration date indicated on the tube or pellet container.

# III.	SPECIMEN COLLECTION AND SPECIFICATIONS:

A.   Patient Preparation: None required

1.	Specimen Type: Whole blood samples must be collected in or immediately

transferred to Greiner 2.0 mL partial fill blue top tubes containing 3.2% Sodium

Citrate.  The tube must be filled to its intended whole blood capacity (indicated by

small black line).

B.    Instructions for Sample Collection by Venipuncture:

- Identify a site that is free of any peripheral venous infusions.

- Apply the tourniquet and identify the phlebotomy site.

- Release the tourniquet and cleanse the phlebotomy area.

- Reapply the tourniquet.   It is important to minimize the amount of time that the tourniquet is applied prior to filling the tube in order to avoid activating the platelets while obtaining the sample.

- Perform the phlebotomy and obtain the appropriate volume of blood in the sample tube following a discard tube of at least 2 mL of whole blood.  DO NOT DRAW A SAMPLE FOR ANY PLATELET FUNCTION ASSAY AFTER A TUBE THAT CONTAINS EDTA (PURPLE TOP).  Always draw the tube(s) for the VerifyNow Test first.

- Gently invert the citrated tube containing whole blood 5 times immediately after collection to mix the blood with the anticoagulant and prevent clotting.

- Samples must be kept at room temperature.  Do not refrigerate, freeze or centrifuge the sample to be used for the VerifyNow Test.

- Samples that are difficult to obtain may hemolyze or clot.  Samples that are hemolyzed or clotted should be recollected.

C.   Instruction for Sample Collection from Indwelling Catheters:

- Ensure that the indwelling catheter line is free of clots.

- Obtain a discard of at least 5 mL in order to clear the line.

- Obtain the sample for testing with a second syringe.

- Immediately transfer the blood to the appropriate collection tube.

- Immediately gently invert the citrated tube containing whole blood at least 5 times to mix the blood with the anticoagulant and prevent clotting.

- Samples must be kept at room temperature.  Do not refrigerate, freeze or centrifuge the samples to be used for the VerifyNow Test.

- Samples that are difficult to obtain may hemolyze or clot.  Samples that are hemolyzed or clotted should be recollected.

D.     Sample Stability and Storage: The sample must be incubated at room temperature

for 10 minutes prior to testing and can be used up to 4 hours after collection if stored

at room temperature (18 to 25 C).  Do not refrigerate or freeze specimen.

E.     Sample Precautions: Collection of the blood specimen should be performed with

care to avoid hemolysis or contamination by tissue fluids.  Samples with evidence of

clotting must not be used.

F.     Interfering Substances:

- Glycoprotein IIb/IIIa inhibitors (abciximab, eptifibatide, and tirofiban) and antiplatelet agents (cilastazol) inhibit platelet function.  Some degree of platelet inhibition by these drugs was detected by the VerifyNow P2Y12 Test.

- Drugs that affect platelet function may be detected up to 14 days after administration.

- Other classes of commonly used drugs were tested with no significant effect on VerifyNow P2Y12 Test performance (antioxidants, ACE inhibitors, antiarrhythmics, anticoagulants, aspirin, antidepressants, insulin, allopurinol, alcohol, betablockers, bronchodilators, calcium channel blockers, gastrointestinal medications, betamethasone, lovastatin, NSAIDs [including COX-1 and COX-2 enzyme inhibitors], and the thyroid hormone L-thyroxine).  The thrombolytic agent streptokinase showed no significant inhibition of platelet function, as measured by the VerifyNow P2Y12 Test.

- Assay performance was not affected by hematocrit values between 33-52%, or platelet count values between 119,000 -502,000/µL.

- No significant interference was observed on samples studied with triglyceride concentrations between 41-824 mg/dL.  No significant interference was observed on samples studied with cholesterol concentrations between 98-316 mg/dL.

- No assay interference was observed when samples with fibrinogen levels between 171 and 599 mg/dL were tested with the VerifyNow P2Y12 Test.

IV.	Handling Conditions:

A.	The sample must be incubated for 10 minutes at room temperature prior to testing.  If testing is delayed, the sample can be stored at room temperature (18 to 25 C) for up to 4 hours.  Do not refrigerate or freeze specimen. Do not transport the specimen by a pneumatic tube system. The sample must be handled in accordance with all policies and procedures relating to the collection, processing and disposal of biohazardous samples.

V.	Reagent Storage and Handling:

A.   Reagents Required:

- All required reagents are contained within the individually packaged P2Y12 Test Device.  Each Test Device contains lyophilized fibrinogen-coated beads, ADP, iso-TRAP, bovine serum albumin, PGE1, and buffer.

- Each individually sealed Test Device contains the lot number and expiration date stamped on the foil pouch.

- Once removed from its foil pouch, the Test Device must be handled only by the finger grip and used immediately.

B.    Receipt of Shipment:

- When a shipment of Test kits is opened, check the temperature indicator located on the outside of the box.  If the temperature indicator is activated, this indicates exposure to elevated temperatures and a VerifyNow Level 2 WQC should be performed.

- If the Level 2 WQC result does not fall within the accepted range on the package insert, call Accumetrics Customer Support at 1(800) 643-1640.

C.   Storage Temperature.

- Store Test Devices at 2° C to 25° C (36° - 77° F).

- If refrigerated, allow Test Devices to reach room temperature, 18° C to 25° C (64° -77° F), prior to use.

- Test Device should remain sealed in the foil pouch until ready for use to prevent damage by humidity.

D.     Reagent Quality Control:

- The manufacturer recommends a Level 2 Quality Control procedure be run once each time a new lot or a new shipment of Test kits is received.

- Level 1 and Level 2 Quality Control Test Devices must also be performed at the frequency specified.  Please refer to the Quality Control section below for additional instructions.

VI.	Calibration and Calibration Verification:

- VerifyNow P2Y12 Test Devices are calibrated by the manufacturer at the factory. This calibration information is contained in the barcode on the pouch of each Test Device.

- The barcode must be scanned whenever a new lot of Test Devices is to be tested. The system will not allow a Test Device to proceed without the lot number calibration information being scanned into the system.  If a new lot of Test Devices is being used, the instrument will prompt the user by displaying a barcode icon after the Test Device is inserted.

  - At prompt, place the Test Device pouch approximately one inch in front of the barcode reader found on the left side of the instrument, so that the light shines on the center of the barcode reader.

  - An audible beep will be heard when the instrument receives the required information.

  - The user needs only to perform this action once per lot.

- No additional calibration is performed by the user.

- Calibration Verification is performed by the use of Wet Quality Control materials with every new lot of reagent and at specified time intervals (refer to Quality Control Procedures below), with periodic review of QC results by the laboratory technical supervisor.  Assays of platelet function are non-linear, and no additional calibration verification is required.

VII.	Quality Control Procedures:

A.    Electronic Quality Control.  The Electronic Quality Control Device (EQC) is supplied with each instrument and must be run each day of use prior to reporting patient results.  This device verifies instrument optics, reagent mixing and instrument pneumatics.  Refer to the VerifyNow System User’s Manual for additional information.

- The EQC Device is located in a storage port on the right side of the instrument.

- Choose QC by pressing the QC button from the main screen.

- Open instrument cover.  Remove the EQC Device from the storage port, open instrument cover and insert the EQC Device into the assay port of the instrument.  Close instrument cover.  The EQC will automatically run.

- Open instrument cover and remove the EQC Device and return to the storage port.  Close instrument cover.  The EQC result displays after the EQC Device is removed.

- If the EQC Device result is “FAIL”, repeat the EQC.  If the result is “PASS”, continue with patient testing.  If the EQC result is “FAIL” after the second analysis, use the cleaning cartridge as described in the Maintenance Section below and rerun the EQC.

- Record and print the result, return to the Main Screen.

B.     Wet Quality Control.   The Wet Quality Control (WQC) procedure must be

performed whenever a problem is suspected with the temperature indicator of a

newly received lot of Test Devices, whenever a problem is suspected with the

VerifyNow System.

C.     Sample Preparation:

- All control material should be stored at room temperature (18 to 30 C).

- Do not open the vial containing the Level 2 Control pellet until immediately prior to use.

- Assay Control Level 1 is ready for use as provided.

D.     Wet Quality Control:

- Press the QC Icon key.

- The Insert Cartridge screen will display.

- Open the foil pouch and remove the Test Device.  The device should only be handled by the finger grip.

- Remove the protective sheath from the device needle by pulling directly up on the sheath.  Do not twist the sheath as this may remove the needle.

- Open instrument cover.  Insert the Test Device at the instrument prompt.  If this is a new device lot, the Bar Code prompt will display.  At prompt, place the Test Device pouch approximately one inch in front of the barcode reader found on the left side of the instrument, so that the light shines on the center of the barcode reader.  An audible beep will signal that the instrument has read the bar code, and the test will continue.

- Prepare the WQC Sample.

- Level 2: Remove the stopper from the tube containing the Level 2 Control diluent by twisting and pulling the cap simultaneously.

- Add the Level 2 Control pellet to the Level 2 control diluent tube and replace the stopper by pressing and turning simultaneously.

- Invert the tube gently 5 times to mix.

- Level 1: Control Level 1 is ready for use as provided.

7.   At the prompt, insert the diluent tube onto the device needle.  Close instrument

cover. The testing will automatically begin.  CAUTION:  The sample is under

pressure once it is inserted onto the device needle.  DO NOT REMOVE the Test

Device or control tube from the instrument until the assay is completed.

8.   The instrument will run the test and display the result.

- Record and print the result and return to the Main Screen.

10. Open instrument cover.  Remove the Test Device into the device port by

grasping the device finger grip and pulling straight up.  Do not remove the

tube from the Test Device.  Close instrument cover.

11. Discard the used Test Device and quality control tube as biohazardous waste.

12.  The instrument is ready to test the next sample.

13.  Determine that the WQC result is within the acceptable range of values

printed on the Test Device pouch provided with the test.

14.  If the WQC is in control, proceed with the testing of patient samples.  If the

WQC is out of range, refer to users manual (section 5: Quality Control) for

recommendations as to how to proceed.  Document all actions taken in

response to the “out of range” condition in the instrument appropriate log.

E.    System (Internal) Quality Control for Each Sample Tested:

- Each time a Test Device is run on the VerifyNow Instrument, the instrument verifies the Test Device expiration date, sample filling, optics performance, correct fluid transfer, and proper mixing.

- The system controls prevent the operator from running an expired Test Device.

- The system also detects certain other operator errors, such as placing the Test Device or the sample in the instrument at the wrong time, or removing the Test Device before the testing is complete.  These controls prevent reporting of an inaccurate test result.

- The Test Device internal controls in VerifyNow P2Y12 Test can detect failures of the reagent system due to improper storage or handling conditions.

- The internal controls will flag an improperly collected or mishandled blood sample, or a blood sample with certain types of interfering substances.

- The Test Device internal controls detect errors from the reagent system, adverse environmental conditions, and additional types of operator errors.

##### F.    Recall of EQC or WQC Results:

- The last 100 EQC and WQC results can be recalled at any time.  You may scroll by using the arrow keys on the keypad.

- Press the Maintenance button from the Main Screen.

- Press the Next arrow 3 times to advance to the Maintenance Submenu where folders are displayed with the words EQC (2nd button) and WQC (3rd button).

- Press the appropriate button to recall the EQC or WQC results.

- Press the back arrow to return to the Main Screen.

VIII.	Testing Patient Samples:

- Refer to the VerifyNow System User Manual for complete operating instructions.

## Power on the instrument.  This will initiate the following startup checks:

- A system program and data memory check to ensure memory integrity;

- A system temperature check to ensure the test warming plate reaches and maintains the proper temperature;

- A system check of proper operating voltages; and

- A system intra-communication validation.

- Perform Electronic Quality Control if one has not been performed within the required timeframe.  The following checks are performed:

- Instrument optics.

- Pneumatics system that draws the sample into the Test Device for reaction and measurement.

- Reagent mixing parameters and sample data acquisition.

- Correct calibration parameters.

- Enter the specimen accession number under patient ID.

- Open the foil pouch and remove the Test Device. Test Devices should only be handled by finger grip.

- Verify that the humidity sensor is not the same color on both sides.

- Remove the needle’s protective sheath by pulling directly up on the sheath. Do not twist the sheath as this may remove the needle.

- Open instrument cover.  Insert the Test Device at the instrument prompt.  If this is a new Test Devices lot, the Bar Code prompt will display.  At prompt, place the Test Device pouch approximately one inch in front of the barcode reader found on the left side of the instrument, so that the light shines on the center of the barcode reader.  An audible beep will signal that the instrument has read the bar code, and the testing will continue.

- At the instrument prompt, invert the sample tube at least 5 times, and insert onto the needle in the Test Device. Close instrument cover.  CAUTION: Sample is under pressure. Do not remove sample tube from Test Device. Only remove the Test Device from the instrument after testing is completed.

- The instrument will run the test and display the result in less than three minutes.

- Record and print the sample result.

- Enter patient results in SoftLab, Resulting Worklist template: VFYNW.

- Open instrument cover.  Remove the Test Device by grasping the device finger grip and pulling straight up.  Do not remove the tube from the Test Device.  Close instrument cover.

- Dispose of the entire Test Device/sample tube in appropriate biohazard waste container.

- The instrument is ready to test the next sample.

## Recall Patient Results:

## In order to enable the recall of specific patient results, the instrument must be configured to require a Patient ID to be entered before each test is performed.

- If required, enter Operator ID and Password.

# Choose the second option (Button next to file folder) from the Main Screen

- Enter the Patient ID, and press the arrow for Next.

- A P2Y12 test result is indicated by a “y” and a P2Y12 baseline result by an “x”, and the % inhibition indicated by %.

- Record or print the results.  The most recent result is displayed, along with the date stamp indicating the date and time the test was performed.  To toggle to other results, use the left and right arrow key.

- If you choose the PRINT ALL icon, everything displayed below the blinking cursor will print.

X.	Reported Results and Calculations:

The VerifyNow P2Y12 Test no longer reports three results for each test. (Mandatory software change by the manufacturer). Only the PRU is reported:

P2Y12 Reaction Units (PRU) report the amount of P2Y12 receptor mediated aggregation specific to the platelet and are calculated as a function of the rate and extent of platelet aggregation in the ADP channel.

XI.	Assay Limitations and Method Notes:

A.	To minimize problems during specimen handling, test performance and reporting of test results, the area where testing is performed must contain the proper workbench space, ventilation, utilities, and supplies necessary for conducting the type and volume of testing performed.

B.	Place the VerifyNow Instrument on a clean, firm, level bench top, which is free of excessive vibration from equipment such as a centrifuge. Provide adequate space around the instrument to access instrument components and be sure the area is free from exposure to unusual temperature fluctuations.

C.	Do not locate the VerifyNow instrument in an area next to a source of heat, air conditioning or in direct sunlight.  Do not place the instrument under an incandescent light source.

D.	The VerifyNow Instrument operates at ambient temperature (18-32°C or 64-90°F) and up to 85% humidity without condensation.

## The lyophilized agent is hygroscopic and can degrade after prolonged exposure to room air. Therefore, the Test Device should be used shortly after removal from the foil pouch. Test Device Humidity Indicator for All Tests:  Each Test Device contains an internal humidity sensor.  If the Test Device has been degraded due to exposure to excessive humidity, a color change of the humidity indicator will be observed. The user is instructed to check the color of the indicator before each test, and to discard a Test Device where a color change is evident.  The Test Device Humidity Indicator detects errors due to adverse environmental conditions. Because of the built-in humidity indicator on each VerifyNow System Test Device, it is not necessary to record the humidity of the room on a daily basis.

## Store reagents and quality control material according to the package directions.  The Test Devices can be stored either at room temperature or in the refrigerator.

## Delays in testing or difficulty of specimen collection may result in spurious values. Do not test any sample that is clotted, too old, hemolyzed or that has been mishandled or mislabeled.

## When results are not with the expected limits, the possibility of improper sample collection or handling should be investigated. Repeat the test using a new Test Device and sample.

## Patients with inherited platelet disorders such as von Willebrand Factor Deficiency, Glanzmann Thrombasthenia and Bernard-Soulier Syndrome have not been studied with the VerifyNow P2Y12 Test. The VerifyNow P2Y12 Test is not intended for use with these types of platelet disorders.

## Patients with a known history of platelet counts <100 x109/L have not been studied.

## Do not test patients who have taken a GP llb/llla inhibitor (tirofiban (Aggrastat®) or eptifibatide (Integrilin®) within two days, or abciximab (ReoPro®) within two weeks.

## The VerifyNow P2Y12 Test results should be interpreted in conjunction with other clinical and laboratory data available to the clinician.

XII.	Diagnostic Error Display Messages:

A.	The instrument displays informational, error and alarm messages.  Refer to Chapter 8, troubleshooting, of the User Manual for additional explanation and information.

## XIII.	Reportable Range:

A.	The P2Y12 results are reported by a proprietary algorithm.  Prior to displaying the result, the software confirms that the result is within the reportable range.  Any numerical value that is displayed is reportable, and will be within the range of 1 – 999.  Any result that is outside of the reportable range will be displayed as an error, and no numerical value will be reported.

B.	Reference Range:

P2Y12 Test Result Baseline: 194 – 418

## C. 	Maintenance

- The Maintenance and Setup Icon is used to setup users, change the date and time, recall QC results, network the instrument, set patient ID requirements and for troubleshooting.  This function should only be performed by users with proper training and authority.  Refer to the users guide for additional information.

- Keep the surface of the instrument clean with any commonly used laboratory disinfectant.  Accumetrics recommends that the exterior surface be cleaned no less than monthly.

- Cleaning Device:  Twice a month, use a cleaning device or if the EQC Device fails twice.  REMOVE THE TAPE FROM THE CLEANING DEVICE AS INDICATED.  Place the cleaning device in the assay port for 5 seconds and remove.  Dispose of the single use cleaning device.  Rerun the EQC Device.  If the EQC does not pass, call Accumetrics Technical Service at 1-800-643-1640.

- If the EQC Device becomes contaminated, wipe with a damp cloth.

- Replace the fan filter yearly (see Section 10.2 of the User Manual for additional information).

- Other than fuse replacement, the external power cord, or parts discussed above no other user-serviceable parts exist on the VerifyNow Instrument. Contact Accumetrics Customer Support at 1-800-643-1640 if further service is required.

XIV.	References:

- VerifyNow Test WQC, Package Insert, P/N 14349

- VerifyNow P2Y12 Test, Package Insert, P/N 14317

### VerifyNow System User Manual, P/N 14340

XV.	HISTORY:

H-1	This procedure was written by N Kriz 10/18/2010.

H-2	This procedure was reviewed by N Kriz 2/24/2011.

H-3	This procedure was reviewed by N Kriz 2/6/2012.

H-4	This procedure was revised by N Kriz 2/13/2013.