Procedure Change/Revision (IMM 36-A)

Procedure:  Kappa and Lambda Free Light Chains  (Binding Site) on the  Siemens BNII Nephelometer (Doc# IMM 199)

Change/Revision: In the Interpretation of Results Section, Part A Reporting Results

The following change was made regarding retesting of New patients 			with abnormal results:

- If the patient does not have any results in Test History, and the Kappa, Lambda or Kappa/Lambda Ratio is out of the normal range, repeat the testing at a 1:400 dilution.

- Go to Lab Journal

- Open the sample by double clicking on the Order#

- Click Dilution

- Click on 1:400, then OK

- Reload the sample on the BNII

If the 1:400 result is less than 5 times 1:100 result, report the 1:100

If the 1:400 result is greater than 5 times 1:100 result, report the 1:400.

Date of Revision: 2/20/13

Director Signature/Date:

Technologist Signature/Date: SEE MTS