- Purpose:

To define the Hematology laboratory staff procedure to follow during scheduled and “crash” LIS computer downtime to maintain timely specimen processing and reporting.

- Policy:

LIS computer downtime (scheduled and crashes) must remain silent to the

Medical/hospital staff.  Specimen processing and reporting must continue uninterrupted regardless of computer support functions.

- Downtime Procedure:

A.  	Scheduled Downtime

- Prior to downtime - call pending reports for all areas

B.  	LIS Computer Crash

- Notify computer operations (LCODO beeper 688-1619), may have to hit #(pound) key if no response.

- If no resolution is made within 30 minutes, Computer Operations will invoke an official SCC downtime

- Notify Lab Manager if Computer Operations invokes an official SCC downtime

- Downtime < 1 hour:

-Only stat specimens will be processed

- Downtime > 1 hour:

-Hourly updates will be provided via the paging system

-Prioritized routine samples will begin to be processed along with stat

samples

C.  	Quality Control handling

- QC is to be run before any patient samples are run

- Perform necessary QC

- CBC: Run patient controls and commercial controls per usual; confirm results are acceptable before running patient samples.

- Urinalysis: Printout QC results from Atlas and compare with acceptable ranges in QC Lot # manual.  Record confirmatory QC results on Atlas printout.

- Coagulation: Look at QC results from BCS Easy Q libraries and confirm that results are acceptable.  Rerun QC, if necessary; do not run patient samples unless QC is acceptable.

- Fluids: Perform QC counts and record on log sheet and compare with acceptable ranges in QC Lot # manual

D.  	Specimen Processing (for both scheduled and crashes)

- Samples will arrive at the Pick-up bench.  Samples with a Soft label that have not been collected/received will be in racks labels “not collected/received.”  Samples with downtime labels will arrive with the patients name, ward and tests ordered hand written on the label. These samples will have #99 numbers. Our copy of the requisition will be filed alphabetically at front bench. For extra labels, use Downtime Link : http://flask.ynhh.org.cgi-bin/labeldup/labeldup.pl

E.  	Samples to be Faxed or Called

- Stat samples, ER’s, ICU’s, Medonc, HMISM and Transplant will need to have downtime result sheets filled out.  Designated supervisor or tech will fax instrument printouts or Downtime sheets to Stat locations.

- Samples with critical values will need to be called to floors, regardless of priority.  Write who was spoken to, the time and techs initials on instrument printout to be entered into Soft when the system is back up.

Sapphire

-Bloods that were run on the Sapphires should remain in racks in front of

instrument to keep track of bloods until system is back up. Tube Tracker

or green Coag racks can be used if Downtime is extended or volume of

samples is high.

- Run all stat samples on one instrument. Chartable printouts called on Stat CBCs and Downtime labels will be placed on printouts to be faxed to Stat locations. Demographic info transcribed on labels.

- Run all routine samples on the other 2 instruments if Downtime > 1 hour.

- Review instrument computer for review data/flags to qualify if sample to be repeated, smears made, other tests indicated.

- If sample is QNS or clotted; print an instrument printout, notify sending location for repeat and record on Sapphire printout for later canceling and comments.

- If sample is normal, keep in instrument racks until system is back up.

- If sample data needs to be reviewed:  Print 2 Sapphire printouts

-Write what needs to be reviewed on each printout

-Forward one printout to misc along with the blood and keep the other at the Sapphire in alphabetical file.

- If sample needs to be called due to critical values or stat sample be sure to write on the printout who the results were called to, the time and the initials of the person making the phone call.  Write this on the printout to be kept at the Sapphire.

  - Check labels for ESR, RETIC and BSINT

Miscellaneous

- Make smears and perform other tests as indicated on the requisition

-Use ESR log sheet to record ESR controls and ESR results. Use

Manual HCT log to record Spun HCT, buffy coats and QBC results

- Give bloods back to person who’s instrument it was ran on

- Give Sapphire printouts to the diff area

Diff Area

- Perform scans/smears/differentials/platelet counts/WBC counts, NRBC count or review as noted on the printout on the manual counters

- Record results on printout (include manual diff and morphology, for scan write “ok” next to autodiff)

-Note morphology on printout:

A – Abnormal, M – Marked Abnormal and/or P – Present

- Make any necessary calls for critical data (who the results were called to, the time and the initials of the person making the phone call.  Write this on the printout)

- File Sapphire printouts in alphabetical order in a folder to be kept in the diff area

Coagulation

- Check for QNS samples.  If sample is QNS, log on a sheet of paper the Soft number or downtime number and patient name, who the result was given to, the time and the initials of the person making the phone call.

- Spin samples

- Run samples on BCS

-If a Soft label is on the tube the BCS will run what was ordered

-If a downtime label is on the tube the BCS needs to be programmed

to run what’s indicated on the label

-Follow protocol for what needs to be repeated

- Be sure to check samples for a clot when appropriate; if clotted, notify sending location for repeat and record on log sheet for later calling and comments.

- Call critical values.  Record who the results were called to, the time and the intials of the person making the phone call on instrument printout.

- After samples are run on the BCS, they should be racked in a green rack by each BCS in alphabetical order.

Urinalysis

- Run all samples in small batches on the Atlas

- Review results from instrument printout to determine if confirmatory testing needs to be performed ( > 7.0 pH do SSA, ICTO, >= 1.035  specific gravity do Clinitest and/or Microscopic Analysis).

- Record confirmatory tests on clipboard

- Spin urines that need to have a microscopic performed

- Keep unspun urines that have been run in a separate rack

- Perform microscopic and record results on manual urine worksheet using macroscopic results from Atlas ( cut printout and attach to urine worksheet with Downtime labels.)

Fluids

- Follow fluid procedure for chamber counts and cyto smears

- Use hand clickers for chamber counts and differentials

- Record results and calculations as usual on log sheets.  Put counts and differentials on slips.

- Call Stats and/or critical results to LIP.  Record who the results were called to, the time and the intials of the person making the phone call on slips.

- Use specialized fluid forms and Downtime labels for Stat samples to fax.

F.  	Recovery after Downtime (for both scheduled and crashes)

All areas- All instruments will need to restart interfaces (Follow Downtime Procedure) before data can be released.

- Samples that arrived at the pick-up bench with a Soft label and in the “not collected/received” racks will need to be collected/received by the LAs before results released.

-Go into PhlebotomyReceiving and scan in samples with Soft label.  Click Save when done. (This can be done by technologist in the event, LAs are not available)

- Samples that arrived at the pick-up bench with a downtime label will be accessioned at the front end.

-Once the sample is accessioned it will appear on the appropriate pending list(s)

- Enter QC results into Soft Total QC

- Make sure all samples from pendings called prior to downtime have been resulted.

- Follow up on any missing samples.

Sapphire: Recovery after Downtime:

- Supervisor or Lead tech should restart all Sapphire interfaces (including YPB.)

- Look at samples in racks in front of Sapphires - determine what first sample number was run when first went down and see what the Sequence number is.

- Retransmit sequence numbers from the first sample to the last sample run while down.

From the main sapphire screen:

-Click Transmit

-Enter Sequence from:  to:  (Enter first downtime number in from and the last downtime number run in to)

-Click Transmit Data Log Records for Selected Sequence

- Look for patient names on the pending that came with a downtime label

-When the sample is accessioned, it will appear on the pending list

(make sure to open the pending list using the downtime label number range, 990000)

-In instrument menu, patient demographics will appear with the order

and the TAT status column will no longer be white

- Samples that did not autoverify will be pending in the Instrument Menu

- Verify samples from the instrument based on Sapphire printouts( Note: Data from Sapphire needs to be released first, before manual diffs can be entered.)

- Enter called to comments for Stats and critical values called to LIP

- Cancel QNS, clotted samples and enter “Called to…” comments

Miscellaneous: Recovery after Downtime:

- Enter ESR results

- Call pending reports to look for missing samples

Diff Area: Recovery after Downtime:  Note: Do not result differentials until samples are verified from the Sapphire

- Enter results of tests performed (if entering results for another tech, enter internal comment with their initials)

- Be sure to change any scans that needed a differential from ScanManual Diff for test SCOPE

- Call pending report for diff area to assure work is complete

Coagulation: Recovery after Downtime:

- BCS interfaces will need to be restarted by a supervisor or lead tech only

- Retransmit samples with Soft numbers and downtime numbers from each BCS:

-Click Setup

-Choose all samples under Sample Selection

-Choose numeric under Sort Samples By

-Click close

-Highlight test results that are Dark Green (Click on

result and box will highlight in bright green).

-Click Send (Sample results will turn light green when

retransmitted)

- Look for patient names on the pending that came with a downtime label

-When the sample is accessioned, it will appear on the pending list (make sure to open the pending list using the downtime label number range)

-In instrument menu, patient demographics will appear with the order

and the TAT

status column will no longer be white

- Samples that did not autoverify will be pending in the Instrument menu

- Evaluate/verify samples which did not autoverify

- Enter called to comments for Stats and critical values called to LIP

- Cancel QNS, clotted samples and enter “Called to…” comments

- Call pending reports to assure work is complete

Urinalysis: Recovery after Downtime:

- Supervisor or Lead Tech should restart Atlases and Clinitek interfaces.

- Retransmit results from Atlas

Press:

-Main Menu

-Recall Results

-#1 – Routine Test Results

-#1 – Seq # from > Seq # to (seq. # of first sample run to seq. # of

last sample run)

-Search

-Continue

-Send

-All Results

-Computer

- Results will autoverify when transmitted

- Look for patient names on the pending that came with a downtime label

-When the sample is accessioned, it will appear on the pending list (make sure to open the pending list using the downtime label number range)

-In instrument menu, patient demographics will appear with the order and the TAT status column will no longer be white

- Result confirmatory tests performed

- Double check for additional and/or confirmatory testing

- Enter microscopic results from manual urine worksheet

- Call pending reports to assure work is complete

Fluids: Recovery after Downtime:

- Enter cell counts and differentials

- Enter called to comments for Stats and critical values called to LIP

All areas:

- Make sure all samples from pending reports printed prior to a scheduled downtime are resulted

- Follow up on any missing samples and record on pending report

- Sign-off pending report

IV.	HISTORY:

H-1	This procedure was written by Jessica Kaul on 6/27/2011.

H-2	This procedure was revised by Donna Fico on 3/20/2012.

H-3      This procedure was revised by Donna Fico on 1/2013.

Appendix A - Coagulation

Appendix A – H-01-019 Procedure/Policy for Computer Downtime

Appendix B – Urinalysis

Appendix B – H-01-019 Procedure/Policy for Computer Downtime

Appendix C – Fluids

Appendix C – H-01-019 Procedure/Policy for Computer Downtime