If a problem is perceived:

- Take a visual inspection of the inside of the Sapphire check for the following:

- Any leaks

- Look for clots in tubing

- Cracked, or crimped tubing or tubing that may have popped off, or visible leaking

- Check peripump tubing is not flat

- Check for any crusting on syringes

- Check for wet caps

- Look for air bubbles in syringes, are the syringes emptying and draining properly

- Check that syringes are properly installed

- Check the volume of reagent kegs ,crusting on top of the caps, make sure the hole in cap is clear and open

- Check the reagent log if lot number has changed, note the lot numbers

- Dilution cups are not overflowing, check for clots

- Waste Chambers are draining and filling properly

- Check for any obstructions in staging pathway

- Check the maintenance log for a recent maintenance, calibration, or part change

- Are there any unusual sounds during the run cycle?

- Characterize the problem:

  - Does the problem occur with one specimen? Verify that it’s not a problem specific to the sample

- Instrument hardware (ex. Instrument sampler doesn’t move)

- Computer (ex. celldyn computer screen frozen)

- Instrument (ex. WBC mix pot overflowing)

- Reagent (ex. Reagent just changed and results incorrect)

- Unknown (ex. Multiple problems)

  - Instrument Hardware problem

- Is something occluding tubing, is a reagent cube sitting on a reagent line

- Call customer help line 1 800 235 5396 and follow prompts

  - Computer problems

- Bring system down and back up

- Celldyn problem vs lab computer (688-1619 computer operations beeper)

- Review celldyn screen- are all prompts on C screen up

  - Reagents

- Correct reagent, was reagent gently mixed

- All lines attached

- Try another cube

  - Unknown (ex. multiple problems)

- Decide on problem

- Check troubleshoot manual for fixes

- Call celldyn phone support

- Review finding

  - Review the following programs for any violation of CBC parameters

- Background counts - Is the nature of the problem background counts outside specifications, reproducibility, something else?

  - patient references 11 & 12

  - moving averages

Review commercial controls data points and Levey Jennings for any trends or jumps

Refer to cell-dyn matrix grid (a.k.a. cell-dyn sapphire parameter generation reference) for common problems.

- Which parameters are affected?

- Does the problem affect only one parameter?

- Does the problem affect only parameters that are measured using the same reagent

- Does the problem affect only parameters that are measured at the measurement site, i.e. Optical flow cell?

  - Data incorrect on open mode or closed mode, or both?

- Which mode is incorrect

- Check needle, sampling

- Check mixpots

- If  reagent(s) just changed (is it what  should be changed, is reagent line in correctly, is line full of reagent

- Look for clots in tubing, needle, fittings, pots

- Backflush problem line or pot

- Refer to celldyn matrix grid for common problems

- Documentation:

    - List Problem

    - All resolves – including any parts replaced

    - If fixed

    - Which supervisor alerted

    - Was service requested

    - Any troubleshooting suggested by phone support

    - Record problem in instrument log, include reference # from phone support

MCH = HGB x10

RBC

MCHC= HGB x100

HCT

MCV = Direct measurement

| Reviewed | Revised | Reviewer |
| --- | --- | --- |
| 12/2003 |  | N Ortoli-Drew |
| 3/2005 |  | N Ortoli-Drew |
|  | 5-2011 | N Ortoli-Drew |
|  | 3-1-2013 | Elizabeth Degano |

II.     HISTORY: