- P2Y12/ VerifyNow specimens will be delivered by 7:00 AM to the routine coagulation area by courier. DO NOT accept samples delivered via the pneumatic tube system, since this can lead to erroneous results from spurious activation of the platelets. Afternoon samples received before 2:00 PM will be processed at 2:00 PM. At the present time, no samples will be tested after 2:00 PM on the evening or night shifts.

- Routine coagulation technologist should accept and sign in the samples using the form provided on the clipboard in that area. When accessioned, attach labels to the sign-in form.

- DO NOT centrifuge the P2Y12 specimens, but deliver them to the special coagulation technologist.

- The special coagulation technologist will collect and receive each sample, referring to the VerifyNow Template which follows here (or in the Soft desk reference manual, or in the full P2Y12 procedure in the VerifyNow instrument notebooks). Collection time should be provided on the tubes.

- Allow samples to sit at room temperature for at least 10 minutes (and up to 4 hours after drawing) before testing. The time it takes to accession samples and ready the instrument(s) should be sufficient.

- Prepare instrument(s) according to directions in full VerifyNow procedure notebook. Do EQC daily and log in the appropriate section in the notebook.

- Test each sample as follows:

  - Enter specimen order number; press enter (top button); prompt to enter device into instrument will appear on screen.

  - Open foil pouch and take out the P2Y12 test device (handle only with finger grip). Remove the needle’s protective sheath by pulling directly up on it (do not twist)

  - Open instrument cover and insert the device

  - At the prompt, gently invert the sample tube at least 5 times, then insert tube carefully but firmly onto the needle in the test device.

  - Close the cover, and DO NOT OPEN and DO NOT REMOVE the device until the instrument has completed testing, and displays the icon to remove the device.

  - The test runs and the results are displayed within 3 minutes. Print the result.

  - Enter results in Soft, using template VFYNW or P2Y12 –there is only one result (PFP for PRU platelet reactive units) to be entered for each patient sample.

  - Open the cover and—using the finger grip only—remove the test device with the tube attached. Discard the whole unit into a biohazard bin.

  - Close the cover; instrument is ready to run next sample.

  - If you forget to print, and go out of that screen, results can be retrieved by pressing the “Folder” icon, then entering the order number of the sample. The results will display on the screen; press the print icon to print results.

- Accession a VerifyNow P2Y12 sample as follows:

Verify Now Training Template

Orderable Tests:				Resultable Tests:

VFYNW = Verify Now			PFP = Platelet Function P2Y12

VFYNW – Resulting Template for Verify Now

To Order Verify Now – P2Y12

- In Order Entry, type in Test ID VFYNW or P2Y12 and hit Enter

- Coll/Rec sample in Specimen tab

- Click  and click YES to “Do you want to save the changes you made to the order #?”

- Enter results for P2Y12 as follows:

To Verify Verify Now – P2Y12 Results:

- Open Resulting Worklist:  template – P2Y12

status -  Pend and Nonver

from order # - current orders

- Select order # for results to be entered

- Enter numerical result for PFP

- To verify and save results:

-Click Verify All and click YES to “Do you want to save the changes you made to the order #?”

- When the sample is complete there will be a blue checkmark and a pencil by the order #

i.	History:

H-1	This procedure was written by N Kriz on 11/26/2010.

H-2	This procedure was reviewed by N Kriz on 2/24/2011.

H-3	This procedure was reviewed by N Kriz on 2/13/2012.