- PRINCIPLE:

Routine hemoglobin electrophoresis is done at pH 8.6. Hemoglobins are negatively charged and migrate toward  	the positive electrode.  The main locations identified are  those for "Fast", "A", "S", A2 and C hemoglobins.  At pH 8.6 other hemoglobins may 		overlap these:  F moves close behind A; D and G if present, move with S; E and O, if 	present, move with C. The hemoglobins designated as F, D, G, E, and O can be separated from the others (A, S, A2, C) by changing the pH of the medium and electrophoresing at pH 6.2.

II.        SPECIMEN:

Hemolysate from EDTA specimen used for screen. Procedure for preparation of hemolysate included later in this procedure.

III.      REAGENTS:

A.	Citrate Buffer - 0.3 Ionic strength - dilute package to 1000 ml deionized water pH  6.0-6.3.  Stable 1 month in refrigerator.  Cat. #5121 Helena, 10 pkg./box.

B.	Sponge strips (EC Apparatus) - cut 3 strips to 6-1/2" long; then cut one in half lengthwise.

C.	Agar plates - each plate contains 1.5% agar in .05 m citrate buffer - stable for 1 year in refrigerator.  Cat. #2400 Helena - Box of 20.

D.	0.2% Diansidine in methanol (Cat. #5032 Helena) (0.5gm of Diansidine dilute with 250 ml of methanol) (diansidine in cabinet at room temp.) Stable for 6 months stored at room temperature.

E.	1.0% Sodium Nitroferricyanide in water (Cat. #5111 Helena). (1.0gm Sodium Nitroferricyanide in 100 ml deionized water) Stable for 1 year stored at room temperature.

F.	5% acetic acid (5 ml glacial acetic acid in 95 ml deionized H20) Stable for 1 year stored at room temperature.

G.	Hydrogen peroxide 3%, stable 3 months from open date stored at room temperature.

H.	Zip Zone applicator (Cat. #4080 Helena).

I.	Zip Zone sample well plate (Cat. #4081 Helena).

J.	Zip Zone aligning base for agar (Cat. #4083).

K.		Labels (Cat # 5006 Helena)

- 2% Glycerol (2 ml glycerol in 98 ml H2O) stable indefinitely stored at room temperature.

M.       Hemolysate reagent (Cat. #5125 Helena 250 ml/bottle)

- Whatman #3 filter paper for blotting

O.        Epson #4 Matte paper (Office Max  S041257)

IV.      CONTROL:

AFSC Hemoglobin Control (Cat. # 5331 Helena)

Normal Hb AA patient control

V.	PROCEDURE:

A.	Preparing electrophoresis chamber:

1.	Place sponge strip that was cut in half lengthwise on bottom of chamber and another strip on top making it perpendicular to the bottom sponge strip.  Do the same on the other side.

2.	Place citrate buffer equally in both chambers enough to soak the strips well and extra to fill the chamber two thirds full.  Place frozen sponge strips in middle chamber for cooling purposes.

3.	Connect the positive and negative leads of the electrophoresis chamber to 	the plugs in the front panel of the power supply.

4.	Turn on the power switch. A beep will sound and the instrument will run a self check. After approximately 2 seconds, dashes will appear on the displays indicating that the power supply is ready for use.

B.	Preparing Agar Plate:

1.	Remove Titan IV citrate agar plate from plastic bag and remove the tape and cover from the agar plate.

2.	Let plate sit at room temperature for a few minutes to let the moisture

evaporate.

3.	Place agar plate into aligning base so that the plate fits perfectly on it.

C. 	Preparing sample:

1.         Write a patient grid to refer to when adding patient samples and controls 			to the wells.

2.	Label 12 x 75 tubes for each sample and dispense 0.5 ml of hemolysate

reagent to each tube.  (CAP samples: 50µl of sample with 150µl of hem-

olysate reagent.)

3.        Add 90 µl from the red cell layer to the hemolysate reagent.

4.        	Cap tube, vortex and allow specimen to sit for 5 min.

5.	Place hemolysates of patients and controls in sample wells using the

pipette set to 8µl. – Use an AFSC and a normal patient control on every

plate.

6.          To prime the Zip Zone applicator, one should place applicator into sample             well plate for a few seconds, then wipe it off with a kimwipe; place applicator in again and blot on filter paper. This should give uniform                      application on the agar and also give some practice on handling the applicator.

7.	 Once the applicator is primed, place it in the aligning base, it should fit

into the slots on either ends, so that you can apply the samples twice on

the same agar plate if you wish.

8.	The applicator should then rest on the agar plate for 1 minute for the

samples to absorb uniformly.  Do not press down on gel.

  - Place agar plate on sponge wicks, placing agar side down and making sure

the plate is in contact with the sponge wicks.

10.	Press VOLTS enter 50 using the keyboard. When the correct voltage is 			displayed, press time.

11.	The TIME display will always show 10 minutes 00 seconds. Enter 60 			minutes using the keyboard.

12.	Allow the plate to equilibrate. Use a manual timer and wait for at least (5 				minutes) before starting the electrophoresis. This is necessary to ensure 				complete electrical contact.

13.	Press RUN/WAIT to start electrophoresis.  If an error code (E04) appears,

allow plates to equilibrate another minute and  press RUN/WAIT again.

14.	Electrophoresis will continue at the selected voltage and the time

programmed.  At the end of the run, a beep will sound and power will be

turned off.  The display will return to dashes only.

D.	Staining Procedure:

1.	In a flask prepare stain by combining:

10 ml 5% acetic acid

5.0 ml 0.2% Dianisidine

1.0 ml 1.0% Sodium nitroferricyanide

1.0 ml 3% hydrogen peroxide

2. 	Remove agar plate from cell and place it in a staining dish, agar side up.

3.	Puddle the stain over the entire surface of the plate and stain for 10

minutes.

4.	For permanent storage:

a.	Wash in 5% acetic acid for 30 minutes in staining dish.

b.  	Rinse in deionized water for 10 minutes in staining dish.

c.  	Hold the plastic plate under gently running water.

d.  	Cut the agar in half, invert agar, then slide a 3 x 5 piece of Epson 			#4 paper under the stained half of the agar and remove it from the 			plastic 	plate.

e.  	Flood the agar surface with 2% glycerol for 35 minutes.

f. 	Tilt and drain the paper onto a blotter for 2 minutes.

g.  	Lay the paper on a fresh blotter, then dry at 37 C overnight.

h.	Secure the ends of the Epson paper with the plastic racks in the incubator. Otherwise, the ends of the Epson paper will curl and the agar will crack.

i.	Label the agar results-date, FASC normal AA patient controls and

patients.

VI.      INTERPRETATION:

Reported by attending and resident, comparison of various patterns to known

mobility patterns and standards help to determine the hemoglobin type.

VII.  	REFERENCE:

1.	Technical Bulletin, "Titan IV Citrate Hemoglobin Procedure", Helena

Laboratories, Beaumont, Texas.

VIII. 	HISTORY:

H-1	This procedure was written by P. Morris on 10-08-97.

H-2      This procedure was revised by S. Richardson on 3/5/2012.