# I. 	PRINCIPLE:

Body fluids other than urine and blood are analyzed for enumeration and differentiation 	of cells. Cell counts are performed by standard chamber count technique. Hematocrits on 	grossly bloody fluids are performed following microhematocrit procedure. Differentials 	are performed on Wright stained cytocentrifuge or pulled smear preparations.

# II. 	SPECIMEN:

Body fluids other than blood and urine are classified under the general heading of Body 	Fluids. These include cerebrospinal, pleural, peritoneal, pericardial, peritoneal dialysis, 	peritoneal lavage (BAL), amniotic, synovial fluids, and drainage fluids (usually Jackson 	Pratt maybe cyst or abscess, etc). Fluids are processed regardless of the presence of clots; 	it must be noted in the report that clots are present in the specimen and cell count will be 	compromised. If the fluid is a solid clot, it will not be processed further and the floor 	notified. The laboratory should check if there are other specimen tubes (shared samples 	from Chemistry or Microbiology) that would be acceptable for analysis. Cerebrospinal 	fluid should be processed within eight hours of collection. Other fluids should be 	processed within twenty-four hours of collection.

# III.	REQUIRED MATERIALS:

A.	Gilson adjustable pipet (0-1000 ul and 20 - 200 ul) and pipet tips

B.	12 x 75 mm test tubes

C.	Neubauer counting chambers and coverslips

D.	Disposable plastic and glass pipets

E.	Disposable Cytocentrifuge chambers, Thermo Fisher Scientific

F.	3 x 1 inch glass slides

G.	Sterile saline - Change daily working tube

H.	22% Albumin

I.	Cytocentrifuge

J.	Hematek 2000 stainer

K.	Microscope

L.	Streck Cell-Chex levels 1 and 2

M.	General lab centrifuge

N.	Microhematocrit tubes/Microhematocrit centrifuge

# IV. 	QUALITY CONTROL:

## A. 	Daily in A.M

Note: 	both saline and albumin must be organism and particle free to be used; if 			first review is unacceptable (note in log), change reagent and review. If any 			problems still 	exist a supervisor must be informed and reagent not used.

### 1.	Saline:

a.	Fill a clean dry tube with saline

b.	Date tube

c.	Remove 1 drop of saline and place on a 3xl-inch glass slide and 					coverslip.

d.	Review prep under microscope for presence of organisms and 					particles.

e.	If prep is particle free and organism free, saline may be used; if 					particles or organisms are seen, discard saline and open new bottle 					of saline and repeat, review.

f.	Acceptable review is particle and organism free

g..	Document review in "saline/albumin" log.

### 2.	Albumin:

a.	Remove albumin from refrigerator.

b.	Place 1 drop of albumin on a 3xl-inch glass slide and coverslip.

c.	Review prep under microscope for presence of organisms and 				particles.

d.	If particles or organisms are seen, discard albumin and open a new 				bottle and repeat review.

e.	Acceptable review is particle and organism free.

f.	Document review in "saline/albumin" log

### 3. 	Streck Cell -Chex fluid controls (Level  l and 2) (stored in Fluid 					refrigerator. Unopened control material stable until outdate. Open vials 				stable for 30 days. Expiraton dates should be on control vials.

a.	Mix vials thoroughly by rolling between hands 20 - 30 seconds, 			then rapid inversion to resuspend and then gently invert 8 - 10 			times (does not need to come to room temp).

b.	Plate both sides of hemocytometer with level 1 and a second 			chamber with level 2.

c.	Count WBC's and RBC's on both sides of sides of chambers.

d.	Average the WBC and RBC counts and multiply by chamber count 		factor.

e.	Order fluid Q.C. in Soft and enter results. If counts are not within 			             posted ranges, replate and repeat counts. If still unacceptable,			             notify a supervisor immediately.

# V. 	CELL COUNT PROCEDURE:

## A. 	Mix fluid.

## B.	Observe fluid for presence of clots and abnormal color; note in computer if 			present.

## C.	Plate both sides of Neubauer counting chamber and allow to settle.

## D.	Examine chamber under low power to evaluate cellularity and distribution.

E.	Method of Counting Fluids:

1.	Clear or low cellular fluids: Count undiluted fluid

2.	Cellular (counts>400/cumm):

3.	Dilute with saline using the following table as a guide.

| Dilution | Fluid Volume | Saline Volume |
| --- | --- | --- |
| 1:2 | 0.2 ml | 0.2 ml |
| 1:5 | 0.1ml | 0.4 ml |
| 1:10 | 0. lml | 0.9 ml |
| 1:20 | 0.1ml | 1.9 ml |
| 1: 100 | 10 u1 | 0.99 ml |
| 1:200 | 10 ul | 1.99 ml |

F.	Alternate Methods for Counting bloody fluids and cellular joint fluids:

1.   	Glacial Acetic acid( See Appendix E) dilution should be done on bloody

Fluid when nucleated count is low. Brown bottle in fluid area 3% glacial

acetic acid. 1:20 – 50 ul  of fluid diluted with .950 ml of 3% glacial acetic

acid in separate tube.

- OR –

2.	Dilution using Methylene Blue instead of saline if low volume of fluid or 	nucleated cell count is low. \

3.	Dilution with 1:1 saline-Methylene blue solution can be used for cellular 	joint fluids. The Wbcs will stain pale blue.

NOTE: Since joint fluids tend to be very viscid, a 1:2 or 1:10 dilution should be 			made to facilitate counting and preparing the cyto smear.

## G.	Count and differentiate rbc from nucleated cells on 40X power on both sides of  			counting chamber and calculate the red and nucleated cell counts using the 			formula below. Agreement of counts is within or equal to 10%. If discrepant, fluid 		must be recounted.

H.	To determine what area of the counting chamber is counted, the large square on 			upper left hand corner is evaluated for cellularity.

I.	Counting cells:

| Cells/ruled chamber ( large square) | # Large squares to count |
| --- | --- |
| 0- 24 | 9 |
| 25_-99 | 4 |
| 100 - 200 | 1 |
| >200 | dilute or ten small center squares |

Formula used for undiluted and diluted fluids.

See counting chamber illustration next page

J.	Calculating chamber count factor:

1.	Depth x dilution x area counted (area corrected to 1 sq mm)

2.	Counting chamber calculation factors for undiluted fluid:

| # Large squares counted | Factor |
| --- | --- |
| 9 | 10/9 x count |
| 4 | 10/4 x count |
| 1 | 10/1 x count |
| 10 small squares | 25 x count |

Examples:

All 9 large squares counted:

9 RBCs x 10/9 = 10 cells/cu mm		RBCs

2 nucleated cells x 10/9 = 2 cells/cumm	nucleated cells

1:10 dilution with 10 small center squares counted:

95 RBCs x 25x 10 =23,750 cells/cumm		RBCs

108 nucleated cells x 25x10 =27,000 cell/cumm	Nucleated cells

## K. 	When counting, count cells within grid lines (top side and left side to avoid 			counting cells twice). Record both counts and calculations per each side on log 			sheet.

Note: Cells counted x calculation factor x dilution factor= cells/cu mm

# VI.	SMEAR PREPARATION:

## A.	 Smear preparation

### 1.	Cytocentrifuge Operation and Preparation:

a.             Remove centrifuge head and shield of CytoSpin 3 in order to

load slides and specimens. Position labeled (accession #) slides

into disposable cytofunnel and insert into centrifuge head.  Pipette

the sample into double-headed cytofunnel.

b.        Clear fluids- pipet  0.1ml of fluid in one side of double-headed 			            cytofunnel and 0.2 ml in other side.

c.	If tinged with blood or slightly cloudy, cyto smear is made 					according to the following: pipette 0.1 ml of fluid into labeled

tube add 0.5 ml of saline and mix.

d.  	For CSF, dialysate, and peritoneal fluids only, add 1 drop of 22% 				albumin to the chambers before adding fluids. Other fluids do not 				require albumin to “cushion” the cells.

e.         Place cover on centrifuge head and install head into centrifuge.

Close lid and press START on pad of cytocentrifuge. Spin 10

minutes at speed of 1100 RPM.                                                   .

- When centrifuge stops, remove centrifuge head  then fluid slides

and mark fluid spot on the reverse side of the slide.

g.  	Air dry slides and stain with Wright stain.

2.	Drop smears must be made if fluid is bloody or cellular or question yeast 				or bacteria.

a. 	Centrifuge specimen at 1200 RPM.

b. 	Remove supernatant to clean, labeled tube (use extra printer label 				and note on tube supernatant)

c. 	Resuspend sediment in 1 drop of supernatant.

d. 	Gently smear one drop of sediment using a pipette onto a 3 x 1 				inch slide.
			Cellular fluid prepare a slide from the sediment as you would make 			a peripheral smear.

e. 	Label slides with accession number

f.	Air dry slide completely and Wright stain.

3.	Hematocrits on bloody fluids- follow microhematocrit procedure in Main 				Lab manual p.H-39

# VII.  	NUCLEATED CELL DIFFERENTIAL:

## A.	 Cells are identified from Wright stained slide only:

Differentiation of nucleated cells is made using the following classification of 				cells.

Granulocytes

Lymphocytes

Monocytoid (CSF)

Tissue Cells (Body fluids- includes histiocytes and mesothelial cells)

Others - includes eosinophils, basophils, nucleated red cells and blasts.

## B. 	The number of cells counted in the differential is guided by the number of 				nucleated cells/cumm. The minimum number of cells in the differential should be:

| Nucleated cells/cumm | Cells in diff |
| --- | --- |
| 5 | 5 |
| 10 | 10 |
| 25 | 25 |
| >25 | 100 |

## C.	All differentials counts are reported as %  regardless if less than 100 cells are   				counted.

## VII.  	 	PROCEDURAL NOTES:

Process only one type of fluid per patient per 24 hours.

All fluids are checked for microorganisms as part of the analysis.

A.	Peritoneal Dialysate Fluids:

1.	No differentials are done on peritoneal dialysis fluids with nucleated cell 				counts <l00/cumm

2. 	Cytopreps and drop smears should be made and scanned for 					microoganisms

3. 	Only 1 dialysate fluid/day (midnight to midnight) will be analyzed.

B.	CSF Specimens:

1. 	Pedi hem/onc slides are checked for bacteria and reviewed for blasts and

abnormal cells by competent techs or supervisor. ( See policy for techs

passing annual fluid competency.)

2.	Bloody CSF specimens are always centrifuged to evaluated supernatant 				for xanthochromia and hemolysis. If either present, enter canned comment 			“sample hemolyzed” or “ sample xanthochromic” on patient specimen in

computer.

3. 	Presence of ghost cells- only count intact rbcs and put comment against 				rbc count, “cell count maybe inaccurate due to presence of ghost cells.”

4. 	If 2 tubes of CSF received, count lowest number tube first. If cell count 				within normal limits <5 rbc and <10 nucleated cells), the second tube is 				not counted. Report count and differential on the first tube, and "Cancel" 				 cell count on the second tube and comment why count cancelled..

C.	Ascites, Peritoneal and Dialysates:

1.	Absolute granulocyte counts (ANC) must be calculated for ascites fluids. 				Formula:

a.	 Total Wbc(nucleated count) x decimal form of polymorphonuclear 				cells(segs)= ANC.

Example:  Total WBC= 3,344 x .84( 84%) granulocytes(segs)= 					2880/ul(ANC), this is considered a critical value,  > or  = 250/ul. When 				fluid diff is less than 100 cells, % polys is determined by:  # of polys in 				diff divided by total # of cells counted in diff. That number in decimal 				form times by total Wbc= ANC.

D.	BAL (bronchial lavage):

1.	 Differential for BAL: if > 50% of cells on cyto smear are degenerated , 				differential is not reported and comment put into computer; “ unable to 				report diff. due to marked degenerated cells”.

E.        Wound drainage, abscess fluid, cyst fluid or other surgical drainages ex. (Jackson

Pratt)

- Chamber cell counts will not be performed. A spun Hematocrit will be done if bloody and reported. If fluid is clear, a spun Hct will be reported as < 2%.

- A pulled smear will be made, stained and examined for significant and/or abnormal findings. These findings will be commented against the spun Hct.

F.	  All fluid specimens:

a. 	Bacteria and fungi seen on cyto smear
				prepare drop or pull smear and review
				check computer to see if new finding.

(1.)	New finding:
				notify supervisor (or lab resident off shift)
				check to see if micro has paired specimen and results
				call caregiver and note in computer who and when notified
				leave for supervisor/resident director review

(2.)	Old finding:
				enter comment into computer
				do not leave for review

b.	Abnormal Cells/ Critical Findings( first time)

any fluid with blasts, malignant or "atypical" cells must be referred 				to supervisor directly and reviewed by lab resident/director prior to 				reporting. These abnormal smears are retained in some cases for

staff education.
				CSF- Nucleated cell count > 1000/ml must be referred to lab 					resident.

ANC  of > or = to 250/ul in ascites is a critical value- notify 					caregiver and/or lab resident if multiple times.

c.	All Body Fluids (other than CSF):
				Differentials are sent out by competent techs or supervisor.( Refer

to policy of guidelines to fluid reporting).

d.	Fluid Specimen Storage:
				    1. Specimens and cyto smears are retained for 7 days.
				    2. Completed fluids forwarded to other labs must be noted in

“ sign-out” log.
				    3.  All dilution tubes must be discarded after count completed.
				    4.  Label all remaining specimen tubes and identify as

supernatant, sediment or original specimen.

# IX. 	NORMAL RANGES:

CSF:

WBC < 6 cells/ul

RBC   < 5 cells/ul

Lymphs	0 – 10	(Child)

0 - 6	 (Adult)

See table for significance

A.	Pleural Fluid:

1. 	Normal volume: < 20 mI. (Normally not tapable.)

2.	Appearance:

a. 	Normal - clear, pale yellow

b. 	Chylous - milky, white

c.	 Inflammatory - cloudy, turbid (increased leukocytes)

3.	Differential on nucleated cells may be useful to differentiate between 				mycobacterium vs bacterial infection.

B.	Pericardial Fluid

1.	 Normal fluid: 10 - 50 ml

2. 	Appearance: interpretation similar to Pleural Fluid.

3. 	Differential on nucleated cells:

a. 	Over 1000 leukocytes/ul with a predominance of neutrophils 				suggests bacterial infection.

b.	Viral pericarditis, tuberculosis, and post-myocardial infarction 				syndrome usually show a predominance of lymphocytes.

C.	Peritoneal Fluid:

1.	Normal volume: < 50 ml

2.	 Appearance: Similar to Pleural Fluid. Green, bile-stained fluid is 					occasionally seen;
			Confirm with spot test for bilirubin (IctoTest).

3. 	Cell count and differential:

a.	 < 300 leukocytes/ul considered normal.

b. 	< 25% neutrophils is considered normal.

(1.)	Absolute granulocyte counts over 250/ul correlates 							strongly with bacterial peritonitis and is a critical value 						which needs to be communicated to a caregiver.

(2.)      High percentage of lymphocytes may suggest TB.

D.	Synovial Fluid:

1.	 Normal volume: <3.5 ml

2. 	Appearance: clear, yellow fluid.

3. 	Cell count and differential:

a. 	< 200 leukocytes/ul considered normal.

b. 	< 25% neutrophils considered normal.

c. 	Crystals: None is normal (See "Examination of Crystals in						Synovial Fluid")

E.	Amniotic Fluid:

1. 	Appearance: Colorless to pale straw; clear to slightly cloudy

2. 	Cell count and differential:

a.	No granulocytes or bacteria should be present.

b.	Squamous cells are considered normal and not counted in the	differential.

X. 	REFERENCES:

1.	Kjeldsberg, Carl and Knight, Joseph. Body Fluids. Laboratory Examination of 				Amniotic.

Cerebrospinal. Seminal. Serous and Synovial Fluids. ASCP Press, 3rd edition. 				1993.

2.	Spriggs and Boddington. The Cytology of Effusions. Grone and Stratton, 1968.

3.	Todd, Sanford, Davidsohn. Clinical Diagnosis and Management by Laboratory 			Methods. Saunders, 1979. Chapter 18, pp. 635 - 679.

4.	Romero R, Emamian M, Quintero R, et al." The Value and Limitations of the 			Gram Stain in the Diagnosis of Intraamniotic Infection." Am J Obstet Gynecol. 			1988; 159: 114-119.

5.        Cytospin 3 manual

X.  	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 11-1999.

H-2	This procedure was revised by N. Ortoli-Drew on 9-2005.

H-3	This procedure was revised by D. Fico on 9-2007.

H-4      This procedure was revised by D.Fico on 11-2010.

H-5      This procedure was revised by D. Fico on 3-14-2011.

H-6      This procedure was revised by D. Fico on 9/2012.

Appendix E

3% Glacial Acetic Acid- Reagent for counting bloody body fluids to lyze rbcs to facilitate enumerating nucleated cells.

Preparation:  15 ml conc. Glacial Acetic Acid( kept in acid cabinet in Spec. Heme)

diluted into 485 ml of distilled water.

Stability:       Indefinitely

Reagent needs to be properly labeled.