- I.	PRINCIPLE:

- STA-Stachrom Plasminogen, (Cat # 0658), is intended for the quantitative

- determination of the plasminogen level by the synthetic chromogenic substrate method.

Under the influence of tissue or plasma activators, plasminogen is converted into

plasmin.  The main role of plasmin is to degrade fibrin and secondarily fibrinogen,

leading to the production of degradation products of stabilized or non-stabilized fibrin

and fibrinogen, respectively.

The test procedure comprises of two steps:

A.   	an excess of streptokinase is added to the test sample whose

plasminogen is to be determined.  The streptokinase and the

plasminogen form a complex and this has a plasmin-like activity.

B.   	the quantity of streptokinase-plasminogen formed is assessed by its

action on the synthetic substrate CBS 30.41.  Quantification is

achieved by the amount of p-nitroaniline released, measured at

405 nm.

- Under these conditions, the reaction is insensitive to plasma inhibitors.

- II.	SPECIMEN:

A.  	 Citrated blood 9:1 (blood to anticoagulant) 3.2% or 3.8% sodium citrate.

- Follow NCCLS guidelines H3-A3 and H21-A2.  No other anticoagulant

is acceptable.

B.  	 Centrifugation:	20 minutes at 3000 g

C.  	 Plasma storage:	6 month at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

D.   	Samples that are short draws, clotted or hemolyzed may yield incorrect results.

- III.	REAGENTS AND EQUIPMENT:

- Centrifuge

Pipettes

Pipette tips

STA/STA Compact

Cuvette roll (contains 1000 cuvettes)

- IV.	REAGENTS:

A.  	STA Unicalibrator: Freeze-dried citrated human plasma.  This plasma

is intended for use as a calibration plasma for functional assays. Reconstitute

each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at room 				temperature.

Swirl gently.  Reconstituted stability on the STA/STA Compact is 4 hours.

- B.   	STA System Control N + P: Control N is citrated normal human plasma and

Control P is citrated abnormal human plasma, both freeze-dried.  Reconstitute

each vial with 1.0 ml reagent grade water.  Let set 30 minutes at room 				temperature.

Swirl gently.  Reconstituted stability on the STA/STA Compact is 8 hours.

C.	Owren-Koller Buffer: Ready to use.  Request the sample drawer to open through 	the MAIN MENU under LOADING and bar code the buffer.  Place the buffer 	into the sample drawer.  Stability on the STA/STA Compact is 120 hours.

D.	Reagent 1: Streptokinase: Approximately 15,000 IU/bottle.  This streptokinase is 	stabilized with human albumin.  Reconstitute each vial with 3.0 ml reagent grade 	water.  Let set 60 minutes at room temperature.  Swirl gently.  Request the reagent 	drawer to open through the MAIN MENU under LOADING and bar code the 	reagent.  Place the reagent into the reagent drawer, RD 1.  Reconstituted stability 	on the STA/STA Compact is 7 days.

E.	Reagent 2: Chromogenic substrate CBS 30.41, freeze-dried.  Reconstitute each 	vial with 3.0 ml reagent grade water.  Let set 60 minutes at room temperature then 	swirl gently.  Request the reagent drawer to open through the MAIN MENU 	under LOADING and bar code the reagent.  Place the reagent into the reagent 	drawer, RD 2.  Reconstituted stability on the STA/STA Compact is 7 days.

- V.	CALIBRATION:

A	STA Unicalibrator: After the 30 minute reconstitution period, request the sample 	drawer to open (STA) or the reagent drawer to open (STA Compact) through the 	       	MAIN MENU under LOADING and bar code the calibrator.  Place the calibrator

into the appropriate drawer.  The assay value for Plasminogen is transferred to 			the STA/STA Compact when the calibrator is loaded and prompted for the bar 			code information.

B.	All standard dilutions are automatically prepared by the STA/STA Compact by 	diluting the STA Unicalibrator with Owren-Koller Buffer according to the 	parameters entered in the Test Set-up.

C.   	If calibration is necessary, order through the MAIN MENU by selecting CALIB/CONTROL.

Select CALIBRATION and enter.  Place the cursor on PLG and enter.  Select

PLG by pressing F1 and then F10 to run the calibration.

D.    	STA System Control N + P are used by the STA/STA Compact to automatically

validate the calibration. (Refer to the QUALITY CONTROL section).

E.	To examine the calibration curve: Through the MAIN MENU under  	CAL/CONTROL select CALIBRATION and enter.  Place the cursor on PLG  and 	enter.  The calibration curve will appear on the STA/STA Compact screen

F.	To verify the calibration curve is valid: Check that the correlation coefficient (r 	       	value) of the curve is 0.90 or greater.  Compare the actual vs the interpolated 	values of the curve.  If a data point difference is greater than 10% for the actual 	value vs the interpolated value, the point can be deleted form the curve data.  	When this is done, the STA/STA Compact will automatically re-draw the curve, 	re-calculate the r value and re-run the controls.  If any controls are outside the 	expected range, the control must be re-run.  If the value is still outside the 	expected range, a new calibration must be run along with fresh controls.

G.	To print the calibration curve: While examining the calibration curve on the 		STA/STA Compact screen, press ESC for options.  Select PRINT.  The STA/STA 	Compact cannot print a calibration curve while the STA/STA Compact is running.

- VI.	QUALITY CONTROL:

A.  	STA System Control N + P: After the reconstitution period, request the sample

drawer to open (STA) for the reagent drawer to open (STA Compact) through the

- MAIN MENU under LOADING and bar code the controls.  Place the controls

into the sample drawer.

B.  	QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC if the

patients’ samples have been loaded.  The QC can also be ordered manually 			before loading patients’ samples.

C.  	All control ranges are monitored automatically by the STA/STA Compact.  If any

controls are outside the ± 2 SD range, the STA/STA Compact will audibly and

visually alarm the operator.  Otherwise, the results can be found in the AUTO

CONTROL file (STA only) and the individual QC files.  Control results are

automatically filed in the STA/STA Compact QC file.  All results for a 24 hour

period will be reduced to a “mean” value at midnight.  This mean is used in the

statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- VII.	PROCEDURE:

A	Refer to the START-UP procedure for the STA/STA Compact before running 	patient samples on the STA/STA Compact at the start of each shift.

B.  	Run the calibration if needed.

- C. 	Request quality control if using a stored curve: Through the MAIN MENU

under CAL/CONTROL select QUALITY CONTROL and enter.  Place the

cursor on PLG.  Select PLG by pressing F1 and then F10 to run the QC.

- D.  	Load patients’ samples: Access the sample drawer(s) through the MAIN MENU

under LOADING.  After the drawer opens, identify the type of specimen, such as

micro-sample, with F8, or stat, with F12.  Identify the sample by bar coding or

manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

- E.  	 In MANUAL MODE, the operator must order the test form the menu, and

then press F10 to save.

F	In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	       	selected in the AUTO MODE profile.

G.	In TELELOADING MODE, the STA/STA Compact will query the host computer  	and download the test as well as assign the status, i.e. stat.

H.   	As soon as the sample drawer closes, the TEST STATUS screen will appear.

- If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

- I.   	All dilutions of the calibrator, controls and patients’ samples are automatically

repared by the STA/STA Compact according to the parameters entered in 	the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

dilution provided that the option has been entered in the Test Set-up.

- J.   	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

K.    	For results in question that needs operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left hand corner of the screen,

i.e. re-run test.

- VIII.	CALCULATIONS:

A.  	The STA/STA Compact automatically converts the results from a standard curve

(lin-lin) to % activity.

B. 	The STA/STA Compact uses a dilution of 1:20, sample to buffer.  Following pre-

set parameters, the STA/STA Compact will automatically dilute the sample lower 		if the result falls below the lowest data point on the calibration curve.  The 		STA/STA

Compact will also automatically dilute the sample at a higher dilution if the 			results	falls above the highest data point on the calibration curve.

C. 	Linearity of the assay is 10-160%.

IX.	REFERENCE RANGE:

_70____ - _140____ % Each laboratory should determine its own reference range.

- X.	REPORTING RESULTS:

- The results for Plasminogen are reported out to the nearest whole number %.  (Example: 	103%).

- For a result below the linearity of the assay, enter “<10”.

- XI.	NOTES:

A.   When the STA/STA Compact re-tests a patient sample at a more appropriate

dilution, (as pre-set in the Test Set-up) the results in the TEST PANEL screen

will appear in blue and have already been corrected by the STA/STA Compact for

the dilutional difference.

- B.   Since most patients’ results are higher than 100% and will probably be

re-tested at a more appropriate dilution, it is more efficient to run the

patient samples initially at a higher dilution, i.e. 1:20.  In order to do this,

a Dependent Test with the selected dilution must be created. (See the

Operators Manual - Dependent Test).

- XII.	LIMITATIONS OF THE PROCEDURE:

1.   This assay is insensitive to plasma inhibitors and fibrin or fibrinogen

degradation products.

- 2.   During thrombolytic treatments it is of interest to measure the plasminogen

level to monitor the hepatic regeneration level and to control and adjust the

perfusion rate if plasminogen is being given to the patient.

- 3.   The plasminogen level in newborn babies is usually low, particularly

if they are premature.

- XIII.	REFERENCES:

1.   STA-Stachrom Plasminogen, Colorimetric Assay of Plasminogen by STA

Analyzers.  Package insert 23477-Revised March 1997.

- 2.   STA Unicalibrator, Calibration Plasma for Functional Assays of Coagulation

Parameters on the STA.  Package insert 26344 - Revised May 1996.

- 3.   STA System Control N + P, Control Plasmas for Assays of Coagulation

Parameters on STA.  Package insert 26384 - Revised May 1997.

- STA/STA Compact Operators Manual.

XIV.  	HISTORY:

H-1	This procedure was written by N. Kriz on 01-06-05.

H-2	This procedure was reviewed by N Kriz on 10/12/2010.

H-3	This procedure was reviewed by N Kriz on 4/1/2011.

H-4	This procedure was reviewed by N Kriz on 2/10/2012.

H-5	This procedure was reviewed by N Kriz on 3/12/2013.