I.	PRINCIPLE:

Plasma which is deficient in any of the factors in the extrinsic pathway will yield a prolonged prothrombin time ((PT)). Known factor-deficient plasma can be used in the identification and quantification of the factor deficiency in the patient plasma. A mixture of the deficient plasma and the patient plasma is tested by a PT and the result is interpreted using a reference curve obtained with dilutions of Standard Human Plasma or a normal plasma pool mixed with the deficient plasma.

II.		SPECIMEN:

A.	Type:

Mix nine parts of freshly collected blood with one part of 3.2% sodium citrate anticoagulant.

Invert the tube gently three or four times immediately after venipuncture to ensure proper mixing 	of blood and anticoagulant. A syringe or evacuated tubes (blue top) may be used for collection. If  multiple specimens are collected,, the coagulation sample should be the second or third tube collected. If blood is drawn from an indwelling catheter, the line should be flushed with 5.0 mL 	saline and the first 5 mL of blood or six dead space volumes of the catheter discarded or used for

other laboratory tests. The citrate concentration must be adjusted in patients who have hematocrit values above 55%. Specimens that are clotted, collected in the wrong tube, or have less than the 90% expected fill should be rejected.

B.	Handling Conditions:

The whole blood specimen is checked for clot formation by gentle inversion and observation.

Centrifuge the capped blood specimen to produce platelet-poor plasma (platelet count <10x109/L for 10 minutes at 3000 g.  Patient plasma should be tested within 4 hours . If immediate testing is to be done, the plasma may remain on the packed cells.

For special coagulation testing, spin samples 20 minutes at 3000 g, separate plasma into plastic tubes, label and freeze at –70C until ready to test. A frost-free freezer should not be used. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately after thawing. If testing is delayed, the sample may be held for 2 hours at 4C until tested.

III.		EQUIPMENT AND MATERIALS:

A.	Equipment:

BCS  Coagulation Analyzer

Rotors

Waste, rinse and disinfectant containers

BCS Sample Cups

B.	Materials:

Siemens Dade Innovin prothrombin time reagent

Coagulation Factor II- (Siemens OSGR13)

Coag Factor V-, VII-, and X- (George King) Deficient plasmas

Control N (normal control)

Control P (abnormal control)

NERL water

Imidazole-buffered saline (IBS)

BCS Wash Solution

C.	Preparation:

1.	Dade Innovin-- A lyophilized human recombinant human thromboplastin

a.	Reconstitute each vial of lyophilized Innovin with 10.0 mL

NERL water

b. 	Immediately after addition of water, mix contents of vial gently before use to assure complete suspension.

c.	Allowed to stand 15 minutes, then swirl gently before loading

d.	Store at 2-8C. Do not freeze.

Stability after reconstitution:   5 days at 2-8C stored stoppered

48 hours on board the instrument at 15C

2.	Coagulation Factor II-Deficient Plasma—lyophilized, immunoadsorbed

human plasma with less than 1% residual factor activity

a.	Reconstitute  II-Deficient Plasma with 1 mL distilled or deionized

water.

b.	Restopper vial and allow to stan15 minutes until dissolved.

c.	Invert gently to mix. Do not shake.

Coagulation Factor II-Deficient Plasma stable 8 hours at 15-25C

- d.	Do not store at 2-8C

- e.	Reconstituted deficient plasma can be frozen AT -70C and thawed

once (USE WITHIN ONE MONTH). Plasma should be rapidly

frozen in original vial.  Thaw at 37C  1C within 10 minutes.  Thawed plasma should be used within two hours at 15-25C

3.	Coagulation Factor V-,VII, and X-deficient plasmas, frozen, George King Biomedical

4.	Siemens Standard Human Plasma—lyophilized normal plasma pool.

Used for calibration, if needed.

5.	Control N (lyophilized)—reconstitute with 1.0 mL NERL water

Used for Quality Control-Normal

6.	Control P (lyophilized)—reconstitute with 1.0 mL NERL water

Used for Quality Control-Abnormal

D.	Quality Control:

Siemens Control N—assayed Normal Control

Siemens Control P –-assayed Abnormal Control

1.	For specific control plasma values see assigned values (on insert) for the respective lot

2.	Controls should be tested at least once every 8 hours for any assays run for patient 	testing during that interval.

3.	Controls should also be tested upon reagent changes, and after each new calibration curve

4.	Controls should be run in the same manner as the test samples.

5.	Corrective action when tolerance limits are exceeded:

a.	Verify reagent performance.

b.	check instrument performance.

c.	document actions taken to identify and correct the problem before reporting any patient data.

d.	Recalibration may be necessary if control values are still out of range.

IV.	PROCEDURE - STEPWISE

A.	Instrument Set-Up and Calibration (minimum every 6 months, and if consistent problems with quality controls or patient samples):

1.	Press Calibration Icon and choose New key.

2.	Click the required assay in the for this assay list box.

3.	Select the lot number of reagents and standards on board for the

calibration curve.

4.	Press Measure curve key to begin curve calibration.

5.	Place the reagents and standards required for the calibration on the

analyzer.

6.	Click OK to confirm the message.

7.	When the reagents and standards have been positioned, the analyzer starts calibration.

8.	Acceptable and valid calibration curves will display a BCS symbol next to the curve and date.  Unacceptable curves will display an “exclamation point”.

9.	View the curve for validity. Click the Calibration button. Click the desired reference curve in the Curve overview column.

10.	Click the Show curve key to view the reference curve. The curve will be displayed.

11.	To re-measure one or more points displayed in the graphic, click the point to be corrected. The point will be highlighted. Click the Repeat key and close the dialog. The point will be re-measured and entered in the graphics.

# B.	Loading Patient Samples On The BCS

C.	Bar-Coded Samples WithoutHost:

1.	Load bar-coded samples into any sample rack with the bar-code label visible through the window of the rack position.

2.	Insert the rack into any open lane.(4-14)

3.	Push the rack until you feel it stop

4.	Open the JOB LIST icon, if not yet open.

5. 	The JOB LIST screen displays the samples loaded with a small green analyzer next to the sample ID.  The small green analyzer icon indicates the sample is currently loaded in the processing area

6. 	To assign tests requests to the samples that have been loaded follow the steps below:

a. 	Locate the test requests required for each sample listed.

b.	Double click on the box that intersects between the sample identification the test column.  A red “X” appears in the intersection.  The “X” denotes that a 	test request for this sample is in process.

7. 	Sample processing begins automatically.

D.	Assigning A Stat Test Request:

1.	 Prior or after to assigning test requests, double click on the sample ID number

2.	Click on the STAT box.  The box will be marked with a “X”.

3.	Confirm the entry by clicking OK

4. 	Locate the test requests for each test required for each STAT sample listed.

5.	Double click on the box that intersects between the sample identification number and the test column.  A red “X” appears in the intersection.  The “X” denotes that a test request for this sample is in process.

6. 	Sample processing begins automatically.

V.	NON-BAR-CODE SAMPLES WITHOUT HOST:

A. 	Load the samples into the sample racks into consecutive positions.

B.	Click on the LOADING icon.

C. 	Click on the number that corresponds to the sample rack (the rack number is highlighted)

D. 	Click on the Select rack button (a picture of the sample rack is displayed)

E. 	Click on the position number of the sample rack in which the sample has been placed.

F. 	Click on the Sample ID box

G. 	Enter the sample identification number.

H.	At the question “Is the sample STAT?” answer either yes or no.

“Yes” answer will result in a red display of the sample ID in the lab journal.

proceed to the next step.

I. 	To request Individual Assays:

1. 	Locate the assay box

2. 	Click on the assay

3. 	Repeat the above steps until all assays are requested.

J.	Click on the INSERT button and the sample ID appears in sample position of the rack and the next position of the sample rack is highlighted.

K. 	To request another:

1.	 Sample in the same rack:   Repeat steps 6-11

2.	 Sample Rack:  Click on OTHER RACK button and repeat steps 3-11.

L.	Insert the sample racks into any open lane (An open lane is one without the LED illuminated or blinking.

M.	Push the rack until you feel it stop.

N.	Repeat steps 12 and 13 until all sample racks have been loaded.

O. 	The JOB LIST screen displays the samples loaded with a small green analyzer next to the sample ID.

The small green analyzer icon indicates the sample is currently loaded in the processing area.  A red “X” appears in the intersection between the sample ID number and the test column and denotes that a test request for the sample is in progress.

P. 	Sample processing begins automatically.

VI.	LOADING AND PROGRAMMING CONTROLS ON THE BCS:

A. 	Place opened control bottles into the 5 ml reagent rack with the bar-code label on the bottle visible through the window of the rack position.

B. 	Insert the rack into any open lane from 4-14 (an open lane is one that does not have a blinking or illuminated LED.)

C.	Push the rack until you feel it stop

D. 	Click on the Control Journal icon.

The small green analyzer icon indicates that the Control lot number is currently loaded in the processing area

E. 	Double click on the box that intersects between the control identification and the test column.  A red “X” appears in the intersection and denotes that a test request for the control is in progress.

F. 	Control processing begins immediately.

G. 	Repeat step 5 until all assays for each level of control have been requested.

H. 	All factor assays controls for the BCS must be requested in Soft and posted there when acceptable. Refer to Soft computer bench manual for details about Total QC.

VII.	REPORTING RESULTS:

Results of the factor assay testing should be reported in % normal.  These results should be related to the reference interval for the factor.

Record results in the computer system; most results autoverify. Post through Instrument Menu referring to Soft bench manual as needed.

Hemolyzed, lipemic, or icteric samples must be noted with the result.

A.	Reference Interval:

Reference interval values determined for PT (extrinsic) factors (II, V, VII. And X) = 50-150%

B.	Procedures for Abnormal Results:

Factor levels less than 40% of normal (first time) must be phoned to the patient’s nursing physician as a Critical Value.

On occasion, a circulating inhibitor such as a strong lupus anticoagulant or—less

frequently—a strong inhibitor to a factor other than the one being analyzed, will

affect the factor assay results. This is an artifact of the assay system, and the

results will appear low (even markedly low) at the 1:10 dilution, somewhat higher

at the 1:20 dilution, and normal or near-normal at the 1:40 dilution (as the

inhibitor is diluted out). In this case, the highest level obtained should be reported

out. In an extreme case, it may be helpful to pre-dilute patient plasma 1:2 with

IBS and rerun the assay, multiplying the final result by 2.

VIII.		PROCEDURE NOTES:

Overall performance of factor assay testing is dependent on reagent and instrument performance.

Acceptable variability (imprecision) should be such that the total coefficient of variation (CV) of the analytic system is less than 20% on the same lot of control plasma.

The measuring range is defined by the concentration of the calibrators used and the extrapolation limits set.

VIX.		REFERENCES:

1.	Innovin package insert, Siemens, Marburg, Germany, June 2009

2.	Coagulation Factors II, V, VII, X package insert, Dade Behring, Marburg, Germany, January 1999

3.	Standard Human Plasma package insert, Dade Behring, Marburg, Germany, January 1999

4.	Control Plasma N package insert, Dade Behring, Marburg, Germany, May 1998

5.	Control Plasma P package insert, Dade Behring, Marburg, Germany, January 1999

6.	National Committee for Clinical Laboratory Standards. Collection, transport and processing of blood specimens for coagulation testing and general performance of coagulation assays. Approved Guideline, 3rd Edition. NCCLS Publication H21-A3. Wayne, PA, 1998.

7.	Dade-Behring BCS Operator’s Manual

8.	Factor II, V, VII and X Application Sheets: 11350501E, 11360501E, 11370501E, 11380501E

X.  	HISTORY:

H-1	This procedure was written by N. Kriz on 7-05-06.

H-2	This procedure was revised by N Kriz on 10/12/2010.

H-3	This procedure was revised by N Kriz on 4/18/2011.

H-4	This procedure was revised by N Kriz on 12/1/2012.

H-5	This procedure was reviewed by N Kriz on 3/12/2013.