- I.	PRINCIPLE:

The Rotachrom Heparin kit, (Cat #0612), is intended for the quantitative determination

of the plasma levels of unfractionated heparins (UFH) by the measurement of their

anti-Xa activity in a competitive system using a synthetic chromogenic substrate.

The normal function of a molecule of factor Xa, as soon as it appears in plasma,

is to cleave its natural substrate, prothrombin, to generate thrombin, the enzyme

responsible for the fibrin clot formation.  In the presence of heparin, competition

occurs between this mechanism and the inhibitory mechanism exerted by the

heparin-antithrombin III complex, this inhibition being largely responsible for the

action of heparin.  The proposed method is a one-step reaction based on a similar

principle: as soon as factor Xa is added to the plasma-substrate mixture, two reactions

take place simultaneously:

-  hydrolysis of the substrate by factor Xa

-  inhibition of factor Xa by the heparin-antithrombin III complex

After the necessary period of time for the competitive reaction to reach equilibrium,

the quantity of pNA that is released is inversely proportional to the concentration

of heparin present in the plasma.

Unfractionated heparins (UFH) are used for the prevention and treatment of

venous thromboembolism.  UFH acts I conjunction with antithrombin III.  The

quantitative determination of plasma UFH levels is a useful tool for monitoring

treatment efficacy.

- II.	SPECIMEN:

A. 	Citrated blood 9:1 (blood to anticoagulant) 3.2% sodium citrate.

No other anticoagulant is acceptable.

B. 	Centrifugation:	20 minutes at 3000 g - room temperature

C.   	Plasma storage:	   6 months at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

D.   	Unacceptable specimens: Samples that are short draws, clotted,

hemolyzed, or icteric may yield incorrect results. These should be noted as a comment.

- III.	REAGENTS AND EQUIPMENT:

- Centrifuge

Pipettes

Pipette tips

STA/STA Compact

Cuvette roll (contains 1,000 cuvettes)

- IV.	REAGENTS:

A. 	STA Hepanorm H: A set of calibrator plasmas intended to be used for

unfractionated heparin assays using the anti-Xa principle.

Reagent 1: STA-Hepanorm H 0: human plasma free of UFH, freeze-dried.

Reagent 2: STA-Hepanorm H 3: human plasma containing a well-defined

quantity of heparin (see enclosed Assay Value Insert), freeze-dried.

Reagent 3: STA-Hepanorm H 6: human plasma containing a well-defined

quantity of heparin that is greater than that of Reagent 2 (see enclosed

Assay Value Insert), freeze-dried.

Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at

room temperature.  Do not invert vials.  Swirl gently.  Reconstituted stability

on the STA/STA Compact is 4 hours.

B. 	STA Heparin Control: Two plasmas containing different levels of unfractionated

heparin (UFH) intended to be used for quality control of UFH assays using the

anti-Xa principle.

Reagent 1: STA Heparin Control 2: human plasma containing a well-defined

quantity of UFH (see enclosed Assay Value Insert), freeze-dried.

Reagent 2: STA Heparin Control 5: human plasma containing a well-defined

quantity of UFH that is greater than that of Reagent 1 (see enclosed Assay

Value Insert), freeze-dried.

Reconstitute each vial with 1.0 ml reagent grade water.  Let sit 30 minutes at

room temperature.  Do not invert vials.  Swirl gently.  Reconstituted stability

on the STA/STA Compact is 4 hours.

C.	Reagent 1: Substrate: lyophilized chromogenic substrate CBS 52.44, approximately 4.2 umoles per vial.  Reconstitute each vial with 4.0 ml reagent grade water.  Replace the perforated plastic 	cap on the vial.  Let sit 60 minutes at room temperature.  Do not invert vials.  Swirl gently.  Request the reagent drawer to open through the MAIN MENU under LOADING and bar code the reagent.  Place the reagent into the reagent drawer, RD1.  Reconstituted stability on the STA/STA compact is 7 days.

D.  	Reagent 2: F.Xa: bovine factor Xa, approximately 0.7 IU per vial.  Reconstitute

each vial with the entire contents of one Reagent 3 vial from same kit.  Replace the perforated plastic cap on the vial.  Let sit 60 minutes at room temperature.  Do not invert vials.  Swirl gently.  Request the reagent drawer to open through the MAIN 	MENU under LOADING and bar code the reagent.  Place the reagent into the reagent drawer, RD2.  Reconstituted stability on the STA/STA Compact is 7 days.

E.  	Reagent 3: F.Xa Solvent: Ready to use.  Transfer the entire contents of Reagent 3 into a Reagent 2 vial.

- V.	CALIBRATION (if necessary):

A. 	STA Hepanorm H: After the reconstitution period, request the sample drawer

to open (STA) or the reagent drawer to open (STA Compact) through the MAIN 	 	MENU under LOADING and bar code each vial of calibrator.  Place each vial of 	      	calibrator into the appropriate drawer after it has been bar coded.  The assay value

for each Heparin Level is transferred to the STA/STA Compact when each

calibrator is loaded and prompted for the bar code information.

B.  	All standard dilutions are automatically prepared by the STA/STA Compact by 	 	      	diluting the STA Hepanorm H calibrators with Owren-Koller Buffer according to

the parameters entered in the Test Set-up.

C.  	To order calibration: Through the MAIN MENU select CAL/CONTROL.

Select CALIBRATION and enter.  Place the cursor on HNF COLB and enter.

Select HNF COLB by pressing F1 and then F10 to run the calibration.

D. 	STA Heparin Controls are used by the STA/STA Compact to automatically

validate the calibration.  (Refer to the QUALITY CONTROL section).

E.  	To examine the calibration curve: Through the MAIN MENU under

CAL/CONTROL select CALIBRATION and enter.  Place the cursor on

HNF COLB and enter.  The calibration curve will appear on the STA/STA

Compact screen.

F.  	To verify the calibration curve is valid: Check that the correlation coefficient

(r value) of the curve is 0.90 or greater.  If any controls are outside the expected

range, the control must be re-run.  If the value is still outside the expected range,

a new calibration must be run along with fresh controls.

G.	To print the calibration curve: While examining the calibration curve on the 	  STA/STA Compact screen, press ESC for options.  Select PRINT.  The STA/STA Compact cannot print a calibration curve while the STA/STA Compact is running.

- VI.	QUALITY CONTROL:

A.  	STA Heparin Controls: After the reconstitution period, request the sample drawer

to open (STA) or the reagent drawer to open (STA Compact) through the MAIN 	 	      	MENU under LOADING and bar code the controls.  Place the controls into the 	 	      	appropriate drawer.

B. 	QC will automatically be run when a new calibration curve has been requested.

If using a stored curve, the STA/STA Compact will automatically run QC when

the patients’ samples have been loaded.  The QC can also be ordered manually 		before loading patients’ samples.

C.	All control ranges are monitored automatically by the STA/STA Compact.  If any 	       controls are outside the ± 2 SD range, the STA/STA Compact will audibly and visually alarm the operator.  Otherwise, the results can be found in the AUTO CONTROL file and the individual QC files.  Control results are automatically filed in the STA/STA Compact QC file. All results for a 24 hour period will be converted to a mean value at midnight. This mean value is used in the statistical data and is plotted on the Levy-Jennings chart as a daily mean.

- VII.	PROCEDURE:

A.	Refer to the START-UP procedure for the STA/STA Compact before running patent samples on the STA/STA Compact at the start of each shift.

B. 	Run the calibration if needed.

C. 	Request quality control if using a stored curve: Through the MAIN MENU

under CAL/CONTROL select QUALITY CONTROL and enter.  Place

the cursor on UFH.  Select UFH by pressing F1 and then F10 to run the QC.

D.  	Load patients’ samples: Access the sample drawer through the MAIN MENU

under LOADING.  After the drawer opens, identify the type of specimen, such

as microsample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.  	In MANUAL MODE, the operator must order the test from the menu, and then

press F10 to save.

F.   	In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	 	      	selected in the AUTO MODE profile.

G. 	In TELELOADING MODE, the STA/STA Compact will query the host computer	 	      	and download the test as well as assign the status, i.e. stat.

H.   	As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

I.   	All dilutions of the calibrators, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

dilution provided that the option has been entered in the Test Set-up.

J.    	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

K.   	For results in question, that need operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left-hand corner of the screen,

i.e. re-run test.

- VIII.	CALCULATIONS:

A.  	The STA/STA Compact automatically converts the results off of a standard curve

(lin-lin) to IU/ml.

B.   	The STA/STA Compact uses a dilution of 1:2, sample to buffer.  Following pre-set parameters, the STA/STA Compact will automatically dilute the sample higher if the results fall above the highest data point on the calibration curve.

IX.	REFERENCE RANGE:

The UF Heparin therapeutic range is 0.3 - 0.7 IU/ml

- X.	REPORTING RESULTS:

The results for the Heparin Assay are manually reported out to the nearest

  - IU/ml.  (Example:  0.5 IU/ml).

The detection threshold of the assay is 0.10 U/mL. For a result below this threshold, enter 0 U/mL

- XI.	NOTES:

A.	When the STA/STA Compact re-tests a patient sample at a more appropriate dilution, (as pre-set in the Test Set-up) the results on the TEST PANEL screen will appear in blue and have already 	been corrected by the STA/STA Compact for the dilutional difference.

B. 	The assay’s linearity extends to 0.70 IU/ml.  In order not to report erroneously

low results,  the auto re-test parameters must be pre-set in the Test Set-up.

The parameters are 1:3 dilution > 0.70.  Using these parameter the assay’s

linearity is extended to 1.1 IU/ml.  In order to determine a heparin level on a

sample with a higher level, the operator must:

-  mix the patient sample 1:1 with Normal Plasma

-  re-test the heparin assay on this mixture

multiply the result  X 2 to correct for the dilutional difference

C. 	The lower detection threshold is 0.10 U/ mL.

- XII.	LIMITATIONS OF THE PROCEDURE:

A.   	Any release of platelet factor 4 (PF4), which is a potent heparin inhibitor,

will lead to an under-estimation of the heparin level in the plasma being

tested.  Careful and adequate centrifugation is essential: the higher the

level of residual platelets, the greater the risk of PF4 release.

B.  	The APTT and the quantitative assay of heparin do not always correlate.

As a test of the overall coagulability of the patient, the APTT reflects

the effect produced by several phenomena, such as, the presence of heparin

and the presence of an inflammatory syndrome.  It is noted that for a given

heparin level, different individuals may exhibit different APTT values as

well as circadian variability.

C.	Icteric plasma may render results somewhat lower than actually is due to coloration from elevated bilirubin. Lipemic plasma may be ultracentrifuged before testing, if results are in question. Hemolysis may affect results and, if indicated, test should be repeated on fresh sample.

- XIII.	REFERENCES:

1.   	STA Rotachrom Heparin, Colorimetric Assay of Unfractionated Heparin

(UFH) by STA Analyzers.  Package insert 23313-Revised November 1997.

2.  	STA Hepanorm H, Calibration Plasmas for Assays of Unfractionated Heparins

(UFH)  Using anti-Xa Method on STA.  Package insert 26311-Revised February

1997.

3. 	STA Heparin Control, Control Plasmas for Assays of Unfractionated Heparin

(UFH)  Using Anti-Xa Method on STA. Package insert 26308-Revised April

1997.

4.	STA/STA Compact Operators Manual.

XIV.  	HISTORY:

H-1  	This procedure was written by N. Kriz on 12-23-2004.

H-2	This procedure was reviewed by N Kriz on 10/12/2010.

H-3	This procedure was reviewed by N Kriz on 4/1/2011.

H-4	This procedure was reviewed by N Kriz on 2/13/2012.