- Purpose and Scope:

When analyzed on an automated cell counter, specimens that are markedly icteric, lipemic, hemolyzed, or contain RBC agglutinins may have an elevated MCHC.

Specimens run on the Sapphire with MCHC >37.3, should immediately have a spun HCT performed (only rerun on instrument if suspect instrument sampling problem)

Resolution of elevated MCHC results must be expedited; either rerun immediately or spin a HCT. RBC indices parameters will need to be canceled or recalculated depending upon results of rerun, spun HCT, incubation or plasma color/opacity. Follow the Soft Procedure for defaulting results which follows the Protocol for elevated MCHC.

If repeated MCHC ≤37.3, report the Sapphire results. If the repeat MCHC is still >37.3 proceed as follows:

- Make and stain a smear for review of morphologic findings

# Do spun HCT and observe the plasma

A.	ICTERIC OR LIPEMIC           HCT WILL BE ACCURATE

Report                     WBC, RBC, HCT, PLT, MCV

Soft computer reporting

- In Instrument menu, under Instrument tab F4 results for MCH, MCHC and  HGB

- Click on Lab Results tab

- Click on MCH result and click Def Res Click OK to “Result field contains unexpected characters. Double-check it.” The result will be Invalid Result.

- Repeat steps 3 and 4 for tests MCHC and HGB

- Click on CBC test code result (top of Lab result screen), click Comm,

- Click CanMesg and choose comment @SH1 for icteric or @SH2 for lipemic.

- Click verify all and click yes to “Do you want to save changes to lab results page” pop-up

- Follow the usual resulting protocol from this point on

B. 	HEMOLYZED                       HGB WILL BE ACCURATE

Report                    WBC, HGB, PLT, MCV

Soft computer reporting

1.  In Instrument menu, under Instrument tab F4 results for RBC, MCH, MCHC and HCT

2.  Click on Lab Results tab

3.  Click on RBC result and click Def Res

4.  Click OK to “Result field contains unexpected characters.

Double-check it.” The result will be Invalid Result

5.  Repeat steps 3 and 4 for tests MCH, MCHC and HCT

6.  Click on CBC test code result (top of Lab result screen), click

Comm, click CanMesg and choose comment @SH3 for hemolyzed

7. Click verify all and click yes to “Do you want to save changes to lab

results page” pop-up

8.  Follow the usual resulting protocol from this point on

C. 	CLEAR AND NORMAL COLOR, spun HCT matches HGB

Report                    WBC, HGB, PLT and SPUN HCT (cmnt spun HCT)

Soft computer reporting

1.   In Instrument menu, under Instrument tab F4 results for, RBC, HCT,

MCV, MCH and MCHC

2.   Click on Lab Results tab

3.   Click on RBC result and click Def Res

4.   Click OK to “Result field contains unexpected characters. Double-check it”  			The result will be Invalid Result.

5.   Repeat steps 3 and 4 for tests MCV, MCH and MCHC

6.   Enter spun HCT result

7.   Click on CBC test code result (top of Lab result screen), click

Comm, click CanMesg and choose comment @SH4

- Click verify all and click yes to “Do you want to save changes to lab         results page” pop-up

9.   Follow the usual resulting protocol from this point on

D. 	CLEAR AND NORMAL COLOR, spun HCT matches instrument HCT

and    MCHC >37. 3

Report                     WBC, RBC, HGB, HCT, PLT, MCV, MCH, MCHC

Review              smear for presence of spherocytes (report if present)

Soft computer reporting

1.  In Instrument menu, click on Lab Results tab

2.  Click on CBC test code result (top of Lab result screen), click

Comm, click CanMesg and choose comment @SH5

3. Click verify all and click yes to “Do you want to save changes to lab

results page” pop-up

4.  Follow the usual resulting protocol from this point on

*make sure to make a slide for smear review of morphology*

E. 	LIPEMIC AND HEMOLYZED

Report                   SPUN HCT, WBC.PLT, MCV

Soft computer reporting

1.  In Instrument menu, under Instrument tab F4 results for HGB, HCT, RBC, MCH  and MCHC

2.  Click on Lab Results tab

3.  Click on HGB result and click Def Res

4.  Click OK to “Result field contains unexpected characters. Double-check it.”  The result will be Invalid Result

5.  Repeat steps 3 and 4 for tests RBC, MCH and MCHC

6.  Click on CBC test code result (top of Lab result screen), click

Comm, click CanMesg and choose comment @SH6

7. Click verify all and click yes to “Do you want to save changes to lab

results page” pop-up

8.  Follow the usual resulting protocol from this point on

F. 	CLEAR AND NORMAL COLOR, disproportionally low RBC suspect   RBC agglutination

INCUBATE SPECIMEN 15 MINUTES AT 37’C,

DO NOT INCUBATE  >15 MINUTES. RERUN ON MANUAL MODE

MCHC <37.3          report results

Soft computer reporting

1.  In Instrument menu, click on Lab Results tab

2.  Click on CBC test code result (top of Lab result screen), click Comm, click  CanMesg and choose comment @SH7

3.  Follow the usual resulting protocol from this point on

IF

MCHC >37.3

Report                     WBC, HGB, PLT

Soft computer reporting

1.  In Instrument menu, under Instrument tab F4 results for RBC, HCT, MCH, MCV and MCHC

2.  Click on Lab Results tab

3.  Click on RBC result and click Def Res

4.  Click OK to “Result field contains unexpected characters. Double-check it.”  The result will be Invalid Result

5.  Repeat steps 3 and 4 for tests HCT, MCH, MCV and MCHC

6.  Click on CBC test code result (top of Lab result screen) click Comm, click CanMesg and choose comment @SH8

7. Click verify all and click yes to “Do you want to save changes to lab

results page” pop-up

8.  Follow the usual resulting protocol from this point on

NOTE:

- if specimen is from ED, superstat call floor and explain that specimen is being

further  evaluated for elevated MCHC.

2.   cold agglutinin patients may also be icteric and or hemolyzed in which case

neither Hct nor Hgb will be accurate. Extremely low RBC count or discrepancy

between RBCi and RBCo should alert tech to presence of RBC agglutination.

3.   Patients with cold agglutinins may also exhibit warm agglutinins so repeated

incubation should not be done.

- History:

H-1   This procedure was written by N. Ortoli-Drew and Jessica Kaul 5-2012.

H-2   This procedure was revised by Elizabeth Degano 12-2012.

H-3   This procedure was revised by Lacey Mendillo 3/7/2013.