I.  	Purpose:

To define the Soft computer process when resulting Special Coagulation tests.

II.  	Procedure:

Special Coagulation Orderable Tests:

SPCOG – Resulting Template for Special Coagulation

A2AP – Alpha-2-Antiplasmin Activity

ATIII – Antithrombin Enzymatic Activity

APC – APC Resistance

HA – Anti-Xa Activity, LMWH

HAUFH – Anti-Xa Activity, UFH

ARIX – Anti-Xa Activity, Fondaparinux

INHIB – Inhibitor Assay

PHOSN – Phospholipid Neutralization

PLG – Plasminogen

PRC – Protein C, Functional

PRS – Protein S, Functional

PRSTF – Protein S, Total and Free Antigen

MPT – Prothrombin Time 1:1 Mix

MPTT– Partial Thromboplastin Time 1:1 Mix

RCF – von Willebrand Factor Activity

DRV – Dilute Russell Viper Venom Time

TT – Thrombin Time

VWFAG – von Willebrand Factor Antigen

FII – Factor II Activity

FV – Factor V Acitivity

FVII – Factor VII Activity

FVIII – Factor VIII Activity

FIX – Factor IX Activity

FX – Factor X Activity

FXI – Factor XI Activity

FXII – Factor XII Activity

FXIII – Factor XIII Screen

LUPUS - Lupus Anticoagulant Panel (DRV, PTT and PHOSN)

HYPER - Hypercoagulable Panel (APC, ATIII, PRC and PRS)

VWD - von Willebrand Panel (VWFAG, RCF and FVIII)

F8CH- Factor VIII Chromogenic

*APC, PRSTF, PTM, PTTM and DRV each have multiple test components

APC: APC1 = APC-1, APC2 = APC-2, APCR = Ratio result that gets reported

PRSTF: PRST = Total Antigen, PRSF = Free Antigen

PTM: PTMI = PT 1:1 Mix 0’, PTMC = PT 1:1 Mix Control 0’, PTMI2 = PT 1:1 Mix 60’, PTMC2 = PT 1:1 Mix Control 60’

PTTM: PTTMI = PTT 1:1 Mix 0’, PTTMC = PTT 1:1 Mix Control 0’,

PTTM2 = PTT 1:1 Mix 60’, PTTC2 = PTT 1:1 Mix Control 60’

DRV: DRVC = DRV Confirm, DRVS = DRV Screen, DRVR = Ratio result that gets reported

Special Coagulation Orderable Sendout Tests:

SCOGS – Resulting Template for Special Coagulation Sendouts

ATAG - ATIII Antigen

PAIAC - Plasminogen Activator Inhibitor-1, Activity

PRCA - Protein C Antigen

TPA - Tissue Plasminogen Activator

C4BPR - C4 Binding Protein

PAIAG - Plasminogen Activator Inhibitor-1, Antigen

PAIM - PAI-1, Mutation (4G/5G)

PTFR - Prothrombin fragment 1&2

TATC - Thrombin-Antithrombin complex

VWMA - von Willebrand Multimeric Analysis

VWPAG - vWF Protease Activity w/Rflx Inhibitor

REPT – Reptilase Time w/Rflx Rep Time 1:1 Mix

CRYOF - Cryofibrinogen

HAFGT- Hemophilia A (F8) 2 Inversions

*Call pending reports as needed throughout the day*

Canned Messages:

@HEM - Sample hemolyzed

@ICT - Sample icteric

@LIP - Sample lipemic

@POLY - APTT ____ seconds when repeated using ___ ug/ml polybrene.

@POSH - APTT shortened to ____ seconds when repeated using ___ ug/ml polybrene indicating the presence of heparin.

@POHE - APTT did not correct when repeated using polybrene, indicating heparin is not present.

Cancellation Comments:

|X#01	Sample clotted

|X#02	Quantity not sufficient

|X#04	Laboratory accident

|X#06	Incorrect tube/container type

|X#08	Does not meet clinical time requirements

|X#10	Broken/spilled in transit

|X#11  Duplicate test

|X#13	IV contamination

|XH01	Unable to report due to specimen problems.

|XH09	Test not indicated.

Pending test list:

-Open Resulting Worklist: Template – SPCOG or SCOGS

Status - Pend and Nonver

From order # - leave as default to see yesterdays and today’s pending orders (or can adjust date and order numbers to see only today’s orders or order from a few days ago until current date)

-Any order #’s that appear on the list have pending tests

-Tests that don’t have a V next to it are the tests that are

pending for an order

OR

-Open Resulting Worklist: Tests - Enter a single test or group of tests (Example:

ATIII, PRC, PRS, PRSTF, or LUPUS or HYPER)

Status - Pend and Nonver

From order # - leave as default to see yesterdays and

today’s pending orders (or can adjust date and order

numbers to see only today’s orders or order from a few

days ago until current date)

-Any order #’s that appear on the list have pending tests

-Tests that don’t have a V next to it are the tests that are

pending for an order

To print a pending test list:

-Open resulting worklist to be printed from direction above

-Click on the printer icon  then click on Worklist

-Change the printer to the appropriate laser printer

-Change the Layout to Regular by clicking on the ▽and clicking on Regular

-Click OK to print

To call extra labels from Resulting Worklist (to use for aliquot tubes):

-Select the order that needs labels printed

-Click on the Printer icon

-Click on Selected Label

-Change the printer to the appropriate label printer

-Click on the box next to HBT (Hematology Blue Top) so there’s a check in the box

-If more than one label is neededClick on the Label Properties tab

-Change the # of copies to the # of labels to be printed

-Click OK and labels will print to the label printer specified

Samples to be run and verified on the BCS:

-Sample arrives at Special Coag Bench after being spun for 20 minutes

-Uncap samples, place in BCS rack and run specified tests on BCS

-If results are not flagged by the instrument and are within normal reference value ranges the result will be autoverified

-If results are flagged or outside of the normal range they need to be manually verified:

-Open Instrument Menu

∙  Under Select Instrument: Choose MBCS3

∙  Under Select Date: Choose Loadlist and today’s results

-Tests that are pending are under Test List

∙  If there is a V in the Flags column and by a test result then that test was

autoverified but there is a pending test on the order

-Check results for flags and errors.

-If all results are OK to release on the order then select Post All.

-If only specific results are OK to release:

∙  Click on the Lab Results menu on the bottom of the Instrument Menu screen

∙  Select the result that is going to be verified

∙  Click on Verify and Save

*If the RCF result is >150 then that result is automatically canceled and the RCFH result is what will be available to verify

Verifying Factor Assays:

-Normal results will autoverify showing 1 test on the Instrument Menu

-Abnormal results will not autoverify and may show multiple results in the Instrument Menu

∙  Each dilution point that was run will be viewable ex. F210, F220, F240 for

Factor II

∙  If there is an orange ? with any factor assay result – no result will appear on the instrument menu

-Look at results on the BCS Screen to determine what result will be entered

-Click on the Lab Results menu on the bottom of the Instrument Menu screen and

enter the result from the BCS in the result field of the test

-Select the result and click Post All to verify results

-If all results are not ready to be resulted:

∙  Select (left click) the test that won’t be resulted and press F4.  This will

“cancel” the test result being sent out.  A “C” will appear in the Status column.

∙  To verify the result simply repeat the above and the “C” will no longer be in the

Status column and the result may be verified

Samples to be run and verified on the Stago:

-Sample arrives at the Special Coag Bench after being spun for 20 minutes

-Uncap samples, enter patient name and order number into Stago and order appropriate tests

-Results will need to be manually verified:

-Open Resulting Worklist: Template - SPCOG

Status - Pend and Nonver

From order # - leave as default to see yesterdays and

today’s pending orders (or can adjust date and order

numbers to see only today or previous days)

-Any order #’s that appear on the list have pending tests

-Select (left click) on the test to be resulted

-Enter results in the result box and click Verify to verify only the selected test or Verify All to verify all results on the order

-Click on the Save icon and OK to save the order

-Phosolipid neutralization: ≤8 will go in as a negative result. >8 will go in as a positive result.

-If VWFAG 1/8 is <60 then that result is automatically canceled and the 1/2 result is what will be available to verify

-If PRS is <50 then PRSTF is automatically reflexed on to the order

How to verify manual tests:

-Open Resulting Worklist: Template - SPCOG

Status - Pend and Nonver

From order # - leave as default to see yesterdays and

today’s pending orders (or can adjust date and order

numbers to see only today or previous days)

-Any order #’s that appear on the list have pending tests

-Select (left click) on the test to be resulted

-Enter results in the result box and click Verify to verify only the selected test or Verify All to verify all results on the order

-Click on the Save icon and OK to save the order

III.	HISTORY:

H-1	This procedure was written by Jessica Kaul on 3/7/2010.

H-2	This procedure was revised by Jessica Kaul on 4/12/2011.

H-3      This procedure was revised by Kelly Baker on 3/13/13.