- PURPOSE:

To define for all laboratory technical staff the expectations for set-up, maintenance and 	evaluation of satisfactory Wright stain on the Hematek stainer.

- POLICY:

The Hematek Wright stainers are used for routine Wright Giemsa staining of blood, bone marrow and fluid smears. These stainers are set-up each morning and test smears are evaluated for satisfactory stain before any patient smears are stained.

- PROCEDURE:

A.   Daily Maintenance

        - Remove all reagent containers from stainer

        - Run fresh methanol through all stainer tubing lines for 5 minutes

- Place all cannulas in a container of clean methanol

- Lift operating lever to Prime position

- Run methanol through tubing until methanol running through is clear

- Remove cannulas from methanol and continue to prime until all methanol cleared from lines

- Wipe methanol from platen

        - Empty reagent waste container into Hazardous waste Container

        - Check stainer level using round level placed on buffer mixing area

        - Replace towels under and in front of stainer

        - Clean platen (mixing grooves and rinse trough)

  - turn off  stainer

  - flood platen with methanol

  - wipe methanol off platen working from right to left

  - wipe all raised bar areas on platen and most importantly those just past wash opening

  - turn stainer back on

        - Clean drain opening with methanol and q tip

        - Check reagent volume (there needs to be sufficient volume for a day run), replace in stainer, and insert cannulas. Discard used stain into Hazardous waste container and dispose of  empty containers in hazardous waste bin

        - If reagent has low volume, Replace stain pack and label with opened date

        - Pool any existing rinse reagents together and do the same with the buffer reagents. DO NOT POOL STAIN together. Pierce opening for cannulas with small screwdriver (NOT Cannula)

        - Prime reagents until all reagents are visible on platen

        - Resolve any reagent dispense problems if occur

        - Prepare “test’ smears for each stainer and label TR,TL and TB

        - Stain test smears

        - Review test smears for satisfactory stain

  - None to minimal precipitate

  - No water marking on rbc

- Platelets stain

- Rbc borders are clear

- Rbc stain color is red (not orange or deep red or slightly purple)

- Wbc stain color is normal

- Satisfactory stain for entire smear

- Record smear review in smear review log in diff area

- If satisfactory simply record

- If unsatisfactory, resolve stainer problem; refer to stainer maintenance/troubleshooting procedures. Repeat test smears after problems resolved

        - Once test slides deemed satisfactory patient slide can be stained, if test slides are 	unsatisfactory routine slides cannot be stained until stainer problem resolved.

B.   Monthly Maintenance – to be performed the 1st week of each month

      - Replace 1/2/3 upper tubing and 1/2/3 lower tubing at same time

        - Remove tubing from top of stainer, flush nipple and under tubing with methanol

        - Reattach upper tubing

        - Remove under tubing from nipples on underside of platen

      - Check cannulas for bends

      - Attach syringe to nipple and flush up onto platen

      - Gently rout reagent opening in platen with 3 ml needle (do not push methanol back into stainer)

      - Reattach tubing and re-flush stainer with methanol

      - Note any problems, resolve and parts changed in stainer problem/corrective action  section log

C.  As Needed Maintenance

- Check stainer settings vs. reagent volume ratio

  - Reagent ratio should be 1(stain): 3 (buffer): 5 (rinse) – to be checked only after all monthly maintenance has been performed

  - Detach all lower tubing from only from the metal platen spout underneath the staining platform.

  - Label 3 conical tubes (with marked volumes on them), 1, 2, and 3 and cover the top with parafilm, poke a small hole in the top of each parafilmed conical tube.

  - Feed the lower tubing from the stain into the #1 conical tube

  - Feed the lower tubing from the buffer into the #2 conical tube

  - Feed the lower tubing from the rinse into the #3 conical tube

  - Be sure tubing is all the way into each conical tube, to prevent any leaks

  - Place 10 new slides through the stainer

  - When all slides are completed remove the conical tubes 1, 2, and 3 and measure the volumes in the conical tubes.

  - Adjust the pump volume of reagent as needed

    - (-) Counterclockwise = less volume

    - (+) Clockwise = more volume

    - Stainer and buffer knobs: each evenly labeled # increment changes the volume by .1 ml per 10 slides (0.01ml/slide)

    - Rinse knob: each evenly labeled # increment changes the volume by .5 ml per 10 slides(0.05ml/slide)

        - After adjusting the volumes to a 1:3:5 ratio place the 10 slides through the stainer again to check the volume. Check the volume for tube 1,2,3 the adjusted volume matches the actual volume in the tube.

        - Place cannualas back in the appropriate resgents and make a test smear to check stain quality.

        - Record stainer settings in stainer maintenance log book.

- REAGENTS AND MATERIALS:

3 x1 glass slides

Permanent marker

Hematek stainer upper and lower tubing

Siemens Hematek Stain Pack Wright stain

Methanol

Beakers

Gauze

3 cc syringes

Pump Tube set (upper)

Under platen tubing (lower)

Cannulas

Q-Tips (Cotton Tipped Applicator)

Small Screw-driver

- REFERENCES:

Hematek slide stainer manual

Clinical Hematology laboratory procedure manual

VI. 	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 10-2005.

H-2      This procedure was revised by N. Ortoli-Drew 1-2011.

H-3      This procedure was revised by N. Ortoli-Drew 7-2011.

H-4	This procedure was revised by E. Degano 3-2013.