I.   		PRINCIPLE:

The presence of fat in stool and body fluids can be confirmed microscopically. On 		mixing the sample with Oil Red 0, neutral fats will appear as red or orange droplets.

II. 		SPECIMEN (S):

Acceptable specimens for analysis with stain are stool, urine, and sputum, pleural and peritoneal fluids. If unable to analyze specimen directly, refrigerate specimen until time of analysis. Specimens must be free of preservatives (EDTA anticoagulant

is acceptable in the case of body fluids) collected without any contamination with

mineral oil, laxatives or diluent fluids.

III.  	MATERIAL / REAGENTS:

Alcohol wipes

3 x 1 glass slides

22 x 22 coverslips

Sterile saline

Oil Red 0 stain

Positive and negative controls Microscope

IV.    	QUALITY CONTROL:

Negative control 1 drop of saline and 1 drop of Oil Red 0 stain

Positive control 3 x 1 glass slide pressed to forehead

Ordered in Total Soft Q.C. and resulted before manual entry of patient results.

V.    	PROCEDURE:

Note: all materials, handling and microscopy must be free of oil and care to 			prevent any fat or oil contamination.

A.	Stool, sputum, fluid fat analysis:

1.	Alcohol clean all glass slides and covers lips and allow to air dry

2.	Place 1 drop of specimen on slide (for stools prepare an emulsion of 			stool and Sterile saline), mix with an applicator stick

3.	Add 1 drop of Oil Red 0 stain to specimen and mix

4.	Prepare positive and negative control slides

5.	Coverslip all preparations and allow to stand at room temperature for 			minimum of 5minutes

6.	Examine control and specimen slides for the presence of red-orange fat 			droplets

B.	Urine fat analysis:

1.	Centrifuge an aliquot of urine

2.	Refrigerate specimen for 10 minutes to induce globule formation, if fat 			is present.

3.         Place a portion of urine from meniscus on one slide and portion of the                   			urine sediment on another slide.

4. 	Prepare positive and negative control slides.

5.	Add 1 drop of Oil Red 0 stain to both urine slides, and mix.

5.	Coverslip all slides and allow to stand at room temperature for a 				minimum of 5minutes

6.	Examine specimen and control slides for the presence of red-orange fat 			droplets.

VI.   	REPORTING:

Slides should be read by the technologist and only positives confirmed by a

Supervisor.

No red-orange droplets seen	Negative

Red-orange droplets seen	Positive

(Semi-quantitate slight, moderate, marked)

Results manually entered in Resulting Worklist (Misc.) - see Soft Procedures.

VII.   	SIGNIFICANCE

A.	Stool fat:

1.	Enterogenous- celiac disease, spure, hypogammaglobulinemia, blind 			loop syndrome, various acute diarrhea syndromes and lipoidystrophy

2.	Pancreatogenous- infectious pancreatitis, cystic fibrosis of the 					pancreas, malignancies, lack of specific pancreatic enzymes

3.	Hepatogenous- various liver disorders, obstruction of bile ducts, lack 					of bile 	salts

B.       	Pleural and sputum fat

Seen in fat embolism

C.	Urine fat:

1.	This may be caused by the release of bone marrow into the circulation 		due to 	bone ftacture or to bone deterioration. It can also be seen in 		chyluria, amyloidosis, glomerulonephritis, lupus and lipid nephrosis.

VIII:	 REFERENCE:

Davidshon and Henry, Clinical Diagnosis and Management by Laboratory Methods,                      	18th edition, 1991.

Bronzel, Nancy: Fundamentals of Urine and Body Fluid Analysis, 1994.

Teitz, Norbert: Fundamentals of Clinical Chemistry, 1970.

IX.  	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 9-1997.

H-2	This procedure was revised by N. Ortoli-Drew on 8-30-2005.

H-3      This procedure was revised by D. Fico on 8/2011.

Appendix  D

Oil Red O Reagent-   Used to stain Neutral Fats in Stool Samples:

Preparation:   1 gm Oil Red O mixed with 100 ml Propylene Glycol at 37*C

Chemical name- Propylene glycol 1,2- propanediol

Note: Flash point – 107*C

- Heat to 56-60*C until dissolved.

- Filter at 60*C through Whatman #2

- Cool Overnight

- Refilter  through suction filter.

Stability:         Indefinitely

Reference:  Clinical Diagnosis and Management

Davidsohn and Henry Laboratory Investigation

Vol. 5 #4 July- August 1956

Reagent needs to be properly labeled