I.	PRINCIPLE:

The activated partial thromboplastin time (aPTT) is a screening procedure for  	deficiencies in coagulation factors of the intrinsic coagulation pathway, factors VIII, IX, 	XI, XII, Fletcher factor and Fitzgerald factor. Severe deficiencies of fibrinogen and 	factors II, V and X can also be detected by the aPTT. It is a useful and effective method 	for screening patients with a 	bleeding tendency, for evaluating the effect of therapy in 	procoagulant disorders and as the basis for several specific coagulation factor assay 	procedures. The aPTT is widely advocated as a test for monitoring and regulating heparin 	therapy. The presence of non-specific inhibitors, such as the lupus anticoagulant, may 	prolong the aPTT; this effect is variable and is generally recognized as being related to 	the nature and level of the inhibitor present.

The aPTT consists of recalcifying plasma in the presence of a standardized amount of 	platelet-like phosphatides and an activator of the contact factors of the intrinsic 	coagulation pathway. The aPTT is the time in seconds, required for a fibrin clot to form.

II.		SPECIMEN:

A.	Type:

Mix nine parts of freshly collected blood with one part of 3.2 sodium citrate

anticoagulant. Invert the tube gently three or four times immediately after

venipuncture to ensure proper mixing of blood and anticoagulant. A syringe or

evacuated tubes may be used for collection (blue top). If an evacuated tube system 			is used, the coagulation sample should be the second or third tube collected. If

blood is drawn from an indwelling catheter, the line should be flushed with saline

and the first 5 mL of blood discarded or used for other laboratory tests.

The citrate concentration must be adjusted in patients who have hematocrit values 			above 55%. (See separate procedure “Correction for Polycythemia”.)

Specimens that are clotted, collected in the wrong tube or have less than 90%

expected fill of the collection tubes   (QNS) should be rejected. Specimens which

are hemolyzed may be run, but results should be accompanied by a comment that

all coag results may be affected by hemolysis.

B.	Handling Conditions:

The whole blood specimen is checked for clot formation by gentle inversion.

Centrifuge the blood specimen for 10 minutes at 3000 g as soon as possible after

collection. If immediate testing is to be done, the plasma may remain on the

packed cells or separated. To save plasma for further special coagulation testing,

separate plasma in several aliquots to plastic tubes, label, and freeze at –70C.

Frozen plasma samples must be rapidly thawed at 37C and tested immediately

after thawing. If testing is delayed, the sample may be held for 2 hours at 4C

until tested.

III.		EQUIPMENT AND MATERIALS:

A.	Equipment:

1.	BCS  Coagulation Analyzer

2.	Cuvette Rotors

3.	Waste, rinse and disinfectant containers

B.	Materials:

1.	Dade Actin FSL Activated PTT Reagent

2.	Calcium Chloride, 0.025M

3.	Control material: Dade Ci-Trol Coagulation Controls, Levels 1and 3

4.	Preservative-free distilled or deionized water

5.	Washing Solution

6.	70% Isopropyl alcohol as cleaner

IV.	PREPARATION:

A.	Dade Actin FSL Activated PTT Reagent -purified soy and rabbit brain

phosphatides in 1 x 10-4 M ellagic acid with added buffer , stabilizers and 	preservative.

1.	The reagent is liquid and ready for use.

2.	Mix by inversion immediately before use.

3.	Store at 2-8C when not in use. Do not freeze.

4.	Stable for 48 hours on board the BCS (rack 1 or 2).

- B.	Calcium Chloride (0.025M)

1.	The reagent is liquid and ready to use.

- 2.	Store at 2-8C when not in use.

- 3.	Stable for 48 hours on board the BCS (rack 1 or 2).

C.	Dade Ci-Trol Coagulation Control, Level 1 (Normal) -Lyophilized

preparation of fresh human, normal citrated plasma. Used for Quality Control.

D.	 Dade Ci-Trol Coagulation Control Level 3 (Abnormal)-Lyophilized

preparation of fresh  human, abnormal citrated plasma. Used for Quality

Control.

1.	Reconstitute Ci-Trols with 1 mL distilled or deionized water

2.	Restopper vial and allow to stand until dissolved.

3.	Invert gently to mix. Do not shake.

Stability after reconstitution: Ci-Trol is stable for 8 hours stored at 2- 8C.

WARNING: Human source material. Treat as potentially infectious. The aterials

from which these products have been produced were tested and found non-

reactive for the presence of HBsAg and antibody to HIV. No known test method

can offer complete assurance that hepatitis B virus, HIV or other infectious agents 			are absent. Therefore, all human blood-based products should be handled with

proper laboratory practices using appropriate precautions.

V.	PERFORMANCE PARAMETERS:

Indication of deterioration: Normal plasma or controls will show deviations in results

from the established laboratory range.

VI.		QUALITY CONTROL:

A.	Dade Ci-Trol Control Level 1-Normal Control

B.	Dade Ci-Trol Control Level 3 - Abnormal Control

1. 	Controls should be tested once each 8 hour shift, and after reagent

changes/additions. Request controls in Soft, referring to Soft bench manual for Total QC as needed.

2.	Controls should be run in the same manner as the test samples.

3.	Control tolerance limits--the range is calculated based on 2-21/2

standard deviations (SD) from the mean control value.

4.	Corrective action when tolerance limits are exceeded:

a.	Rerun out-of-range control material.

b.	Check reagents/ change if expired

c.	Check instrument performance (eg: pipettor function)

d.	Document actions taken to identify and correct the problem before 					reporting any patient data.

VII.	PROCEDURE – STEPWISE:

### A.	Instrument Set – Up:

# 1.	CALIBRATION

# Standard Curve (done only with yearly rollover of reagent)

The reference curve is entered manually as follows:

a.	Open the calibration dialog on the main screen

b.	Select “New” to request a new reference curve

“New reference curve” diaglog appears

c.	In section “Create a reference curve for this assay” choose assay 			d.	PTT.AFSL.CA (or aPTT.AFSL.570.CA) and assure that at least 				one lot of Actin FSL and the standard “Input” is listed in section “-				--for these reagent lots”. Select “Enter curve”

e.	In the following dialog click on the button “New” twice

Four boxes appear to enter the reference curve values

f.	Enter the CA reference curve values on the left side and the BCS 				reference curve values on the right side. Do not press enter until all 			values are typed in.

g.	“Save” the new reference curve. The dialog closes automatically.

h.	Close the screen “New reference curve”

i.	In the calibration dialog a valid reference curve for assay 					aPTT.AFSL.CA (or aPTT.AFSL.570.CA) is listed in the “Curve 				overview”.

(1.)	The values are:

(a.)	 CA	   BCS

(b.)	11.0        7.0

(c.)	121.9      160.0

# B.	Loading Patient Samples on the BCS:

1.	AUTO HOST:

a.	Load bar-coded samples into any sample rack with the bar-code 					label visible through the window of the rack position.

b. 	Insert the rack into any open lane.

c.	Push the rack until you feel it stop

d. 	Open the JOB LIST icon.

e. 	The JOB LIST screen displays the samples loaded with a small 					green analyzer next to the sample ID. The small green analyzer 					icon indicates 	the sample is currently loaded in the processing area

f. 	The test requests will automatically be picked up from the host.

2.	ASSIGNING A STAT TEST REQUEST:

a.	 Prior to assigning test requests, double click on the sample ID 					number

b. 	Click on the STAT box.  The box will be marked with a “X”.

c.	Confirm the entry by clicking OK

d. 	Locate the test requests for each test required for each STAT 					sample listed.

e. 	Double click on the box  that intersects between the sample 					identification number and the test column.  A red “X” appears in 					the intersection.  The “X” denotes that a test request for this 					sample is in process.

C.	Non – Bar - Code Samples without Host (eg: for running in sample cups without a barcode label):

1. 	Load the samples into the sample racks into consecutive positions.

2.	Click on the LOADING icon.

3.	Click on the number that corresponds to the sample rack (the rack number 				is highlighted)

4.	Click on the Select rack button (a picture of the sample rack is displayed)

5.	Click on the position number of the sample rack in which the sample has 				been placed.

6.	Click on the Sample ID box

7.	Enter the sample identification number.

8. 	At the question “Is the sample STAT?” answer either yes or no.

“Yes” answer will result in a red “X” appearing in the box.

“No”---proceed to the next step.

9. 	To request Individual Assays:

a. 	Locate the assay box

b.	Click on the assay

10. 	Click on the INSERT button and the sample ID appears in sample 					position of the rack and the next position of the sample rack is highlighted.

11. 	To request another:

a. 	Sample in the same rack:   Repeat steps 6-11

b.	Sample Rack:  Click on OTHER RACK button and repeat steps 					3-11.

12. 	Insert the sample racks into any open lane (An open lane is one without 				the LED illuminated or blinking.

13. 	Push the rack until you feel it stop.

14. 	Repeat steps 12 and 13 until all sample racks have been loaded.

15. 	The JOB LIST screen displays the samples loaded with a small green 				analyzer next to the sample ID.

The small green analyzer icon indicates the sample is currently loaded in 				the processing area

A red “X” appears in the intersection between the sample ID number and 			the test column and denotes that a test request for the sample is in 				progress.

16. 	Sample processing begins automatically.

D.	Loading and Programming Controls on the BCS:

1.	Place control bottles into the 5 ml reagent rack with the bar-code label on 				the bottle visible through the window of the rack position.

2.	Insert the rack into any open lane from 4-14 (an open lane is one that does 				not have a blinking or illuminated LED.)

3.	 Push the rack until you feel it stop

4.	Click on the Control Journal icon.

The small green analyzer icon indicates that the Control is currently 				loaded in the processing area

5.	Double click on the box that intersects between the control identification 				and the test column. A red “X” appears in the intersection and denotes that 			a test request for the control is in progress.

6. 	Control processing begins immediately.

7.	Repeat step 5 until all assays for each level of control have been requested.

VIII.	RESULTS MANAGEMENT (AUTO REPEATING):

If the initial APTT result on a given sample is >120 seconds, or if no clot is detected, the 	BCS instrument will automatically rerun a “PTT long” test, allowing that sample extra 	time for clot formation.

If the initial APTT is unreadable due to icteric or lipemic plasma (“minimal absorbance 	too high”), the BCS will automatically rerun a “PTT 570” test, allowing the APTT of that 	specimen to be read on a different wavelength in order to generate a clotting time result.

Lipemic plasmas may have to be ultracentrifuged before running.

Either of these tests can also be requested manually by the operator.

IX.	REPORTING RESULTS:

Results of the aPTT testing should be reported as the aPTT in seconds. These results

should be related to the reference interval for the aPTT. Hemolyzed, lipemic, or icteric

samples should be noted with the result. Results may be auto- or manually verified, or

direct entered if necessary (refer to separate procedures for verification guidelines.)

Prolonged PTT results, and a short PTT of <22.0 sec which do not autoverify should be checked for a clot before being reported.

A.	Reference Interval:

Reference interval values determined for the current Actin FSL lot number (+/- 	2.0 SD)

are 23.2-31.6 seconds.  (May be different in newborn population.)

B.	Procedures for Abnormal Results:

APTT values greater than 120 seconds must

be phoned to the patient’s caregiver as a Critical Value.

X.	PROCEDURE NOTES:

Overall performance of aPTT testing is dependent on reagent and instrument performance.

Acceptable variability (imprecision) should be such that the total day-to-day coefficient of

variation (CV) of the analytic system is less than 5% on the same lot of control plasma.

XI.	LIMITATIONS OF THE PROCEDURE:

Plasma collection and storage conditions must be carefully controlled. Avoid small plasma

volumes prior to testing which may cause pH changes in the plasma leading to inactivation 	of specific components of the coagulation system.

Testing may be affected by a number of commonly administered drugs. Short aPTT values

during conjugated estrogen therapy in males and oral contraceptive administration in

females may occur. Heparin, coumadin (warfarin), and thrombin inhibitors may increase

the aPTT. Hemolyzed, lipemic, icteric specimens may also affect results. Blood

clotting factor deficiencies which should produce prolonged aPTT values may be 	compensated and appear normal by elevated levels of one or more different clotting

factors. The presence of active intermediates (eg activated Factor VII) which would

normally shorten the clotting time may also mask conditions that would lead to 	prolongation of the aPTT. Mild or minor deficiencies in several factors may have an

additive effect on increasing the aPTT. Unexpected abnormal aPTT results should be

followed by additional coagulation studies to determine the cause of abnormal results.

XII.	REFERENCES:

1.	Dade ActinFSL Activated PTT Reagent package insert, Dade Behring,

Marburg, Germany September, 1998.

2.	Dade Ci-Trol Coagulation Control, Levels 1, 2 and 3 , Dade Behring,

Marburg, Germany, July, 1998

3.	Control Plasma N and P package inserts, Dade Behring, Marburg, Germany June,

1998

4.	National Committee for Clinical Laboratory Standards. Collection, transport and

processing of blood specimens for coagulation testing and performance of

coagulation assays. Approved Guideline, 2nd Edition. NCCLS Publication H21-

A2. Wayne, PA, 1991.

5.	National Committee for Clinical Laboratory Standards. One-stage prothrombin

time test (PT) and activated partial thromboplastin time test (APTT). Approved

Guideline. NCCLS Publication H47-A. Wayne, PA, 1995.

6.	Dade-Behring BCS Operator’s Manual

XIII.	HISTORY:

H-1 	This procedure was written by N. Kriz on 7-5-2006.

H-2	This procedure was revised by N Kriz on 10/15/2010.

H-3	This procedure was reviewed by N Kriz on 7/7/2011.

H-4	This procedure was reviewed by N Kriz on 4/24/2012.

H-5	This procedure was reviewed by N Kriz on 6/11/2013.