- PRINCIPLE:

The sulfosalicylic acid method for urine protein is used to confirm the dipstick or 	Clinitek Atlas method when the protein is positive and the PH is alkaline. Highly 	buffered alkaline urines may give false positive results if the buffer system of the 	reagent pad is overcome. The sulfosalicylic acid method is a precipitation technique for 	semiquantitative determination of protein which is unaffected by urine PH.

- SPECIMEN REQUIREMENTS:

Fresh random urine tested within 2 hours of collection. If testing will be delayed 	specimen should be refrigerated. Specimen must be at room temperature for testing.

- QUALITY CONTROL:

On each shift a sulfosalicyclic method is performed on negative and positive urine controls. See Urine Soft Q.C. Procedure. If control results are not acceptable, patient results must not be reported and a supervisor notified.

- REQUIRED MATERIALS:

3% Sulfosalicylic acid ( 3 gr. sulfosalicylic acid dissolved and diluted to 100 ml with distilled water, stable indefinitely at room temperature)

Abnormal Urine Control, Count-IO-trol III, Myers-Stevens Group

Normal Urine Control, Count-IO-trol I, Myers-Stevens Group

Transfer pipets

Test tubes

# PROCEDURE:

A.	Mix equal parts of 3% sulfosalicylic acid and urine sample or control well.   Let stand at room temperature for 10 minutes.

B. 	Read the resultant turbidity against the native urine turbidity/clarity of the urine 			being tested.

C.	Interpret as follows:

| Result | Appearance | Approx. Protein Conc. |
| --- | --- | --- |
| Negative | No turbidity / no increase in turbidity. | <5 mg/dl |
| 1+ | Perceptible to distinct turbidity but no granulation. | 10 - 50 mg/dl |
| 2+ | Turbidity with granulation but no flocculation. | 200 mg/dl |
| 3+ | Turbidity with granulation and flocculation. | 500 mg/dl |
| 4+ | Clumps or solid precipitate | 1000 mg/dl |

# EXPECTED RESULTS:

In most cases this method will correlate with the dipstick or Clinitek Atlas results. In some cases when urine is alkaline the dipstick method will give a false positive and this method will yield a true negative. If this method agrees with the dipstick as positive the original quantitation from the dipstick (or Atlas) should be reported. If the dipstick (or Atlas) has a positive result and this method is negative, a negative result should be reported.

# LIMITATIONS OF PROCEDURE:

If the urine is cloudy the sample should be centrifuged before performing the 	  		sulfasalicylic acid method.

# REFERENCES:

## 1.	Modern Urine Chemistry, Bayer Corporation, Tarrytown, NY, 1996

IX.	  HISTORY:

H-1	This procedure was written by Natalie Ortoli-Drew on 09-2003.

H-2	This procedure was revised by Natalie Ortoli-Drew on 7-2005.

H-3 	This procedure was revised by D. Fico on 9/2011.