I.	PRINCIPLE:

A plasma deficient in any of the factors in the intrinsic pathway will result in a

prolonged activated partial thromboplastin time (aPTT). Factor deficient 	plasma can be

used in the identification and quantification of the factor deficient in the patient plasma.

A mixture of the deficient plasma and the patient plasma is tested by an aPTT and the

result is interpreted using a reference curve obtained with dilutions of Standard Human

Plasma or a normal plasma pool mixed with the deficient plasma. The percent of factor

activity present in plasma can be determined by the degree of correction obtained with

the aPTT when the test plasma is added to specific factor deficient plasma.

II.		SPECIMEN:

A.	Type:

Mix nine parts of freshly collected blood with one part of 3.2% sodium citrate

anticoagulant. 	Invert the tube gently three or four times immediately after

venipuncture to ensure proper mixing of blood and anticoagulant. A syringe or

evacuated tubes (blue top) may be used for collection. If multiple specimens are

collected,, the coagulation sample should be the second or third tube

collected. If blood is drawn from an indwelling catheter, the line should be

flushed with 5.0 mL saline and the first 5.0 mL of blood or six dead space

volumes of the catheter discarded or used for other laboratory tests. The citrate

concentration must be adjusted in patients who have hematocrit values above

55%. Specimens that are clotted, collected in the wrong tube, or have less than 90% of the expected fill  (QNS) should be rejected.

B.	Handling Conditions:

The whole blood specimen is checked for clot formation by gentle inversion and

observation. Centrifuge the capped blood specimen for 10 minutes at 3000 g as

soon as possible after collection. Patient plasma should be tested within 4 hours.

If immediate testing is to be done, the plasma may remain on the packed cells or

separated. For special coagulation assays, centrifuge for 20 minutes at 3000 g,

then separate plasma into several aliquots in plastic tubes and freeze at -70C until

ready to test. A frost-free freezer should not be used. Frozen plasma samples must 			be rapidly thawed at 37C while gently mixing and 	tested immediately after

thawing. If testing is delayed, the sample may be held for 2 hours at 4C

until tested.

III.		EQUIPMENT AND MATERIALS:

A.	Equipment:

1.	BCS  Coagulation Analyzer

2.	Cuvette Rotors

3.	Waste, rinse and disinfectant containers

4.	 BCS Sample Cups

B.	Materials:

1.	Dade Behring Actin FSL Activated PTT Reagent, B4219

2.	Calcium Chloride, 0.025M-ORHO33

3.	Standard Human Plasma (as calibrator)-(lyophilized)

4.	Control Plasma N—Normal control (lyophilized)

5.	Control Plasma P--Abnormal control (lyophilized)

6.	Factor VIII-, IX-, XI-,  or XII- Deficient Plasma-George King                        					King (frozen plasma)

7.	NERL water

8.	Imidazole-buffered saline (IBS)

9.	Wash Solution

10.	70% Isopropyl alcohol disinfectant

C.	Preparation:

1.	Dade Actin FSL Activated PTT Reagent--purified soy and

rabbit brain phosphatides in 1 x 10-4 M ellagic acid with added

buffer, stabilizers and preservative.

a.	The reagent is liquid and ready for use.

b.	Mix by inversion immediately before use.

c.	Store at 2-8C when not in use. Do not freeze.

d.	Stable for 48 hours on board the BCS at 15C

- 2.	Calcium Chloride (0.025M)

a.	The reagent is liquid and ready to use.

- b.	 Store at 2-8C when not in use.

- c.	 Stable until the date indicated on the vial label.

- d.	 Stable for 48 hours on board the instrument at 15-25C

These products are for in vitro diagnostic use.

3.	Factors VIII, IX, XI, XII Deficient Plasmas—frozen 							1-ml aliquots. Thaw in 37C waterbath approx 5 min

- Coagulation Factor VIII Deficient Plasma is stable 8 hours at 15-	25C

- Coagulation Factor IX Deficient Plasma is stable 8 hours at 15-25C

- Coagulation Factor XI Deficient Plasma is stable 4 hours at 15-25C

Coagulation Factor XII Deficient Plasma is stable 2 hours at 15-25C

4.	Control Plasma N-- Used for Quality Control-Normal

5.	Control Plasma P--Used for Quality Control-Abnormal

IV.	QUALITY CONTROL:

A.	Controls should be tested at least once every 24 hours for any assays run for patient

testing during that interval.

B.	Controls should also be tested upon reagent changes, and after each new calibration curve

C.	Controls should be run in the same manner as the test samples.

D.	Corrective action when tolerance limits are exceeded:

1.	Recalibration may be necessary if control values are outside the target range.

2.	Verify reagent performance.

3.	Check instrument performance.

4.	Document actions taken to identify and correct the problem before reporting any

patient data.

V.	PROCEDURE – STEPWISE:

A.	Instrument Set-Up and Calibration (only if necessary):

1.	Press Calibration Icon and choose New key.

2.	Click the required assay in the for this assay list box.

3.	Select the lot number of reagents and standards on board for the calibration curve.

4.	Press Measure curve key to begin curve calibration.

5.	Place the reagents and standards required for the calibration on the analyzer.

6.	Click OK to confirm the message.

7.	When the reagents and standards have been positioned, the analyzer starts 					calibration.

8.	Acceptable and valid calibration curves will display a BCS symbol next to the 				curve and date.  Unacceptable curves will display an “exclamation point”.

9.	View the curve for validity. Click the Calibration button. Click the desired 				reference curve in the Curve overview column.

10.	Click the Show curve key to view the reference curve. The curve will be 				displayed.

11.	To re-measure one or more points displayed in the graphic, click the point to be 				corrected. The point will be highlighted. Click the Repeat key and close the 				dialog. The point will be re-measured and entered in the graphics.

Rack lanes 4-14 should be used for on board position of deficient plasmas, controls and standards.

Rack lane 1 or 2 should be used for on board position of Actin FSL and Rack lane 1 or 2 for Calcium Chloride

# VI.	LOADING PATIENT SAMPLES ON THE BCS:

A.	Bar-Coded Samples without Host:

1. 	Load bar-coded samples into any sample rack with the bar-code label visible 				through the window of the rack position.

2. 	Insert the rack into any open lane.(4-14)

3. 	Push the rack until you feel it stop

4. 	Open the JOB LIST icon, if not yet open.

5. 	The JOB LIST screen displays the samples loaded with a small green analyzer 				next to the sample ID.

The small green analyzer icon indicates the sample is currently loaded in the 				processing area

6. 	To assign tests requests to the samples that have been loaded follow the steps 				below:

a.	Locate the test requests required for each sample listed.

b. 	Double click on the box that intersects between the sample

identification 	number and  the test column.  A red “X” appears

in the intersection.  The “X” denotes that a test request for this sample is 					 in process.

7. 	Sample processing begins automatically.

B.	Assigning A Stat Test Request:

1.	Prior or after to assigning test requests, double click on the sample ID number

2. 	Click on the STAT box.  The box will be marked with a “X”.

3. 	Confirm the entry by clicking OK

4. 	Locate the test requests for each test required for each STAT sample listed.

5. 	Double click on the box that intersects between the sample identification number 				and the test column.  A red “X” appears in the intersection.  The “X” denotes that 				a test request for this sample is in process.

6. 	Sample processing begins automatically.

C.	Non-Bar Code Samples without Host:

1.	Load the samples into the sample racks into consecutive positions.

2.	Click on the LOADING icon.

3.	Click on the number that corresponds to the sample rack (the rack number is 				highlighted)

4. 	Click on the Select rack button (a picture of the sample rack is displayed)

5. 	Click on the position number of the sample rack in which the sample has been 				placed.

6. 	Click on the Sample ID box

7. 	Enter the sample identification number.

8. 	At the question “Is the sample STAT?” answer either yes or no.

“Yes” answer will result in a red “X” appearing in the box.

proceed to the next step.

9. 	To request Individual Assays:

a. 	Locate the assay box

b. 	Click on the assay

c. 	Repeat the above steps until all assays are requested.

10. 	Click on the INSERT button and the sample ID appears in sample position of the 				rack and the next position of the sample rack is highlighted.

11.	To request another:

a. 	Sample in the same rack: Repeat steps 6-11

b. 	Sample Rack:  Click on OTHER RACK button and repeat steps 3-11.

12. 	Insert the sample racks into any open lane (An open lane is one without the LED 				illuminated or blinking.

13.	Push the rack until you feel it stop.

14.	Repeat steps 12 and 13 until all sample racks have been loaded.

15. 	The JOB LIST screen displays the samples loaded with a small green analyzer 				next to the sample ID.

The small green analyzer icon indicates the sample is currently loaded in the 				processing area

A red “X” appears in the intersection between the sample ID number and the test 			column and denotes that a test request for the sample is in progress.

16. 	Sample processing begins automatically.

VII.	LOADING AND PROGRAMMING CONTROLS ON THE BCS:

A. 	Place control bottles into the 5 ml reagent rack with the bar-code label on the bottle 			visible through the window of the rack position.

B. 	Insert the rack into any open lane from 4-14 (an open lane is one that does not have a 			blinking or illuminated LED.)

C.	Push the rack until you feel it stop

D. 	Click on the Control Journal icon.

The small green analyzer icon indicates that the Control lot number is currently loaded

in the processing area

E. 	Double click on the box that intersects between the control identification and the test 			column. A red “X” appears in the intersection and denotes that a test request for the 			control is in progress.

F. 	Control processing begins immediately.

G. 	Repeat step 5 until all assays for each level of control have been requested.

VIII.	REPORTING RESULTS:

A.	Results of the factor assay testing should be reported in % normal.  These results

should be related to the reference interval for the factor.

1.	Record results in the computer system.

2.	Record quality control values.

3.	Hemolyzed, lipemic, or icteric samples should be noted with the result.

B.	Reference Interval:

1.	Reference interval values for Factor VIII = 50-200%

2.	Reference interval values for Factors IX, XI, and XII = 50-150%

C.	Procedures for Abnormal Results:

1.	Factor levels less than ___40%____ of normal for the first time should be phoned to the patient’s caregiver as a critical value.

2. 	On occasion, a circulating inhibitor such as a strong lupus anticoagulant or—less frequently—a strong inhibitor to a factor other than the one being analyzed, will affect the factor assay results. This is an artifact of the assay system, and the results will appear low (even markedly low) at the 1:10 dilution, somewhat higher at the 1:20 dilution, and normal or near-normal at the 1:40 dilution (as the inhibitor is diluted out). In this case, the highest level obtained should be reported out. In an extreme case, it may be helpful to pre-dilute a patient plasma 1:2 with IBS and rerun the assay, multiplying the final result by 2.

D.	Procedure Notes:

Overall performance of factor assay testing is dependent on reagent and instrument

performance. Acceptable variability (imprecision) should be such that the total

coefficient of variation (CV) of the analytic system is less than 15% on the same

lot of control plasma.

The measuring range is defined by the concentration of the calibrators used and the

extrapolation limits set.

IX.		LIMITATIONS OF THE PROCEDURE:

Partial activation of the coagulation factors (VIII) due to incorrect sample handling can lead to 		falsely elevated single factor determinations

X.		REFERENCES:

1.	Actin FSL package insert, Dade Behring, Marburg, Germany, March 1999

2.	Coagulation Factors VIII, IX, XI XII package insert, Dade Behring, Marburg, Germany,

January 1999

3.	Standard Human Plasma package insert, Dade Behring, Marburg, Germany, January

1999

4.	Control Plasma N package insert, Dade Behring, Marburg, Germany, May 1998

5.	Control Plasma P package insert, Dade Behring, Marburg, Germany, January 1999

6.	National Committee for Clinical Laboratory Standards. Collection, transport

and processing of blood specimens for coagulation testing and general performance

of coagulation assays. Approved Guideline, 3rd Edition. NCCLS Publication H21-A3.

Wayne, PA, 1998.

7.	Dade-Behring BCS Operator’s Manual

8.	Factor VIII, IX, XI and XII Application Sheets: 11391401E, 11401401E, 11411401E,

11421401E,

XI.		HISTORY:

H-1	This procedure was written by N. Kriz on 07-05-06.

H-2	This procedure was revised by N Kriz on 10/21/2010.

H-3	This procedure was revised by N Kriz on 7/7/2011.

H-4	This procedure was revised by N Kriz on 12/1/2011.

H-5	This procedure was reviewed by N Kriz on 6/11/2013.