I.	PRINCIPLE:

The quantitative determination of free Protein S is accomplished by a one-step 	“sandwich” technique known as enzyme-linked immunosorbent assay (ELISA). The 	internal wall of a plastic microwell is precoated with the first monoclonal antibody to free 	protein S. The second anti-free protein S monoclonal antibody that is coupled with 	peroxidase is added to the precoated microwell at the same time as the plasma whose free 	protein S is to be determined. The free protein S of the plasma being tested is 	simultaneously captured on the one hand by the first monoclonal antibody that is 	precoated on the microwell wall, and on the other by the second monoclonal antibody-	peroxidase conjugate (hence, the “sandwich”). Next, the bound peroxidase is revealed by 	its activity in a predetermined time on the TMB substrate. After stopping the reaction 	with a strong acid, the intensity of the color produced is directly proportional to the free 	protein S concentration initially present in the plasma sample.

II.	REAGENTS AND PREPARATION:

A.	Reagent 1:

1.	Pouch containing 2 strips, each of 16 wells coated with the first 					monoclonal anti-human free protein S. Allow to stand at room temperature 			30 minutes before opening, then use as soon as strips are removed from 				packet.

B.	Reagent 2:

1.	Second monoclonal anti-human free protein S antibody (different from the 			first one) coupled with peroxidase, freeze-dried. Reconstitute one vial of 				this reagent with 2 ml of reagent 4 and allow to stand 30 minutes at room 				temperature. Vortex before use. After reconstitution, Reagent 2 is an 				irritant, and is stabile for 4 hours at 15-20o C.

C. 	Reagent 3:

1.	TMB solution, ready for use. Allow to stand 30 minutes at room 					temperature, then use immediately.

D.	Reagent 4:

1.	TRIS buffer, ready for use. Allow to stand 30 minutes at room 					temperature before use. Stabile 15 days at 2-8o C.

E.	 Reagent 5:

1.	20-fold concentrated washing solution. Dilute 1:20 with distilled water for 			use. For 2 strips, mix 15 ml reagent 5 + 285 ml distilled water.

NOTE: The presence of crystals in 5 will not affect its quality. If necessary, warm at 37o C until all crystals are dissolved. Do not use any of the reagents if they show obvious signs of contamination.

F.	 Reagent 6:

1.	Lyophilized human plasma containing a free protein S level of 					approximately 100% (for exact value see Assay Value insert provided 				in each kit). This is the “t” value for assay calibration. Reconstitute the 				vial of reagent 6 with exactly 0.5 ml distilled water and allow to stand for 			30 minutes at room temperature. Vortex  before use.

G. 	Reagent 7:

1.	Lyophilized human plasma containing a known free protein S level (see 				Assay value inserts provided in each kit). This is for assay control. 				Reconstitute the vial with exactly 0.5 ml of distilled water and allow to 				stand at room temperature for 30 minutes. Vortex before use.

H. 	1 M sulfuric acid : (not provided in kit):

1.	Prepare 3M H2SO4 as described in Protein S Total antigen procedure. 				Further dilute as follows: 4.0 mL distilled water + 2.0 mL 3M H2SO4. 				Prepare fresh each run.

III.	PREPARATION OF STANDARD CURVE:

A. 	Prepare a 1:21 (“t”) dilution (most concentrated) by adding 50 ul of reconstituted 			reagent 6 + 1.0 ml reagent 4. Using reagent 4, make doubling dilutions to obtain 			the other points, as follows:

0.5 ml “t” dilution + 0.5 ml reagent 4 = t:2

0.5 ml t:2 dilution  + 0.5 ml 	“      = t:4

0.5 ml t:4 dilution  +  0.5 ml	“      = t:8

0.5 ml t:8 dilution  +  0.5 ml	“      = t:16

NOTE: If, for example, the assay value insert assigns the calibrator a “t” value of 106%, 	then the points of the standard curve would be:

t	=	106%

t:2	=	 53%

t:4	=	 26.5%

t:8	=	 13.3%

t:16	=	   6.6%

IV.	PREPARATION OF PATIENT OR CONTROL PLASMAS:

Control plasma (reagent 7) is run in duplicate at 2 dilutions, prepared in a manner 	identical to the first 2 points of the standard curve. Patient plasmas should be thawed in 	waterbath for at least 15 minutes, and should be run in duplicate whenever possible, with 	dilutions prepared in the same manner as  the standard curve.

A.	TEST PROCEDURE (reagents have been reconstituted and plasmas 				diluted):

1. 	Cut open 1 pouch of reagent 1; insert into plate frame, being sure to 				match notches in frame with those in the strips. Include the following for 2 			strips:

a.	reagent 4 TRIS buffer, as blank, in duplicate

b. 	each calibrator point, in duplicate

c. 	control at 2 dilutions, each in duplicate

d. 	patient samples in duplicate (8 patients can be run conveniently)

2.	For antigen-and-immunoconjugate immobilization step, pipet into each 				well:

a. 	50 ul of reconstituted reagent 2 (using multi-channel pipettor, set 				to 50 ul) then pipet:

b.	200 ul of each reagent blank, calibrator, control, or test sample.

Cover the strips and incubate for exactly 1 hour at room temperature. Do not allow wells 	to remain in light during incubations (cover with cardboard, paper, etc).

3. 	At 1 hour, wash all wells 5 times with reagent 5, as follows, referring to 				operator’s manual for more details, if needed:

a.	Remove probe/filter from distilled water wash bottle and place 				into flask containing Protein S wash solution.

b. 	Prime 1 or 2 times

c. 	Ensure that program 3 (P3) is in use, and set # of rows to 4.

d.	Place frame with strips onto instrument and press “Start” to begin 					wash procedure.

e.	After the last wash, remove plate, and onto a thickness of paper 				towels, bang out any remaining moisture from the wells, to prevent 			dilution or contamination of the test.

4.	Change the multi-channel pipettor to 200 ul, and, immediately 					following the washing procedure, pipet 200 ul reagent 3 into each well. 				Using a stop watch, incubate this for exactly 5 minutes at room 					temperature. Do not wash wells this time.

5. 	Change the multi-channel pipettor to 50 ul, and immediately following 				the above color development step 4, pipet 50 ul H2SO4 into each well. 				Swirl plate gently.

6. 	Let stand covered for 15 minutes, then read the absorbance of each well on 			plate reader set at 450 nm, as follows, referring to operator’s manual for 				more details, if needed:

a. 	Turn on printer and plate reader. Reader will perform self-test.

b. 	On main menu, press “Read”; select appropriate assay (04 = Prot 					S free)

c. 	Place plate with wells to left; press “Read” key at bottom of key 					pad

d. 	After plate is read, results will print automatically (curve and 					results in both matrix and column form; column can be discarded)

V.	RESULTS:

A. 	Subtract absorbance value of blank from each reading of calibrators, controls, and 		patients’ samples. Using log-log graph paper (provided, or in folder in lab) plot 			the calibrator values (%) on the X-axis and their corresponding absorbance values 		on the Y-axis to obtain a standard curve. Interpolate the absorbance value of each 			control and patient sample from the curve, being sure to multiply by the 				appropriate dilution factor. The 2 results from each sample can then be averaged 			to derive the final free protein S level (%).

B.	To ensure a valid test, control values obtained should be within the range 				indicated on the assay value insert of each kit. If the control is out of range, the 			test run should be considered invalid, and the patient samples repeated.

C.	Detection limit of this assay system is 2% free Protein S. For a result below this 			threshold enter “<2”.

D.	The working range of this assay is 2-130% free Protein S; however, for a given 			lot, the upper limit corresponds to the value for Reagent  6 as stated in the Assay 			Value insert.

VI.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-23-04.

H-2	This procedure was reviewed by N Kriz on 10/21/2010.

H-3	This procedure was reviewed by N Kriz on 7/7/2011.

H-4	This procedure was reviewed by N Kriz on 4/24/2012.

H-5	This procedure was reviewed by N Kriz on 6/11/2013.