I.	PROCEDURE:

Vacutainers for coagulation testing contain 3.2% sodium citrate, intended for a ratio of 9:1 whole blood to anticoagulant. Above hematocrits of 50%, error is introduced by excessive citration and dilution. In cases where a patient's hematocrit is greater than 50%, it is necessary to scale down the amount of anticoagulant proportional to the hematocrit. The following formula is used to determine the appropriate amount of anticoagulant needed:

S= V (100-H)

----------

(640-H)                  		S= amount of anticoagulant

V= total volume of blood and anticoagulant

H= hematocrit

Tubes can be provided by the coagulation laboratory containing the proper amount of 3.2% sodium citrate for a given hematocrit, keeping V at 5 ml as in the standard vacutainer (blue stopper).

(N.B.: 3.8 gm sodium citrate + 100 ml distilled water = 3.2%)

HEMATOCRIT                AMOUNT NaCit       	HEMATOCRIT            AMOUNT Nacit

50			0.42				65			0.30

51			0.42				66			0.30

52			0.41				67			0.29

53			0.40		 		68			0.28

54			0.39				69			0.27

55			0.38				70			0.26

56			0.38				71			0.25

57			0.37				72			0.25

58			0.36				73			0.24

59			0.35				74			0.23

60			0.34				75			0.22

61			0.34				76			0.21

62			0.33				77			0.20

63			0.32				78			0.20

64			0.31				79			0.19

NOTE: Amounts must be multiplied by 3/5 if smaller (2.7 ml) coagulation tubes are used.

II.	REFERENCE:

Komp,D.M. and Sparrow,A.W. Quantitation of secondary fibrinolysis in cyanotic heart disease. J.Pediat. 77:679-682, Oct.1970.

III.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-16-2004.

H-2	This procedure was reviewed by N Kriz on 10/22/2010.

H-3	This procedure was reviewed by N Kriz on 8/23/2011.

H-4	This procedure was reviewed by N Kriz on 10/19/2012.

H-5	This procedure was reviewed by N Kriz on 6/12/2013.