I.	PRINCIPLE:

The quantitative determination of total Protein S is accomplished by a one-step 	“sandwich” technique known as enzyme-linked immunosorbent assay (ELISA). The 	internal wall of a plastic microwell is precoated with the first monoclonal antibody to 	total protein S. The second anti-total 	protein S monoclonal antibody that is coupled with 	peroxidase is added to the precoated 	microwell at the same time as the plasma whose 	total protein S is to be determined. The total 	protein S of the plasma being tested is 	simultaneously captured on the one hand by the first monoclonal antibody that is 	precoated on the microwell wall, and on the other by the second monoclonal antibody-	peroxidase conjugate (hence, the “sandwich”). Next, the bound peroxidase is revealed by 	its action  on  the  TMB substrate. After stopping the reaction with a strong acid, 	the 	intensity of the color produced is directly proportional to the concentration of total 	protein S initially present in the plasma sample.

II.	REAGENTS AND PREPARATION:

A.	Reagent 1--pouch containing 2 strips, each of 16 wells. coated with the first 			monoclonal anti-human total protein S. Allow pouch to stand, unopened, for 30 			minutes at room temperature before using. Do not open until immediately 			before use.

B. 	Reagent 2--second monoclonal anti-human total protein S antibody (different 			from the first one) coupled with peroxidase, freeze-dried. Reconstitute each vial 			of this Reagent 2 with 2 ml of Reagent 4. Allow reagent to stand at room 			temperature for 30 minutes; then swirl gently before use.

C. 	Reagent 3—ready-for-use tetramethylbenzidine (TMB <1%) solution. Allow 			reagent to stand at room temperature (18-25o) for 30 minutes. Then, use 				immediately. (Avoid all contact with metallic compounds and oxidizing agents.)

D.	Reagent 4—ready-for-use phosphate buffer. Allow bottle to remain at 				room temperature for 30 minutes before use. Stability after opening is 15 				days at 2-8C, if free of contamination.

This reagent may cause skin irritation; refer to product insert for complete safety 	information.

E.	Reagent 5--20-fold concentrated washing solution. For 2 strips, mix 15 ml 			reagent 5 + 285 ml distilled water. If necessary, warm first to dissolve any crystals 		present.

F.	Reagent 6--Lyophilyzed human plasma containing a total protein S level of 			approximately 100% (for exact value see Assay Value insert provided in each 			kit). Stable 4 hours at 15-25C.

This is the “t” value for assay calibration. Reconstitute the vial of reagent 6 with 		exactly 0.5 ml distilled water and allow to stand for 30 minutes at room 				temperature. Vortex before use.

G. 	Reagent 7-- lyophilyzed human plasma containing a known total protein S level 			(see Assay value insert provided in each kit). This is for assay control. 				Reconstitute the vial with exactly 0.5 ml of distilled water and allow to stand at 			room terperature for 30 minutes.Vortex before use. Stable 4 hours at 15-25C.

NOTE: The presence of cystals in Reagents 4 and 5 will not affect their 	quality. If necessary, warm at 37o C until all crystals ard dissolved. Do not 	use any of the reagents if they show obvious signs of contamination.

H.	3 M sulfuric acid --(not provided in kit). Add 2.0 ml concentrated H2SO4 to 10.0 			ml distilled water. (Larger amounts can be made, and stored in flask at room 			temperature, in hood.)

III.	PREPARATION OF STANDARD CURVE:

A.	Prepare a 1:101 (“t”) dilution (most concentrated) by adding 20 ul of 				reconstituted reagent 6 + 2.0 ml reagent 4. Using  reagent 4, make 					doubling dilutions to obtain the other points, as follows:

1.0 ml “t” dilution + 1.0 ml reagent 4 = t: 2

1.0 ml t: 2 dilution + 1.0 ml 	“      = t:4

1.0 ml t: 4 dilution + 1.0 ml	“      = t:8

1.0 ml t: 8 dilution +  1.0 ml	“      = t:16

NOTE: If, for example, the assay value insert assigns the calibrator a “t” value of 106%, 	then the points of the standard curve would be:

t	=	106%

t:2	=	 53%

t:4	=	 26.5%

t:8	=	 13.3%

t:16	=	   6.6%

IV.	PREPARATION OF PATIENT OR CONTROL PLASMAS:

Control plasma (reagent 7) is run in duplicate at 2 dilutions, prepared in a manner 	identical to the first 2 points of the standard curve. Patient plasmas should be run in 	duplicate whenever possible, and dilutions prepared in the same manner as the standard 	curve.

A.	Test Procedure (reagents have been reconstituted and plasmas diluted):

1.	Cut open 1 pouch of reagent 1; insert into plate frame, being sure to match 			notches in frame with those in the strips. Include the following for 2 strips:

a.	reagent 4 TRIS buffer, as blank, in duplicate

b. 	each calibrator point, in duplicate

c. 	control at 2 dilutions, each in duplicate

d. 	patient samples in duplicate (8 patients can be run conveniently)

2. 	For antigen-and-immunoconjugate immobilization step, pipet into each 				well:

a. 	50 ul of reconstituted reagent 2 (using multi-channel pipttor, set 					to 50 ul)

then pipet:

b.	 200 ul of reagent blank, calibrator, control, or test sample.

Cover the strips and incubate for exactly 1 hour at room temperature. Do not 			allow wells to remain in light during incubations (cover with cardboard, paper, 			etc).

3.	At 1 hour, wash all wells 5 times with reagent 5, as follows:

a.	Remove probe from DW wash bottle and insert into wash solution 					flask.

b. 	Set instrument to program 3 (P 3) using select pad.

c. 	Press Prime pad to fill lines with wash solution.

d. 	Place frame with strips onto instrument.

e. 	Ensure that number of rows to be washed is set at 4.

f. 	Press start to begin wash procedure.

After the last wash, remove plate, and onto a thickness of paper towels, bang out 			any remaining moisture from the wells, to prevent dilution or contamination of 			the test.

4.	Change the multi-channel pipettor to 200 ul, and ,immediately 					following the washing procedure, pipet 200 ul reagent 3 (TMB) into each 				well. Using a stop watch, incubate this for exactly 5 minutes at room 				temperature. Do not wash wells this time.

5. 	Change the multi-channel pipettor to 50 ul, and immediately following 				the above color development step 4, pipet 50 ul 3M H2SO4 into each well. 				Incubate for 10 minutes, using stopwatch, and do not wash wells or 				discard liquid. Gently swirl the plate to mix contents.

6. 	Read the absorbance of each well on plate reader set at 450 nm, as 					follows:

a.	 Turn on printer and plate reader. Reader will perform self-test.

b. 	On main menu, press “Read”; select appropriate assay (02= Prot S 				total)

c. 	Place plate with wells to left; press “Read” key at bottom of key 					pad

d. 	After plate is read, results will print automatically (curve and 					results in both 	matrix and column from; column can be discarded)

V.	RESULTS:

A.	Subtract absorbance value of blank from each reading of calibrators, controls, and patients’ samples. Using log-log graph paper plot the calibrator values (%) on the X-axis and their corresponding absorbance values on the Y-axis to obtain a standard curve. Interpolate the absorbance value of each control and patient sample from the curve, being sure to multiply by the appropriate dilution factor. The 2 results from each sample can then be averaged to derive the final total protein S level (%). NOTE: This is done automatically by the reader, and is stated here, in case a reader is not available. There is a reader in Virology, which can be programmed and used.

B. 	To ensure a valid test, control values obtained should be within the range 				indicated on the assay value insert of each kit. If the control is out of range, the 			test run should be considered invalid, and the patient samples repeated.

VI. 	HISTORY:

H-1	This procedure was written by N. Kriz on 12-17-04.

H-2	This procedure was reviewed by N Kriz on 10/21/2010.

H-3	This procedure was reviewed by N Kriz on 8/19/2011.

H-4	This procedure was reviewed by N Kriz on 4/24/2012.

H-5	This procedure was reviewed by N Kriz on 6/11/2013.