I.	PRINCIPLE OF TEST:

A.	Patient plasma is clotted with an equal volume of 0.025 M CaCl2, and after incubation for 30 minutes at 37oC, 5M urea is added. After 48 hours incubation at 37oC, patients who are deficient in Factor XIII will have a clot that is completely urea-soluble due to lack of cross-binding by adjacent gamma chains of fibrinogen.

II.	REAGENTS:

A.	5m urea--made up just before performing the test--as follows:

3 gm urea into 10 ml distilled water (use template provided for 10-mL volume)

B.	GK Pooled normal plasma (frozen in 1.0-mL aliquots at -70C)

C.	CaCl2 (0.025M)

D.       GK Factor XIII-deficient plasma as positive control (frozen in 1.0- mL aliquots at

-70oC)

III.      PROCEDURE:

NOTE: Use only glass 12 x 75 mm tubes throughout

A.	To 0.5 ml PATIENT plasma add 0.5 ml CaCl2.

B. 	To 0.5 ml NORMAL plasma add 0.5 ml CaCl2 as NEGATIVE control.

C. 	To 0.5 ml XIII-deficient plasma add 0.5 ml CaCl2 as POSITIVE control.

D. 	Place all tubes in 37oC waterbath and note time when samples have clotted.

E.	At 30 minutes after ALL patient and control samples have clotted, add 1.0 ml 5M 			urea solution to each tube. The plasma clots should float upwards; if they do not, 			agitate gently with applicator stick. Incubate overnight at 37oC.

F. 	Observe all tubes after 48 hours incubation and note whether clot in each tube has 			dissolved or remained stable.

1. 	Clot dissolved = deficiency of FXIII = ABSENT

2. 	Clot remains = presence of FXIII = PRESENT

IV.	NOTES:

A.	Results are valid only if the XIII-deficient (POSITIVE) control has been completely 			dissolved and the NORMAL (NEGATIVE) control has remained stable.

B. 	This test CAN be done on plasma which has been thawed and refrozen.

REF: Duckert,F. et al. Thromb.Diath. Haemorr.  5:179, 1960.

V.	HISTORY:

H-1	This procedure was written by N. Kriz on 12-22-2004.

H-2	This procedure was revised by N Kriz on 10/22/2010.

H-3	This procedure was revised by N Kriz on 8/23/2011.

H-4	This procedure was reviewed by N Kriz on 6/26/2012.

H-5	This procedure was reviewed by N Kriz on 6/12/2013.