I.		Purpose:

If the laboratory uses more than one instrument to test for a given analyte, the instruments must be checked at least twice per year for correlation of results.

II.		Routine Coagulation Procedure:

- Patient plasmas are preferred for comparisons of routine coagulation testing on the BCS instruments 1 and 2 (PT/INR, PTT, fibrinogen, d-dimer).

- For comparison of routine control results from BCS (1) and BCS (2), submit statistical data on line at Siemens.com (see separate procedure for this). This submission should be done monthly, and the appropriate reports printed out at that time. Refer to monthly peer group analysis on those reports for interpretation of the QC data.

- To compare sets of patient results, enter the appropriate patient values for a given parameter into an Excel spreadsheet. Plot the data and print the regression data (R value) on the curve.

III.		Special Coagulation Procedure:

- STA Compact (1) SN 8041308 is used for the following tests:

  - LMWH

  - ARIX

  - UFH

  - AT activity

  - Protein C activity

  - Protein S activity

  - Alpha-2 AP

  - Plasminogen

  - APC

- STA Compact (2) SN 3104095 is used for the following tests:

  - dRVVT

  - STAclot LA (phospholipid neutralization)

  - vwFag

- BCS (3) SN 161792 is used for Factor assays; BCS (1) 261818 is used as back-up for Factor VIII assays in the event of downtime for BCS (3).

- No interinstrument comparisons have to be done for the Stago instruments.

- Interinstrument Factor VIII comparisons will be done on BCS (3) and (1).

- IV.		Tolerance Limits:

- For comparison of patient results (PT/INR, PTT, and fibrinogen) from BCS (1), (2), and (3), the R value should be >0.85.

V.		History:

H-1	This procedure written (as “Tolerance Limits for Agreement of Inter-Instrument Parameters”) by N Kriz 1/28/2005.

H-2	This procedure was revised (to “Comparability of Instruments and Tolerance for Agreement of Inter-Instrument Parameters”) by N Kriz 11/4/2010.

H-3	This procedure was reviewed by N Kriz on 8/23/2011.

H-4	This procedure was revised by N Kriz on 10/19/2012.

H-5	This procedure was reviewed by N Kriz on 6/12/2013.