I.	PRINCIPLE:

Lupus anticoagulants (LA) are autoantibodies against negatively-charged 				phospholipids or complexes of phospholipids with either beta-2-glycoprotein 1 or 			clotting factors, such as prothrombin. They occur in various conditions, especially 		autoimmune diseases, and are now considered to be a significant hypercoagulable 			risk factor in patients with otherwise unexplained thrombosis and are often 			present in  women who have recurrent fetal loss. LA have traditionally been 			detected using phospholipid-responsive tests such as APTT, kaolin clotting time, 			and dRVVT where they have an anticoagulant effect.

Russell’s viper venom present in LA-Screen reagent initiates plasma clotting by directly 	activating Factor X. LA antibodies prolong the LA-Screen clotting time. The normal 	reference range for LA-Screen may, by itself, rule out the presence of LA.  LA-Confirm 	reagent is similar to LA-Screen but contains a higher phospholipid concentration. If LA 	is present in a patient’s plasma, the extra phospholipid counteracts the LA antibody and 	largely corrects the clotting time as compared to the LA-Screen result.

A   patient LA-Screen/ patient LA-Confirm   ratio is reported, with LA-negative being 	<1.20.

II.	REAGENTS:

A.	LA-Screen reagent (Gradipore) - reconstitute each vial used with 2 ml distilled water. Mix by gentle inversion to ensure complete re-suspension of the lyophilized material. May be frozen at -70C for 1 month in original vial for one additional use.

B.	LA-Confirm reagent (Gradipore) - reconstitute each vial with 2 ml distilled water. Mix by gentle inversion to ensure complete re-suspension of the lyophilized material. May be frozen at -70C for 1 month in original vial for one additional use.

C.	George King Lupus Anticoagulant Positive plasma —frozen at –70C; thaw just 		prior to use.

D.	George King pooled normal plasma —frozen at –70C; thaw just prior to use.

E. 	LA-1/LA-2 controls—Lyophilized normal and elevated controls, same as used for STAclot LA testing

III.	PROCEDURE:

A.	Prepare reagents as described above. Allow to stand at room temperature 20 			minutes.

B.	Thaw one vial of GK lupus positive plasma and 2 vials of GK pooled normal 			plasma. Place the positive LAC plasma into a sample microtainer. In another 			sample microtainer make a 1:4 dilution (low positive) as follows: 0.2 ml LA 			positive + 0.6 ml PNP. Place the remaining pooled normal plasma into a reagent 			microcup.

C.	Swirl gently and load reagents onto STA Compact using the following designations:

1.	DRVVT-S       DRVVT-SCR        2 (ml)            8hrs)           1234(lot#)*

2.	DRVVT-C      DRVVT-CON       2 (ml)            8(hrs)          1234(lot#)*

3. 	NORMAL	NORMAL	       1(ml)	        8(hrs)	  2222(lot#)

*These reagents require a stirbar, and therefore must be placed in a reagent well which has a stirrer under it. To accomplish this, put reagent into a microcup, add a stir bar, and place microcup into a holder which is wrapped in parafilm to allow it to fit in the appropriate stirring well.

D.	Allow reagents/controls to equilibrate on instrument for 10 minutes, and then 			request control for NORMAL plasma, referring to operator’s manual for complete 		details, if needed. Control should fall within specified values; calculate the LA-			Screen/LA-Confirm ratio and record normal control values on the log sheet.

E.	When the normal plasma control has run, load the GK LA+ and 1:4 dilution and the LA-1 and LA-2 controls in microtainers onto the Compact as patient samples, and request DRVVT-S (LA screen) and DRVVT-C (LA confirm) and RVVRATIO for each. Assure that the ratios for these control results are within the acceptable ranges. Record these results on the control log sheet.

F.	Thaw patient samples and place on STA Compact instrument. Request DRVVT-S 	(LA-Screen) and DRVVT-C (LA-Confirm) and RVVRATIO on all samples.

G.	On the Resulting Worlist, enter “.” for DRVS and for DRVC and enter the RVVRATIO result for DRVR; verify and save entries.

Normal RVVT ratio = <1.20

IV.	NOTES:

A.	Every sample which has a Lupus anticoagulant work-up requested MUST have a PTT reported on it, in addition to the RVVT. Check to be sure PTT result is in Soft.

B.	In the case of a specimen with prolonged PT and PTT, it is essential to rule out the presence of a thrombin inhibitor (an alternate form of anticoagulation when the patient has exhibited heparin-induced thrombocytopenia). Check with the floor/ physician to determine the anticoagulant status of the patient, and if a thrombin inhibitor (argatroban, hirudin, lepirudin, etc) is present, testing for lupus anticoagulant cannot be done until the anticoagulant has been stopped, since any clotting test will be affected.

C.	LA-Screen and LA-Confirm (dRVVT) will be run on every LAC request, in order 		to generate the RVVRATIO, including those with coumadin and/or Lovenox.

D.	Since the reagents for LA-Screen and LA-Confirm contain polybrene, specimens 			containing therapeutic levels of heparin can be run. However, the mixing study 			with PNP will not be performed on heparinized samples.

- The following flow chart for handling lupus anticoagulant testing is included for information only. STAclot LA is run on all lupus anticoagulant samples:

## PTT normal		      		 PTT elevated

                                                   		 

RVVRATIO                                          ______polybrene PTT_________

n↓         abn↓                                                ↓        		 ↓

↓            ↓                                  corrects (hep)*     	does not correct (no hep)

no LAC	       LAC+		      ↓	    		      ↓

RVVRATIO        	_________PTT mix_________

n  ↓     abn ↓        corrects (defic)		does not

correct (inhib)

no LAC      LAC+             ↓ 				↓

RVVRATIO		   RVVRATIO

n ↓             abn↓                     n↓             abn↓

↓                 ↓

no LAC            LAC+   STAclot LA               LAC+

?factor defic         	          <  8 sec       >8 sec

n↓           abn↓

no LAC           LAC+

? factor inhib

* If PTT does not correct completely to normal, or is too prolonged to be able to 	demonstrate correction, and/or the PT is also prolonged due to heparin, the sample may 	have been drawn through a heparinized line. In this case, the sample should be invalid for 	lupus anticoagulant analysis, and another sample should be requested from the floor. On 	an inpatient, it is also permissible to look for a more current sample which does not 	contain the same large amount of heparin, and perform the lupus anticoagulant testing on 	that sample.

V. 	HISTORY:

H-1	This procedure was written by N. Kriz on 12-21-04.

H-2	This procedure was revised by N Kriz on 10/22/2010.

H-3	This procedure was reviewed by N Kriz on 8/23/2011.

H-4	This procedure was reviewed by N Kriz on 6/26/2012.

H-5	This procedure was reviewed by N Kriz on 6/12/2013.