- I.	PRINCIPLE:

The Rotachrom Heparin (Cat #0612), is intended for the quantitative determination

of the plasma levels of low molecular weight heparin (LMWH) by the measurement of their anti-Xa activity in a competitive system using a synthetic chromogenic substrate. This assay can also be used to measure the anti-Xa activity of Fondaparinux sodium (Arixtra), a synthetic and specific inhibitor of activated Factor X (Xa). Only Fondaparinux can be used to calibrate this anti-Xa assay. The compound can reduce thrombin generation in vivo via inhibition of factor Xa and does not inactivate thrombin (IIa).	The normal function of a molecule of factor Xa, as soon as it appears in plasma, is to cleave its natural substrate, prothrombin, to generate thrombin, the enzyme responsible for the fibrin clot formation.  In the presence of heparin, competition occurs between this mechanism and the inhibitory mechanism exerted by the heparin-antithrombin III complex, this inhibition being largely responsible for the action of heparin.  The proposed method is a one-step reaction based on a similar principle: as soon as factor Xa is added to the plasma-substrate mixture, two reactions take place simultaneously:

-  hydrolysis of the substrate by factor Xa

-  inhibition of factor Xa by the heparin-antithrombin III complex

After the necessary period of time for the competitive reaction to reach equilibrium, the quantity of pNA that is released is inversely proportional to the concentration of heparin present in the plasma.  Arixtra is used as a prophylaxis against thromboembolic events following hip and knee replacement surgery.

- II.	SPECIMEN:

A.	Citrated blood 9:1 (blood to anticoagulant) 3.2% sodium citrate.

No other anticoagulant is acceptable.

B.   	Centrifugation:	20 minutes at 3000 g - room temperature

C.   	Plasma storage:	  2 hours for 20 C

6 months at -70 C.  Frozen plasma should be

thawed only once at 37 C for 5 minutes.

D.   	Unacceptable specimens: Samples that are QNS or clotted. Hemolyzed samples 			may be run, but hemolysis should be noted and commented that results may be 			affected.

- III.	REAGENTS AND EQUIPMENT:

- A.	Centrifuge

B.	Pipettes

C.	Pipette tips

D.	STA/STA Compact

E.	Cuvette roll (contains 1,000 cuvettes)

- IV.	REAGENTS:

A.	Arixtra Calibrators:

Obtain a syringe of Arixtra (fondaparinux) from the 	pharmacy. If the concentration 	in the syringe is 2.5 mg/ 0.5 mL it is equivalent to 5000 mg/L.

If the concentration in the syringe is other than 2.5mg/0.5 mL, calculate the mg/L and adjust these dilutions accordingly.

Prepare dilutions for calibration as follows:

0.5 ml of 5000 mg/L + 4.95 ml saline = 50 mg/L

0.1 ml of 50 mg/L + 0.4 ml saline = 10 mg/L

(ARIX 1)  0.1 ml of 10 mg/L + 0.9 ml PNP  = 1.0 mg/L

(ARIX 2)  0.5 ml of 1 mg/L + 0.5  ml  PNP  =  0.50 mg/L

(ARIX 3) 0.5 ml of  0.50 mg/L + 0.5 ml PNP  = 0.25 mg/L

(ARIX 4) 0.5 ml of  0.25 mg/L + 0.5 ml PNP  = 0.125 mg/L

(ARIX 5)                                      0.5 ml PNP = 0 mg/L

B.	Controls:

Two plasmas containing different levels of anti-Xa activity should be used:

A  0.50 mg/L and the 0.25 mg/L  should be run as control material to validate the 	calibration curve.

To prepare controls, use reconstituted LMWH controls (as for Lovenox and Fragmin anti-Xa  assays), but dilute each 1:2 with GK PNP. The LMWH 8 control should yield a result of approx 0.25 mg/L; the LMWH 14 control should yield approx 0.50 mg/L.

C.	Reagent 1:

Substrate: lyophilized chromogenic substrate CBS 52.44, approximately 4.2 	umoles per vial.  Reconstitute each vial with 4.0 ml reagent grade water; discard 	rubber stopper. Let sit 60 minutes at room temperature.  Do not invert vials.  Swirl 	gently; install an STA-reducer and replace the perforated plastic cap on the vial.  	Request the reagent drawer to open through the MAIN MENU under LOADING 	and bar code the reagent.  Place the reagent into the reagent drawer.  	Reconstituted stability on the STA/STA Compact is 7 days.

D.   	Reagent 2:

F.Xa: bovine factor Xa, approximately 1.1 IU per vial.  Reconstitute

each vial with the entire contents of one Reagent 3 vial from same kit; discard 	rubber stopper.  Let sit 60 minutes at room temperature. Do not invert vials.  	Swirl gently; install an STA-reducer and replace the perforated plastic cap on the 	vial. Request the reagent drawer to open through the MAIN MENU under 	LOADING and bar code the reagent.  Place the reagent into the reagent drawer.  	Reconstituted stability on the STA/STA Compact is 7 days.

E.  	Reagent 3:

F.Xa Solvent: Ready to use.  Transfer the entire contents of Reagent 3

into a Reagent 2 vial.

- V.	CALIBRATION (if necessary)

A.	Arixtra calibrators:

1.	Request the product drawer to open.  Manually enter the 	calibrator 			information as stated in the Test Setup. Place each calibrator in 				the product drawer after the information is entered.

B.	All standard dilutions are automatically prepared by the STA/STA Compact by 	diluting the calibrators with Owren-Koller Buffer according to the parameters 	entered in the Test Set-up.

C.   	To order calibration:

1.	Through the MAIN MENU select CAL/CONTROL.

Select CALIBRATION and enter.  Place the cursor on “ARIXROTA” and 		select 	by pressing F1 and then F10 to run the calibration.

D.  	Controls are used by the STA/STA Compact to automatically validate the 	calibration.  (Refer to the QUALITY CONTROL section).

E.   	To examine the calibration curve:

1.	Through the MAIN MENU under CAL/CONTROL select 						CALIBRATION and enter.  Place the cursor on ARIXROTA and enter.  				The calibration curve will appear on the STA/STA Compact screen.

F.   	To verify the calibration curve is valid:

1.	Check that the correlation coefficient (r value) of the curve is 0.90 or 				greater.  If any controls are outside the expected range, the control must be 			re-run.  If the value is still outside the expected range, a new calibration 				must be run along with fresh controls.

G.	To print the calibration curve:

1.	While examining the calibration curve on the STA/STA Compact screen, 			press ESC for options.  Select PRINT.  The 	STA/STA Compact cannot 			print a calibration curve while the STA/STA Compact is running.

- VI.	QUALITY CONTROL:

A.	Controls:

1.

a.	After the reconstitution period, request the sample or the reagent drawer to open through the MAIN MENU under LOADING.  Place the controls into the appropriate drawer, entering information as in the Test Setup.

2.  	QC will automatically be run when a new calibration curve has been 				requested, if it is in place on the instrument

a.	If using a stored curve, the STA/STA Compact will automatically 					run QC when the patients’ samples have been loaded.  The QC can 				also be ordered manually before loading patients’ samples.

3.	Controls placed in the product drawer are monitored automatically by the  			STA Compact.  If any controls are outside the ± 2 SD range, the 				STA/STA Compact will audibly and visually alarm the operator.  				Otherwise, the results can be found in the individual QC files.  Control 			results are automatically filed in the STA Compact QC file.  All results for 		a 24 hour period will be converted to a mean value at midnight.  This 			mean value is used in the statistical data and is plotted on the Levy-			Jennings chart as a daily mean.

a.	Controls run as patient samples should be evaluated by the 					operator, and should closely approximate the expected values.

- VII.	PROCEDURE:

A.	Refer to the START-UP procedure for the STA Compact before running patient samples at the start of each shift.

B.	Run the calibration if needed.

C.   	Request quality control if using a stored curve. (Otherwise, calibrate and run 			controls.)

D.   	Load patients’ samples: Access the SAMPLE drawer through the MAIN MENU

under LOADING.  After the drawer opens, identify the type of specimen, such

as microsample, with F8, or stat, with F12.  Identify the sample by bar coding

or manually entering on the keyboard the patient identification number and then

placing the sample into the drawer.

E.   	In MANUAL MODE, the operator must order the test from the menu, and then

press F10 to save.

F.  	In AUTO MODE, the STA/STA Compact will automatically order the test if it is 	 	selected in the AUTO MODE profile.

G.   	As soon as the sample drawer closes, the TEST STATUS screen will appear.

If there is not enough reagent(s) to run the test, the suspect reagent(s) will appear

in red with the amount of deficiency.  This deficiency will BLOCK the sample

pipettor.  When this occurs, add the deficient reagent(s) to run the samples.

H.   	All dilutions of the calibrators, controls and patients’ samples are automatically

prepared by the STA/STA Compact according to the parameters entered in the

Test Set-up.  If the patients’ results fall outside the assay reportable range, the

STA/STA Compact automatically re-tests the sample in question at an appropriate

dilution provided that the option has been entered in the Test Set-up.

I.   	All patient results are displayed on the TEST PANEL screen and automatically

print out and transmit if selected.

J.   	For results in question, that need operator intervention, cursor to the identification

number in the TEST PANEL screen and enter.  This will bring up the FILE

PROCESSING screen.  Follow the options in the left-hand corner of the screen,

i.e. re-run test.

- VIII.	CALCULATIONS:

A.	The STA Compact automatically converts the results off of a standard curve

(lin-lin) to mg/L.

B.  	The STA Compact uses a dilution of 1:2, sample to buffer.

IX.	REFERENCE RANGE:

A.	The therapeutic range for Arixtra (fondaparinux), determined on a specimen approximately 3 hours post dose, is 1.20-1.26 mg/L per manufacturer’s information.

B.	Prophylactic range is 0.39-0.50 mg/L per manufacturer’s information.

C.	Therapeutic / prophylactic levels needed may change due to clinical 	considerations specific to the patient’s condition.

D.	Detection threshold for the assay is 0.10 mg/L

- X.	REPORTING RESULTS:

A.	The results for the Heparin Assay are reported out to two decimal places

(Example:  0.50 mg/L).

B.	For results below the detection threshold of the assay, manually enter “0” mg/L

XI.	NOTES:

A.   	The assay’s linearity extends to 1.00 mg/L.  In order to determine a result on a

sample with a higher level, the operator must:

1.	Mix the patient sample 1:1 with Normal Plasma (as the only source of ATIII, which is required for this assay)

2.	Re-test the heparin assay on this mixture

3.	Multiply the result x2 to correct for the dilutional difference.

- XII.	LIMITATIONS OF THE PROCEDURE:

A.	The lab must know which heparin or LMWH is being administered.

The draw time should also be noted, as therapeutic levels are based upon a sample 		drawn approximately 3 hours post dose.

B.	Any release of platelet factor 4 (PF4), which is a potent heparin inhibitor, will lead to an under-estimation of the heparin level in the plasma being tested.  Careful and adequate centrifugation is essential: the higher the level of residual platelets, the greater the risk of PF4 release.

C.	Icteric plasma (bilirubin >6.6 mg/dL) may render result lower than it should be, due to intense yellow coloration; the result, therefore, may be invalid for monitoring purposes.

D.	Lipemic plasma may be ultracentrifuged prior to use on the instrument. Hemolyzed plasma may be run, but reported with a comment.

XIII.	REFERENCES:

1.   	STA- Rotachrom HBPM/LMWH Colorimetric Assay of Low Molecular Weight 			Heparin

(LMWH) by STA Analyzers.  Package insert 26377-Revised March 1997.

2.   	STA-Hepanorm HBPM/LMWH, Calibration Plasmas for Assays of Low 				Molecular

Heparin (LMWH) using anti-Xa Method on STA. Package insert -Revised 	October 2000.

3.	STA-Quality HBPM/LMWH Control Plasmas for Assays of LMWH using Anti-	Xa Method on   STA Compact. Package insert—revised October 2000.

4.	STA/STA Compact Operators Manual.  Revised 3/30/98

5.	Warkentin, TE Heparin-Induced Thrombocytopenia: Towards Consensus.

6.	Fondaparinux sodium package insert, 12/2001, Organon.

XIV.	HISTORY:

H-1	This procedure was written by N. Kriz on 3-13-2007.

H-2	This procedure was reviewed by N Kriz on 10/22/2010.

H-3	This procedure was reviewed by N Kriz on 8/23/2011.

H-4	This procedure was reviewed by N Kriz on 6/26/2012.

H-5	This procedure was reviewed by N Kriz on 6/12/2013.