- PURPOSE:

  - To define what test will be ordered on HPC specimens

  - To describe how specimens are to be accessioned

  - To define where specimens should be forwarded

II.	POLICY:

## Specimens on HPC will be delivered to Department of Laboratory Medicine

Specimen - Receiving / processing area. Specimens will consist of 1 tube for flow 	cytometry, 2 tubes (one concentrate and one 1:5 dilution) for main hematology. Each 	specimen (s) must be accessioned, appropriate comments entered and forwarded to 	testing areas.

III.	PROCEDURE:

- Each specimen will be labeled with patient information and accompanied by a requisition slip (requisition slip is to be scanned in by specimen processing).

- Dependent on the sample type the following tests are ordered

- HPCA: Hematopoietic Progenitor Cells, Apheresis (1:5 dilution)

- HPCCB: Hematopoietic Progenitor Cells, Cord Blood

- HPCM: Hematopoietic Progenitor Cells, Marrow

  - 1 tube of concentrate and 1 dilution (if applicable) x5 is delivered to Sapphire area

F.  Specimen analysis:

- Scan the order number of the patient sample in Open tube mode

- Run 1:5 dilution specimen through Sapphire (For HPCCB and HPCM samples run     straight concentrate tube (no calculation needed).

- Print result sheet

- Multiply the 1:5 diluted samples (HPCA) x 5 for the WBC and platelet counts with a calculator.

- In instrument menu Shift Press F4 on instrument tab in Soft to withhold all results.

- Enter the WBC and PLT results in Lab results tab: (only enter the calculated WBC and PLTI if the 1:5 dilution was run) and verify all (man diff is automatically added).

- Forward concentrated specimen and result slip to smear area

- Make smear from concentrated specimen.

- Perform 200 cell differential from smear of undiluted specimen

- Verify results into computer unless abnormal cells seen and differential needs to be reviewed by supervisor.

IV.	HISTORY:

H-1	This procedure was written by N. Ortoli-Drew on 6-4-2003.

H-2       This procedure was revised by N. Ortoli-Drew on 3-27-2010.

H-3       This procedure was revised by N. Ortoli-Drew on 5-2011.

H-4       This procedure was revised by N. Ortoli-Drew on 4/2012.

H-5	This procedure was revised by Elizabeth Degano on 12-17-12

H-6	This procedure was revised by Elizabeth Degano on 7-8-13.